## 中文
序号 | 分类编码 | 产品名称 | 产品描述 | 类别 | 备注 | 新
1 | 0 | 一类医疗器械目录中灭菌后升二类的医疗器械产品 | 一类医疗器械产品目录中灭菌后升为二类的产品，灭菌后不产生新风险或灭菌后风险程度降低的，可免于进行临床评价。 | Ⅱ
2 | 02 | 无源手术器械 | 02子目录下Ⅱ类无源产品均免于进行临床评价。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型作用机理、药械组合产品除外；新型结构设计不产生新风险或风险程度降低的，可免于进行临床评价。 | Ⅱ
3 | 03 | 神经和心血管手术器械 | 03子目录下Ⅱ类无源产品均免于进行临床评价。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型作用机理、药械组合产品除外；新型结构设计不产生新风险或风险程度降低的，可免于进行临床评价。 | Ⅱ
4 | 11 | 医疗器械消毒灭菌器械 | 11子目录下Ⅱ类产品均免于进行临床评价。 | Ⅱ
5 | 16 | 眼科器械 | 16子目录下Ⅱ类无源产品均免于进行临床评价。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型作用机理、药械组合产品除外；新型结构设计不产生新风险或风险程度降低的，可免于进行临床评价。 | Ⅱ
6 | 18 | 妇产科、辅助生殖和避孕器械 | 18子目录下Ⅱ类无源产品均免于进行临床评价。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型作用机理、药械组合产品除外；新型结构设计不产生新风险或风险程度降低的，可免于进行临床评价。 | Ⅱ
7 | 01-02-02 | 医用激光光纤 | 由激光器连接头、光纤传输体及应用端组成，与激光器连接后，用于传输激光,供激光手术用。光纤出光端为直射平切端面。 | Ⅱ
8 | 01-02-02 | 一次性使用鼻腔内照射光纤头 | 由激光器连接头、光纤 、病人端组成，病人端为U型。与适用的激光器连接后，将激光传输至鼻腔，对鼻腔内的毛细血管进行照射。 | Ⅱ
9 | 01-03-01 | 高频电灼器 | 高频电灼器由主机、多种治疗器、附件组成，通常额定频率、输出功率远低于高频手术设备，可按设计、型式、技术参数、预期用途等不同分为若干型号；仅用于皮肤科、耳鼻喉科、妇科和肛肠科浅表部位的手术中，对相应组织进行凝固、使组织变性和/或坏死。 | Ⅱ
10 | 01-03-01 | 高频手术设备 | 通常包括高频发生器主机及其他相关附件（脚踏开关、手术电极、中性电极等）；利用高频电流（频率范围200k—5MHz）直接流经人体产生热效应从而对人体组织进行常规切割和凝血；用于传统电外科手术中对人体组织进行常规切割和凝血。高频发生器频率范围应仅限于GB 9706.4中所规定范围200k—5MHz；适用范围仅限于常规的组织切割和凝血手术（普通外科、耳鼻喉科、神经外科、妇科、腹腔镜手术等）；不包括特殊临床应用或使用方式，如等离子切割凝血、大血管闭合等。 | Ⅲ
11 | 01-03-03 | 无菌管路高频连接仪 | 无菌管路高频连接仪采用高频辐射加热方法使两个高分子制无菌管路实施无菌连接；仪器通常由高频发生器、输出和控制器组成，借助模具，可热合不同管径的无菌管路；供临床现场热合输液管路等无菌管路用。 | Ⅱ
12 | 01-03-03 | 氩气控制器 | 包括控制器和氩气喷笔等附件；通过气体减压阀及氩气控制模块将高纯氩气瓶中的高压氩气转化为工作需要的可调节低流量氩气，与高频刀合用成为氩气刀；与高频电刀配合使用，在电外科手术中实现氩气环境下的凝血，也可用于切割模式下隔绝空气（不具有氩气电离效果）。 | Ⅱ
13 | 01-03-04 | 等离子手术电极 | 等离子手术设备的双极附件。通常由电极头、电极杆、手柄、电缆线、连接器、液体管路组成。与等离子手术设备主机及灌注泵配合使用，在开放手术或内窥镜手术下，适用于在生理盐水或其它电解质溶液环境下，对软组织进行切割、凝血操作。产品性能指标采用下列参考标准中的适用部分，如：GB 9706.4医用电气设备 第2-2部分：高频手术设备安全专用要求、YY/T 1409-2016等离子手术设备。 | Ⅱ/ Ⅲ |  | 新
14 | 01-03-04 | 一次性使用高频手术设备用阴极板 | 一次性使用高频手术设备用阴极板通常由隔离纸、导电胶、铝箔、基衬（如聚乙烯发泡或无纺布涂丙烯酸压敏胶）和电极导线组成，按照导电介质（金属电极导电/离子导电胶导电）、使用人群（成人型/儿童型）、电极数量（单极/双极）及形状（方型/椭圆型）、是否带导线等分为不同的型号和规格；连接人体和高频手术设备，提供低电流密度的高频电流回路，防止人体灼伤用；双极中另一极供高频手术设备报警回路用。产品性能指标采用下列参考标准中的适用部分，如：GB 9706.4医用电气设备 第2-2部分：高频手术设备安全专用要求。 | Ⅱ
15 | 01-03-04 | 高频手术附件 | 高频手术设备的附件，通常由手术手柄、手术附件的电缆、手术连接器和手术电极组成，可含手控开关、可附带降温供水系统。附件可按工作原理、降温等附加功能、技术参数、预期功能等不同分为若干类别、型号和规格；可以以灭菌/非灭菌形式提供；包括单极和双极形式。与高频手术设备配套，供临床各类开放性外科手术中分别用于组织切割、分离、电凝止血、组织凝固等。不包括等离子手术电极和用于血管闭合的闭合器械。产品性能指标采用下列参考标准中的适用部分，如：GB 9706.4医用电气设备 第2-2部分：高频手术设备安全专用要求。 | Ⅱ | 修改产品描述。
16 | 01-03-04 | 高频内窥镜手术器械 | 用于在内窥镜下完成手术操作的高频电极，属于高频手术设备的应用部分；在内窥镜手术下，通过内窥镜器械孔道或其他器械通道进入人体；在内窥镜手术中用于对人体组织进行常规切割和凝血。 | Ⅲ
17 | 01-03-04 | 射频消融用针状电极 | 射频消融用针状电极通常配合穿刺针、套管等成套使用，属于射频消融设备的应用部分。消融针上含有测温传感器；把射频发生器产生的高频电流传递至指定的人体组织，从而实现对其消融/凝固/坏死等目的；与射频发生器（射频控制器）配合使用，用于人体组织的消融。 | Ⅲ
18 | 01-03-04 | 电容式中性电极 | 产品通常由外接导线、绝缘膜、高分子凝胶和导电网组成。与高频手术设备配合使用，利用电容结构原理设计，在患者和电极间形成一个交变电场，使流经人体的高频电流以场的形式通过电容，为高频电流提供返回通道，以防止在患者组织中产生过高的温升或不希望的灼伤。 | Ⅱ
19 | 01-03-05 | 灌注泵 | 一般由主机、控制器和电缆组成，与射频消融设备配套使用，用于降低消融区域与患者接触部分的温度。 | Ⅱ
20 | 01-07-03 | 手术室设备整体控制系统 | 通常由主机、控制模块、连接第三方设备的通讯装置、扩展装置、连接线缆组成。用于对手术室内相关设备（包括医疗器械和非医疗器械）进行集中控制和显示。系统不改变被控设备的原有结构功能，所实现的控制功能和显示参数不超出被控设备自身原有功能和参数调节范围。 | Ⅲ
21 | 01-08-00 | 一次性使用脊柱外科照明光纤 | 产品由光源连接头、光纤、头端组成。以无菌形式提供。与照明光源连接后，向手术部位传输光能，为脊柱微创手术提供照明。 | Ⅱ
22 | 01-08-01 | 手术无影灯 | 手术无影灯由固定座、横梁、平衡体、平衡弯管、灯头弯管、灯头、调光手柄、拉手组成，可按光源、结构型式、技术参数、配置、附加功能等不同分为若干型号，供手术照明用。产品性能指标采用下列参考标准中的适用部分，如：YY 0627-2008 医用电气设备 第2部分：手术无影灯和诊断用照明灯安全专用要求。 | Ⅱ
23 | 01-09-01 | 内窥镜手术动力系统 | 产品一般由主机、手柄、刀头组成，通过电机马达带动手柄及刀头进行机械运动，用于内窥镜手术中实现绞碎或切除组织等手术功能（鼻腔部位手术除外）。适用或参考行业标准YY/T 0955《内窥镜手术设备 刨削器》。 | Ⅱ
24 | 01-09-01 | 鼻窦镜手术动力系统 | 产品一般由主机、手柄、刀头组成，通过电机马达带动手柄及刀头进行机械运动，用于鼻窦镜手术中实现切除鼻腔部位组织等手术功能。适用或参考行业标准YY/T 0955《内窥镜手术设备 刨削器》。 | Ⅱ
25 | 01-10-00 | 外科术前备皮器 | 一般由电动手柄主机和电池充电器组成。与一次性使用的刀片联合使用，用于外科手术前去除患者身体和头部的毛发。 | Ⅱ
26 | 01-10-04 | 手术动力系统 | 通常由主机、控制装置、电动马达、手柄和各类切割器组成。用于开放性手术时切割/切开、削磨、钻孔、敲击等操作。 | Ⅱ
27 | 01-10-05 | 植皮制网机 | 由制网座、滚刀组件、锁紧组件、无菌载片等组成，通过外接的有源器械驱动，用于烧伤治疗扩展皮片。不同规格的载片可实现皮片不同比例的扩展。 | Ⅱ
工作原理：
将待扩展的皮片放在无菌载片上，载片放置于制网座上，然后驱动连接于本产品的外部有源器械（电机），使电机动力传递到本产品上，最终使本产品的滚刀部件以旋转运动的方式实现对载片的碾压。
28 | 01-10-05 | 植皮机 | 通常由主机和手持件组成。用于皮肤的取皮、植皮或擦皮。 | Ⅱ
29 | 02 | 重复使用内窥镜手术器械 | 重复使用内窥镜手术器械用不锈钢制成，由手柄、钳口、剪刀头等组成；可按材质、设计、技术参数、工作部位、预期用途等不同分为若干型号及规格；非灭菌提供，可重复使用；配合宫腔镜、胸腹腔内窥镜、肾镜及膀胱镜等，通过自然或手术通道进行抓取、咬切、分离组织或异物等用。 | Ⅱ
30 | 02 | 胸腹腔内窥镜用手术器械 | 胸腹腔内窥镜用手术器械为内窥镜下无源组件，用聚砜树脂、不锈钢等材料制成，由管路、穿刺针、穿刺针套管、钳子、打结推杆、肌瘤螺杆、探杆、组织牵开器、棉签施用器、扩张器、手术钩、冲洗装置、封帽、连接器、适配器等组成；可以材质、设计、技术参数、工作部位、预期用途等不同分为若干型号及规格；可以以灭菌/非灭菌形式提供；与胸腹腔内窥镜配合使用,用于患者胸腹腔疾病的诊断和治疗。 | Ⅱ
31 | 02 | 一次性使用颅内吸引管 | 一次性使用颅内吸引管可采用聚氯乙烯等适用材料制成，可设计为单管或套管，管的头端及侧管有光滑圆孔，可按型式、管径等不同分为若干型号和规格，以无菌形式提供；供脑外科手术吸取颅内体液、血液、冲洗液等，清理手术视野用。 | Ⅱ
32 | 02 | 吸引/灌流系统（管） | 本产品可由手柄、吸引灌流管、连接器，密封帽，备用阀管，密封圈等部件组成。本产品经灭菌，一次性使用。与吸引灌流控制主机配合，用于胸腔和腹腔镜手术时进行抽吸和冲洗。 | Ⅱ
33 | 02 | 无源内窥镜下非骨科手术器械类产品 | 无源短期(≤24小时)使用器械，不与有源设备配合使用，市场应用多年，非创新功能设计，非新材料制成，无涂层，无降解、吸收功能，不含药物、动物源性成分，用于辅助完成内窥镜下检查、导引、取出、套扎（吻合器除外）、造影、扩张，夹取、分离、切割组织或取石/碎石。 | Ⅱ
34 | 02 | 肾扩张器/扩张鞘 | 产品一般由聚四氟乙烯、聚乙烯、聚氯乙烯、聚丙烯等材料制成，可按材质、设计、技术参数等不同分为若干型号及规格，可带探针或不带探针。以无菌形式提供，一次性使用。供泌尿外科手术中，建立经皮入路。 | Ⅱ
35 | 02 | 内窥镜连接桥 | 产品材料为不锈钢，为泌尿科内窥镜的附件。用于连接内窥镜和管镜,为手术器械提供通道进行手术治疗或进行观察诊断用。 | Ⅱ
36 | 02 | 内窥镜闭孔器 | 产品由不锈钢制成，插入镜鞘，配合镜鞘使用。泌尿及妇科内窥镜诊断和手术中在内窥镜镜鞘插入人体时封闭镜鞘前端，避免损伤人体组织。 | Ⅱ
37 | 02 | 消化道镜螺旋外管（不含润滑剂） | 本产品由一个末端带有螺旋、顶端带有把持柄的软管组成。软管的顶端还附带旋转领及连接组件。辅助内窥镜插入并推进到目标部位。 | Ⅱ
38 | 02 | 套筒式皮肤扩展器 | 由套筒、内芯和罩筒组成。产品可根据伤面的大小和部位的不同，与牵引线配合，对创面周围的皮肤进行拉伸延展，使创缘距离逐渐缩小或靠拢。 | Ⅱ
39 | 02 | 冲洗管路 | 该产品由冲洗管路组成，用于在手术中进行冲洗和抽吸。 | Ⅱ
40 | 18-01-10 | 宫腔球囊导管 | 该产品由球囊、导管、导管座、止回阀、接头、支撑杆等组成，用于将造影剂注射到子宫和输卵管，进行子宫输卵管造影，或在腹部手术中注射美蓝用于输卵管成形术和/或输卵管吻合术。 | Ⅱ | 修订分类编码。
41 | 02 | 一次性腹腔镜软器械鞘管 | 一次性腹腔镜软器械鞘管（简称软鞘管）是一种单切口腹腔镜手术用器械，可由切口保护套和多通道密封体等部件组成。切口保护套起牵开、保护切口以外卡环为基座，多通道密封体以盖帽形式和基座密封嵌合，密封体的阀门通道作为单孔腹腔镜手术用腔镜、钳、剪等器械进出切口的密封操作通道。产品经灭菌，一次性使用。产品性能指标可采用下列参考标准中的适用部分，如：YY 0672.1-2008内镜器械第1部分腹腔镜用穿刺器。 | Ⅱ
42 | 02 | 一次性使用内窥镜护套 | 一次性使用内窥镜护套可由复式管、塞子、接头、夹子、法兰、喷洒头、带定距管口(镜片)、硅胶管封等部件组成。产品经灭菌，一次性使用。用于上呼吸道、声带或鼻咽部进行柔性内窥镜检查过程中，为使用的内窥镜提供一个无菌的一次性防护层。 | Ⅱ
43 | 02 | 一次性使用切口保护套 | 一次性使用切口保护套一般采用高分子材料制成。可由外环、置入环及通道等部件组成。产品可分为定高型和变高型等形式，按尺寸不同分为若干规格。适用于内窥镜手术及小切口手术，扩展切口术野，保护切口免受损伤，减少切口感染。 | Ⅱ
44 | 02 | 医用冲洗头 | 本产品可由冲洗头座、冲洗头等部件组成，不含冲洗液。采用高分子材料制成。本产品为灭菌提供，一次性使用。用于鼻炎患者的鼻腔冲洗，耵聍、中耳炎患者的外耳道冲洗。 | Ⅱ
45 | 02 | 支架回收器 | 用于在内窥镜直视或X线下捕获和回收输尿管支架。 | Ⅱ
46 | 02 | 外科组织袋和导入器 | 由外科组织袋和导入器组成。外科组织袋由组织袋和拉线组成，导入器由导入器插脚和把手组成。产品经灭菌，一次性使用。外科组织袋用于手术取出和/或粉碎组织过程中或之前分离组织；导入器用于内窥镜外科手术中辅助导入外科组织袋。 | Ⅱ
47 | 02 | 三通旋塞 | 本产品由旋塞座、旋钮和螺旋式接口组成。无菌、无热原。一次性使用。如产品具有特殊性能或结构等不能豁免临床。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
48 | 02-01-01 | 一次性使用手术刀片 | 一次性使用手术刀片采用碳素工具钢或合金工具钢等适合制作刀片的材料制成，可含刀柄；刀片根据型式不同分若干种，每种可按配合刀柄和刀片尺寸等不同分为若干规格；产品以无菌形式提供；供切割软组织用。产品性能指标采用下列参考标准中的适用部分，如：YY 0174-2005手术刀片，GB 8662－2006手术刀片和手术刀柄的配合尺寸。 | Ⅱ
49 | 02-01-03 | 一次性使用无菌备皮刀 | 通常由刀片、刀架和保护盖组成。无菌提供。一次性使用。用于术前准备，去除毛发。 | Ⅱ
50 | 02-03-03 | 一次性使用内窥镜用结扎线剪 | 一次性使用内窥镜用结扎线剪用聚氯乙烯、聚苯砜、不锈钢等制成，通常由手柄、管身、刀头/剪刀收纳部组成；可按材质、设计、技术参数、适用部位等不同分为若干型号及规格；可以灭菌或非灭菌形式提供；与内窥镜配套，供剪断结扎线头等用。 | Ⅱ
51 | 02-04-09 | 内窥镜抓钳 | 内窥镜抓钳可由三爪钳、鞘管和控制手柄等部件组成。本产品无菌状态提供，一次性使用。在内窥镜手术时抓取和夹持组织用。 | Ⅱ
52 | 02-04-09 | 内窥镜用钳 | 通常由头部、杆部或软性导管和手柄组成，通过手柄操作传递、控制头部工作。一般头部采用不锈钢材料制成。手术中在内窥镜下操作，用于钳夹、钳取、分离组织，钳取异物，夹持器械。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
02-04-10
02-04-11
02-04-12
02-04-13
53 | 02-04-10 | 一次性内镜用软管式活组织取样钳 | 一次性内镜用软管式活组织取样钳主要由钳头部分、弹簧外管和操作柄部分组成，可有定位针，可按材质、设计、尺寸等不同分为若干型号及规格；以无菌形式提供；供消化道、呼吸道等内窥镜下活组织取样或钳取和清除异物用。产品性能指标采用下列参考标准中的适用部分，如：YY/T1076－2004内镜用软管式活组织取样钳。 | Ⅱ
54 | 02-04-12 | 内窥镜用异物钳 | 内窥镜用异物钳用不锈钢和塑料等制成，由钳头、外套管、连杆、钳柄组成，钳头类型分为鳄嘴头、锯齿头、麦粒头、取珠式、新锯齿头、杯型头、取笔套式、反二爪、花生米式、三爪等，可按材质、设计、技术参数等不同分为若干型号及规格；可以灭菌或非灭菌状态提供；可与气管镜、食管镜及喉镜等配套，供气管或上消化道等钳取异物用。 | Ⅱ
55 | 02-04-13 | 内窥镜弯针持针器 | 内窥镜弯针持针器采用不锈钢、硅胶和树酯等制成，由持针槽、推送杆、控制手柄组成，可按材质、设计、技术参数等不同分为若干种；可以无菌形式提供；供腹腔内窥镜手术中缝合操作用。 | Ⅱ
56 | 02-04-13 | 内窥镜施夹钳及除夹钳 | 本产品可由钳头，钳杆，转盘，钳柄等部件组成，通常为不锈钢材料制造。可重复使用。本产品作为传送装置，用于通过特定尺寸的套管传递或者移除结扎钉/夹。 | Ⅱ
57 | 02-06 | 微创筋膜闭合器 | 产品包括闭合固件和缝线穿引器两部分。闭合固件由器身、传动器、近端翼、手环、窗口组成,器身上有座盘、导引孔、出线口、指握点、控制板组成。缝线穿引器由针和手柄组成。产品灭菌，一次性使用。用于在腹腔镜手术中收拢组织、经皮缝合、以便闭合手术切口。 | Ⅱ
58 | 02-06-01 | 闭合夹 | 由金属、高分子聚合物或其他材料制成，用于外科手术时临时夹闭血管或组织，术后即刻取出。无菌提供，一次性使用。 | Ⅱ | 修订产品名称。
59 | 02-06-01 | 闭合夹 | 用于标记、止血或闭合管状组织结构。可带有输送器。术后不取出。产品与已获准境内注册产品具有等同性，产品组成材料成熟。可吸收材料、新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ | 修订产品描述。
60 | 02-07-01 | 医用缝合针 | 通常外形呈直形或弧形，针尖有圆、三角、铲形状，针尾带孔。一般采用不锈钢材料制成。用于缝合组织、皮肤。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0043医用缝合针。 | Ⅱ
61 | 02-07-01 | 一次性使用荷包缝合针 | 一次性使用荷包缝合针用金属和缝线制成，由两根缝针、一根非吸收性缝线组成。产品以无菌形式提供，适用于消化道吻合手术中的荷包缝扎。产品性能指标参考YY 0877荷包缝合针标准中的适用部分。 | Ⅱ
62 | 02-07-02 | 一次性使用无菌点刺针 | 一次性使用无菌点刺针采用适用的不锈钢制成，由扁平的柄部和三角形针尖两部分组成；供速发型超敏反应皮肤测试时点刺刺破皮肤，使皮肤上的抗原液渗入皮下用。不包括针内含抗原液的点刺针。 | Ⅱ
63 | 02-07-02 | 一次性使用内窥镜气腹穿刺套管针 | 一次性使用内窥镜气腹穿刺套管针用适用的ABS等高分子材料制成，由外套管及可配套锁止的穿刺内芯组成；可按材质、设计、尺寸等不同分为若干型号及规格；为内窥镜配套用手术器械，通过在腹壁穿刺并固定，供形成气腹注气通道和腹腔镜或其他手术器械的工作通道用。 | Ⅱ
64 | 02-07-02 | 皮肤点刺盒 | 由聚甲基丙烯酸甲酯聚合物等制成。由测试头和盘组成。用于过敏人群过敏原的测试工具。 | Ⅱ
65 | 02-07-02 | 气腹针 | 由针管、握把、活塞、锁定接头、管芯针等组成，不包含Ⅲ类医疗器械组件。用于腹部内视镜手术过程中建立气腹。基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括：使用新型材料，包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等的产品。 | Ⅱ
66 | 02-07-03 | 乳腺组织标记物 | 通常由推送工具（如：套管针）和标记物组成。标记物一般采用不锈钢、钛、或钛合金材料制成，可包含有辅助定位的的高分子材料。无菌提供，一次性使用。标记物接触人体组织时间超过30天。用于乳腺组织定位标记。新材料、新设计、新功能的产品除外。 | Ⅲ |  | 新
67 | 02-07-03 | 一次性使用乳腺定位丝及其导引针 | 产品中的导引针针管应使用GB 18457中的材料制成，定位丝宜使用不锈钢或其他合金材料制成，临床上使用本产品为了便于手术切除乳腺局部微小病灶，在X射线或超声下用乳腺定位丝及其导引针对病灶进行定位，给临床手术医生以明确的引导，以缩小切口，减少手术损伤。产品性能指标参考YY/T 0880 一次性使用乳腺定位丝及其导引针中的适用部分。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
68 | 02-07-04 | 一次性内窥镜吸引活检针 | 通常由头部、杆部或软性导管和手柄组成，头部为针形，通过手柄操作传递、控制头部工作。一般头部采用不锈钢材料制成。无菌提供。豁免情况不包括使用了新材料、新作用机理、新功能的产品，以及用于注射、定位、治疗的产品。 | Ⅱ
69 | 02-08-02 | 内窥镜用组织拉钩 | 内窥镜用组织拉钩采用不锈钢和铝等材料制成，由工作端(拉杆)、器械杆、握把等组成，可按材质、设计、技术参数等不同分为若干型号；以无菌形式提供；供各种内窥镜手术过程中组织的暂时性牵拉。 | Ⅱ
70 | 02-09 | 一次性使用皮肤刮匙 | 由刮匙头、手柄构成，可有防护帽。刮匙头由不锈钢或其他无毒性金属材料制成，手柄、防护帽由塑料材料制成；可按刮匙头内径、手柄形状分为多个规格型号；产品经灭菌，一次性使用；适用于外科手术时刮除坏死组织、皮屑。 | Ⅱ
71 | 02-09-02 | 内窥镜刮匙 | 本产品可由先端部、插入部、手柄等部件构成。产品以无菌形式提供，为一次性使用。经内窥镜插入体腔内，通过本产品的进退操作及刮匙先端部的弯曲操作，用于在体腔内采集组织或细胞。 | Ⅱ
72 | 02-11-00 | 内窥镜腹腔外腔扩张器 | 产品可由管身、扩张囊、管座、密封圈、三通单向注气阀（座）、导入柱、止流夹及护帽等部件组成。用于腹腔镜微创手术前，在腹膜外腔与手术部位中间建立一个手术所需的操作空间。 | Ⅱ
73 | 02-11-03 | 一次性使用尿道扩张器 | 通常为细长设计，由硬质导丝和软质导管组成。无菌提供，一次性使用。用于扩张尿道。 | Ⅱ
74 | 02-11-03 | 食道贲门狭窄扩张器 | 食道贲门狭窄扩张器由扩张条（有X射线显影标志）、导引钢丝、安全软弹簧等组成，扩张条采用软聚氯乙烯塑料制成，导引钢丝采用镍钛合金制成，可按材质、设计、直径/长度等不同分为若干型号和规格；可以以灭菌/非灭菌形式提供；供内窥下食道、贲门狭窄扩张治疗用。 | Ⅱ
75 | 02-12-01 | 一次性使用腹腔镜用穿刺器 | 一次性使用腹腔镜用穿刺器可用合适的金属材料和高分子材料制成，典型结构可包括穿刺针、穿刺套管、注气阀、阻气阀、密封帽构成；可按材质、设计及技术参数等不同分为若干型号及规格；产品以无菌形式提供；供腹腔镜检查和手术过程中，对人体腹壁组织穿刺，建立腹腔手术的工作通道用。产品性能指标采用下列参考标准中的适用部分，如：YY 0672.1-2008内镜器械 第1部分 腹腔镜用穿刺器。 | Ⅱ
76 | 02-12-01 | 腹腔镜穿刺器 | 本产品一般可由穿刺套管和穿刺芯等组成。产品经灭菌，一次性使用。用于与内窥镜配套使用，在内窥镜手术中对人体组织进行穿刺并建立腹腔通道用。豁免情况不包括：使用新型材料；使用非DEHP增塑剂；包含功能性高分子、药物、生物制品等特殊涂层的产品，包含活性成分、新技术、新设计或具有新作用机理，以及具有特殊结构功能、适用范围等的产品。 | Ⅱ
77 | 02-12-02 | 一次性使用皮肤组织钻孔器 | 由环钻头和手柄构成，可有防护帽。无菌，一次性使用；适用于皮肤组织钻孔。 | Ⅱ
78 | 02-12-03 | 泌尿道用导丝 | 泌尿道用导丝由导丝和推管组成，导丝的材质和形状各异：如金属导丝、金属弯头导丝、PE导丝、PTFE导丝、PTFE弯头导丝、PTFE涂覆导丝、PTFE涂覆弯头导丝等，导丝可涂覆亲水/疏水涂层；可按产品材质、设计、技术参数、预期用途等不同分为若干型号及规格；以无菌形式提供；内窥镜下与J型导管和微创扩张引流套件配套使用，起支撑、引导作用。 | Ⅱ
79 | 02-12-03 | 一次性使用输尿管导引鞘 | 一次性使用输尿管导引鞘用尼龙、聚四氟乙烯、不锈钢、聚乙烯、聚丙烯等制成，通常由管身涂有亲水涂层的鞘管和扩张器组成，也可包含导丝，鞘管可为单腔或双腔，可按材质、设计、技术参数等不同分为若干型号及规格；以无菌形式提供；供泌尿外科手术中，建立内窥镜等器械进入泌尿道的通道用。 | Ⅱ
80 | 02-12-03 | 一次性使用膀胱造瘘管 | ：一次性使用膀胱造瘘管通常由造瘘管、导引针和连接管组成；以无菌形式提供；供临床暂时性尿流改道，耻骨上膀胱穿刺造瘘，引流尿液用。 | Ⅱ
81 | 02-12-03 | 非血管用支架推送器 | 用于内窥镜下推送输尿管支架或其他非血管用支架。 | Ⅱ
82 | 02-12-05 | 内窥镜用导引器 | 内窥镜用导引器通常由头部、杆部或软性导管和手柄组成，头部为导引件，通过手柄操作传递、控制头部工作。手术中在内窥镜下操作，用于引导器械，进入腔道或组织。 | Ⅱ
83 | 02-12-05 | 内窥镜导引管鞘 | 本产品可由鞘管部、主体部和密封帽（如有）等部件组成。本产品插入穿刺器或体腔，作为内镜插入通路的管鞘，防止内镜碰触穿刺器先端部和穿刺器内部，保护内镜先端部及弯曲部。 | Ⅱ
84 | 02-12-05 | 消化内窥镜用一次性导丝 | 产品可由导丝本体和保护器等部件组成。通过消化道内窥镜进入腔内，用于引导各种导管到指定位置。 | Ⅱ
85 | 02-13-00 | 垫片/带垫片的不可吸收缝合线 | 配合不可吸收缝合线使用的垫片或带垫片的不可吸收缝合线，在软组织或器官缝合中使用，用于缝合线的支撑，防止缝合线割裂人体组织。垫片的材料为聚四氟乙烯、膨体聚四氟乙烯、聚酯等。新型结构设计、新材料等除外。 | Ⅲ | 新增 | 新
86 | 02-13-01 | 一次性使用自动荷包缝合器 | 一次性使用自动荷包缝合器采用金属、高分子材料、非吸收外科缝线制成，由夹座、钉座、钉和荷包线组成，按材质、设计、技术参数等不同分为若干型号与规格；以无菌形式提供；供消化道开放手术中作荷包缝合用。 | Ⅱ
87 | 02-13-01 | 一次性使用包皮切割吻合器 | 一次性使用包皮切割吻合器采用不锈钢和ABS等材料制成，缝合钉可用不锈钢或适用的金属制成，通常包括环形切割刀、定位弹簧、金属垫片、固定销及活动连片，316L不锈钢为材料的缝合钉，以及以ABS为材料的凹形座、环形刀砧、钉仓保护盖、钉仓、顶体、主体外壳、移动器、主体固定圈、定位调节螺母、主体上下盖、活动手把、保险销。产品应无菌。适用于临床包皮切割缝合手术。 | Ⅱ
88 | 02-13-01 | 一次性使用皮肤缝合器 | 一次性使用皮肤缝合器一般由推片、壳体、手柄、缝合钉组成，可选配拆钉器；可按照设计、缝合钉数量、缝合钉成形后的外形尺寸等不同分为若干型号及规格；产品以无菌形式提供；供创伤及手术切口的表层皮肤缝合用。 | Ⅱ
89 | 02-13-03 | 内窥镜缝合器 | 由进针装置、缝合针、传动机构、控制手柄等组成，缝合器可带缝线或配合线盒使用，适用于微创含自然腔道手术缝合，在软组织中进行间断缝合或连续缝合。一次性使用或可重复使用。带有可吸收缝线或用于血管缝合的产品除外。。 | Ⅱ
90 | 02-13-07 | 非吸收性外科缝线 | 非吸收性外科缝线可分别由蚕丝、聚酰胺6/66、聚对苯二甲酸乙二酯、聚丙烯、不锈钢丝、钛镍记忆合金丝等制成，可带或不带缝合针，可按材质、股数/编织、涂层、带针、束/团、线径、使用部位等不同分为若干型号及规格；可以灭菌或非灭菌状态提供；供人体组织的缝合、结扎用。产品性能指标采用下列参考标准中的适用部分，如：YY 0167 非吸收性外科缝线。 | Ⅱ
91 | 02-14-01 | 冲吸器 | 通常由冲吸管、管路和连接口组成。用于冲洗组织或吸液。 | Ⅱ
92 | 02-15 | 一次性使用手指操作用手术器械 | 一次性使用手指操作用手术器械由聚碳酸酯套管和不锈钢末端执行器组成，可按设计、技术参数和预期用途等不同分为若干型号和规格；以无菌形式提供；用于结肠直肠、泌尿、妇科等腹腔镜手术(HALS)和开放式手术中，进行抓取、移动、分离、切割及缝合组织用。 | Ⅱ
93 | 02-15 | 一次性隆胸手术用漏斗 | 一次性隆胸手术用漏斗由弹性尼龙材料制成，内表面涂有由聚氨酯和聚乙烯吡咯烷酮共混物构成的亲水涂层；以无菌形式提供；在外科隆胸手术时，将植入物置于本器械含有亲水性润滑涂层的内表面，通过挤压，使植入物轻易通过漏斗的小开口进入预定部位，以此缩小患者的切口。 | Ⅱ
94 | 02-15-01 | 一次性使用包皮套扎器 | 通常由套环、夹紧环、连接装置组成。无菌提供。一次性使用。 用于包皮环切，免缝愈合。 | Ⅱ
95 | 02-15-01 | 一次性使用肛肠套扎器 | 一次性使用肛肠套扎器可用适用的金属、塑料和硅胶等制成，可由枪体（内置弹簧）、枪管、硅胶圈、负压吸引接头、绕线轮、套扎器开关和盖帽组成，可带单个或多个硅胶圈，可有辅助照明装置，可按材质、设计、技术参数、预期用途等不同分为若干型号与规格；以无菌形式提供；适用于各期内痔及混合痔或直肠良性息肉的套扎治疗。 | Ⅱ
96 | 02-15-02 | 缝线推结器 | 缝线推结器用金属等材料制成，由头端和把柄组成，头端有缝线穿入孔，可按材质、设计、技术参数等不同分为若干型号及规格；胸腔手术中，用于将体外缝线结推到体表面或胸腔内。无菌提供。 | Ⅱ
97 | 02-15-03 | 支撑棒 | 由尼龙棒和旋转头组成。用于袢式造口手术，穿过肠系膜支撑肠管在固定位置，阻止肠袢返回腹腔。 | Ⅱ
98 | 02-15-04 | 一次性使用手术电极清洁片 | 一次性使用手术电极清洁片可由（海绵）基层、（SiO2等）精细磨料层、压敏胶层、（钡线）显影线、（硅油）隔离纸构成，按形状、尺寸等不同分为若干规格；以无菌形式提供；供手术时清洁高频手术电极上的烧焦组织及其他附着物用。 | Ⅱ
99 | 02-15-04 | 一次性使用无菌瘘管刷 | 一次性使用无菌瘘管刷可由刷柄、刷毛和丝线组成，以无菌形式提供；供识别、清洁直肠瘘管用。 | Ⅱ
100 | 02-15-04 | 刀头清洁片 | 通常由氧化铝层、聚酯层、双面胶和纸质衬垫等组成。本产品为无菌产品，一次性使用。刀头清洁片用于清洁电外科手术刀笔刀头。 | Ⅱ
101 | 02-15-06 | 一次性使用龟头保护套 | 一次性使用龟头保护套由ABS塑料制成，以无菌形式提供；包皮环切术中套在龟头上，防止电刀切割包皮时避免损伤龟头。 | Ⅱ
102 | 02-15-06 | 失禁肛门塞 | 失禁肛门塞以聚氨酯泡沫为原料，由膨胀栓塞、水溶性薄膜和纱布绳组成，适用于排泄失禁患者或肠胃气失禁者栓塞肛门，延长大便滞留时间。 | Ⅱ
103 | 02-15-08 | 医用皮肤记号笔（含墨水） | 由卷包芯、墨水、涤纶纤维笔头、笔套、笔杆、尾帽和标尺组成，其中墨水由水、酒精、乙二醇、龙胆紫和苯甲酸钠组成。应用于外科手术、放射治疗和皮肤治疗时在病人皮肤（无黏膜及皮肤破损处）上作标记和定位。 | Ⅱ
104 | 02-15-13 | 乳房植入体试模 | 通常外壳为硅橡胶，内填充硅凝胶。无菌提供。用于确定乳房植入体时形状和大小。 | Ⅱ | 修订产品描述。
105 | 02-15-14 | 夹子装置 | 与内窥镜配套使用，用于在消化道内放置夹子。夹子用于内窥镜下的标记、消化道组织的止血。无菌提供。可带有输送器。产品与已获准境内注册产品具有等同性，产品组成材料成熟。可吸收材料、新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
106 | 02-15-16 | 腹腔镜线结推送器 | 腹腔镜线结推送器采用不锈钢等材料制成，由线结固定槽、推送杆和手柄组成，头端有缝线穿入孔，可按材质、设计、技术参数等不同分为若干型号及规格；用于将体外打结缝线放入腹腔镜术区中. | Ⅱ
107 | 02-15-17 | 一次性使用内窥镜用取样刷 | 一次性使用内窥镜用取样刷通常由刷柄、管鞘、牵引丝和刷子等组成，可按材质、设计、技术参数、适用部位等不同分为不同型号及规格；以无菌形式提供；供内窥镜检查或配合导航系统使用时从人体中提取组织样本用。 | Ⅱ
108 | 02-15-17 | 一次性使用防污样本刷 | 一次性使用防污样本刷由管路组件、刷头组件、聚乙二醇和保护塞组成，可按材质、设计、技术参数、适用部位、预期用途等不同分为若干型号及规格；以无菌形式提供；供内窥镜下刷取患者病变部位细胞组织或分泌物用。 | Ⅱ
109 | 02-15-17 | 一次性使用样本刷 | 一般由管路组件、刷头组件和保护塞组成，用于内窥镜下刷取患者病变部位细胞组织或分泌物。 | Ⅱ
110 | 02-15-18 | 一次性使用内窥镜取石篮 | 一次性使用内窥镜取石篮通常由带鲁尔接头的手柄、带软鞘导丝杆（或套管）和远端的 360° 金属丝制自展篮组成，可按材质、设计、技术参数、适用部位、预期用途等不同分为若干型号及规格；以无菌形式提供；供消化、泌尿等诊疗时在内窥镜下抓住、操控和取出结石以及其他异物用。 | Ⅱ
111 | 02-15-18 | 内窥镜下机械碎石/取石辅件 | 内窥镜下机械碎石/取石辅件用金属和高分子材料制成，由外鞘管、碎石手柄、冲洗延长管、接头组成；可按材质、设计、技术参数、适用部位、预期用途等不同分为若干型号及规格；可以灭菌/非灭菌状态提供，可重复使用；供其他内镜取石方法失败的情况下进行机械性碎石/取石用。 | Ⅱ
112 | 02-15-20 | 一次性使用取石球囊 | 产品可由球囊、接头、导管等部件组成，其中导管为多孔腔结构的塑料管。本产品经灭菌，一次性使用。与内窥镜配合使用，用于经十二指肠内窥镜进入胆道，进行胆道取石。 | Ⅱ
113 | 02-15-22 | 一次性使用内窥镜给药管 | 一次性使用内窥镜给药管由不锈钢和聚乙烯、ABS等制成，通常由喷洒头、导管、接头等组成，可包括导丝，可按材质、设计、技术参数、适用部位、预期用途等不同分为若干型号与规格；以无菌形式提供；与内窥镜配合，用于灌洗、喷洒药液、造影、移动胆石/胆砂等用。 | Ⅱ
114 | 02-15-23 | 一次性使用内窥镜用套扎器 | 一次性使用内窥镜用套扎器主要由套筒、触发拉线、控制手柄、冲洗接头、装载导管、套圈组成，可按材质、设计、技术参数等不同分为若干型号及规格；以无菌或非无菌形式提供；手术中在内窥镜下操作，用于套扎息肉、或放入可防止或控制出血的结扎环也用于静脉曲张的结扎。 | Ⅱ | 修订产品描述。
115 | 03-00-00 | 心脏瓣膜安装工具套件 | 心脏瓣膜安装工具套件由推结器、瓣膜探测针、瓣膜安置器组成。推结器用于体外缝线结推到体表面或胸腔内；瓣膜探针用于检验机械瓣膜瓣叶的灵活性；安置器用于放置机械瓣膜。 | Ⅱ
116 | 03-00-00 | 神经外科内窥镜微调装置 | 神经内窥镜微调装置可由微调装置、固定配件、固定臂等部件组成。一般采用符合YY/T0294.1的不锈钢材料制成，非灭菌产品。用于支撑和微调节神经外科手术中的神经内窥镜或穿刺器。 | Ⅱ
117 | 03-00-00 | 经颈静脉肝内穿刺器械 | 通常由通芯针、导管、加硬套管、扩张器、导管鞘、扩张器和导丝组成。用于经颈静脉肝内门静脉穿刺，进行门静脉的肝内分流手术，以降低门静脉压。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
118 | 03-01-01 | 神经外科用刀凿 | 通常由刀头和杆部组成。刀头一般采用不锈钢材料、钻石等制成。用于神经外科手术中的组织切断。 | Ⅱ
119 | 03-02-01 | 脑内用剪 | 脑内用剪由金属制成，由手柄与头部组成，单关节，刀头部可设计成直、上下弯等，柄与头部可成一定的二维或三维角度，以适应不同解剖位置的操作，刀头未端可为尖或钝头，表面光滑，不损伤周围组织；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供脑内手术时剪除病骨或坏死神经/脑组织用。 | Ⅱ
120 | 03-02-01 | 脑膜用剪 | 通常由一对中间连接的叶片组成，头部有刃口，柄部带指圈。一般采用不锈钢材料制成。用于剪切脑膜组织。 | Ⅱ
121 | 03-02-01 | 冠状动脉剪 | 产品由一对中间连接的叶片组成，头部有刃口，柄部带指圈。一般采用不锈钢材料制成。非无菌提供。用于剪切冠状动脉或扩大冠状动脉切口。 | Ⅱ
122 | 03-03-01 | 脑内用钳 | 脑内用钳由金属制成，由柄与头部组成，单关节，钳头可设计成直、左右弯等，柄与钳头可成一定的二维或三维角度，以适应不同解剖位置的操作，钳头未端可为圆或卵圆形或其他形状，表面光滑，不损伤周围组织；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供脑内手术时夹取颅内肿瘤和包膜用。 | Ⅱ
123 | 03-03-01 | 瓣膜手术用钳 | 瓣膜手术用钳用不锈钢制成，用于瓣膜手术中抓取和移除心脏结构上的硬物质，如钙化物质和纤维组织。 | Ⅱ
124 | 03-03-01 | 胸腔心血管外科用钳 | 由一对中间连接的叶片组成，头部为钳喙。一般采用不锈钢材料制成。用于胸腔手术时夹持脏器、血管，游离腔静脉、主动脉血管。 | Ⅱ
125 | 03-03-02/05 | 神经外科脑内用钳 | 通常由钳喙、杆部和柄部组成，钳喙有刃口，柄部带指圈。一般采用不锈钢材料制成。用于钳取、咬除脑部组织、异物、增生物或者摘除肿瘤。 | Ⅱ
126 | 03-03-04 | 无损伤动/静脉阻断钳 | 无损伤动/静脉阻断钳用金属制成，单关节，由柄与头部组成，钳头可设计成直、上下弯等，柄与钳头可成一定的二维或三维角度，以适应不同解剖位置的操作，钳头内侧可有半硬性嵌入物，以增加钳夹稳定性与无损伤性，钳可锁止，可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供涉及大血管手术时动静脉阻断用。 | Ⅱ
127 | 03-04-01 | 脑内用镊 | 脑内用镊由金属制成，由柄与头部组成，头部尖端可设计成尖细或圆形，柄与头部可成一定的二维角度，柄上有防滑凹凸设计，表面光滑，不损伤周围组织；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供脑内手术时夹取敷料、脑内细软组织等用。 | Ⅱ
128 | 03-04-01 | 微创心外科手术用镊 | 微创心外科手术用镊用不锈钢制成，在微创心外科手术中用于夹捏和操作心脏结构。 | Ⅱ
129 | 03-04-01 | 胸腔心血管外科用镊 | 由一对柄部叠合的叶片组成，头部为直形或弯形。一般采用不锈钢材料制成。非无菌提供。用于夹持心肌组织、血管。 | Ⅱ
130 | 03-04-01/02 | 神经外科用镊夹 | 通常由一对尾部叠合的枪型叶片组成。一般采用不锈钢材料制成。用于脑外科手术时夹持细软组织或摘除肿瘤。 | Ⅱ
131 | 03-05-01 | 一次性使用无菌头皮夹 | 一次性使用无菌头皮夹可采用聚苯乙烯树脂或聚甲醛等适用高分子材料制成，U形结构，连接的两夹持臂刃片互相咬合，使用时可经由手动或头皮夹钳放置，非植入器械；产品以无菌形式提供，供临床手术时固定头皮及止血用。 | Ⅱ
132 | 03-05-02 | 止血夹 | 用于脑部手术时夹闭小血管和管状组织，以止血。也可用于腹腔等微创手术。术后不取出。可带有输送器。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
133 | 03-07-01 | 脑内用钩 | 脑内用钩由金属制成，由柄与头部组成，头部尖端可设计成半圆形弧状或光滑过渡的钩状，柄与头部可成一定的二维或三维角度，柄上可有防滑设计，表面光滑，不损伤周围组织；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供脑内手术时牵拉神经或脑内组织等用。 | Ⅱ
134 | 03-07-01 | 主动脉拉钩 | 产品由头部和柄部组成。头部带弯钩。一般采用不锈钢材料制成。非无菌提供。用于牵拉主动脉。 | Ⅱ
135 | 03-08-01 | 脑刮匙 | 脑刮匙由金属制成，由柄与头部组成，头部尖端可设计成匙状，柄与头部通常共轴，柄可有防滑设计，头部表面光滑，不损伤周围组织；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；通常为可重复器械；供脑内手术时刮除病骨及坏死组织用。 | Ⅱ
136 | 03-08-01 | 神经外科用刮匙 | 通常为细长设计，近端有手柄，远端为匙形。一般采用不锈钢材料制成。用于颅脑手术时剥离、刮除病社。 | Ⅱ
137 | 03-09-01 | 脑内剥离器 | 脑内剥离器由金属等制成，由柄与头部组成，头部尖端可设计成宽窄不一的各种形状，柄与头部通常共轴，柄可有防滑设计，头部表面光滑，不损伤周围组织；可按材质、设计、技术参数等不同分为若干型号及规格；通常为可重复器械；供脑内手术时剥离脑膜用。 | Ⅱ
138 | 03-09-01 | 胸腔心血管外科用剥离器 | 可由剥离头、杆部和柄部等部件组成。一般采用不锈钢材料制成。用于剥离软组织，或剥离静脉血管内的血栓等，使静脉畅通。 | Ⅱ
139 | 03-10-01 | 肋间软组织牵开器 | 肋间软组织牵开器由牵开环、牵开壁和胶带组成。牵开环可用聚酯等弹性材料,牵开壁可用聚醚聚氨酯等适用材料制成,胶带的背衬为无纺布,胶粘层为丙烯酸粘胶；可按材质、设计、技术参数等不同分为若干型号与规格；以无菌形式提供；心脏外科手术中放入肋间，牵开组织形成手术操作口。 | Ⅱ
140 | 03-10-01 | 心房牵开器 | 心房牵开器由不锈钢制成，由三角支架、牵开叶片和镊子组成。用于瓣膜手术中，牵开右心房前壁以暴露二尖瓣，便于观察内部和外部心脏结构。 | Ⅱ
141 | 03-10-01 | 一次性使用颅内牵开固定导管 | 通常由引导器和扩张器组成，采用聚碳酸酯高分子材料制成。无菌提供，一次性使用。适用于临床颅内手术中扩开术野周围软组织，在脑部手术中提供一个可视化的工作通道。 | Ⅲ | 新增 | 新
142 | 03-10-03 | 二尖瓣扩张器 | 二尖瓣扩张器用适用的金属材料制成，由带撑开杆的手柄、器身和扩张器头部组成，扩张宽度可调节，可按材质、设计与技术参数等不同分为若干型号与规格；通常可反复使用；供心脏手术时扩张二尖瓣用。 | Ⅱ
143 | 03-10-03 | 血管扩张器 | 通常由头部、杆部和柄部组成，一般可采用不锈钢或钛合金制成，用于心胸外科或血管外科手术中扩大血管口径。产品基本原理、适用范围、性能和结构组成与已上市产品相同。豁免情况不包括：使用新型材料，包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等产品。 | Ⅱ
144 | 03-10-03 | 经皮肾扩张器 | 用于经皮肾盂穿刺时，扩张皮下组织、肌肉、肾实质组织，以建立和保持经皮通路。 | Ⅱ
145 | 14-00-00 | 经皮肾造瘘导管及附件 | 产品由经皮造瘘导管、扩张器、穿刺针、导丝、导管鞘、球囊扩张导管、压力泵和附件组成。用于经皮穿刺肾造瘘术。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅱ | 修订分类编码、产品名称和产品描述。 | 新
146 | 03-11-01 | 打孔器 | 通常由打孔头和柄部组成。无菌提供。一次性使用。接触中枢神经系统或血液循环系统。用于介入手术中在组织上打孔，建立通路。 | Ⅲ
147 | 03-11-01 | 血管打孔器 | 血管打孔器分为一次性使用和反复使用两类。一次性使用血管打孔器用不锈钢和ABS等适用材料制成,以无菌方式提供；反复使用血管打孔器用符合ASTM F899的不锈钢材料制成,非无菌方式提供；两类产品均可按材质、设计、技术参数、适用血管等不同分为若干型号与规格；供外科手术时在动脉/静脉血管壁上打一个圆形开口,以便后续血管移植/吻合等用。 | Ⅱ
148 | 03-11-01 | 神经外科用钻头 | 一般由不锈钢材料制成。可重复使用。用于神经外科开颅手术。无菌提供。 | Ⅱ
149 | 03-11-01 | 电动骨钻 | 通常由钻头和驱动装置组成。钻头一般采用不锈钢材料制成。用于组织钻孔，建立通路。 | Ⅱ
150 | 03-12-01 | 脑室液抽吸管 | 脑室液抽吸管用金属制成，由管身与鲁尔接头组成，管身为细长管状，头部可有端孔或侧孔，可按材质、设计、技术参数等不同分为若干型号及规格；通常为可重复器械；供脑内手术时抽吸脑室液作检验或引流脑室液用。 | Ⅱ
151 | 03-12-01 | 心内吸引头 | 心内吸引头用金属制成，由头端、管身和接头组成，头端可有多种角度与形状设计，以适应不同解剖位置的吸引，可按材质、设计与技术参数等不同分为若干型号与规格；可反复使用；与吸引器配套，供心脏手术时吸引心包/心内液体用。 | Ⅱ
152 | 03-12-01 | 血管冲洗/吸引管 | 血管冲洗/吸引管用不锈钢等金属材料制成，由头端、管身和接头组成，可按材质、设计、技术参数等不同分为若干型号与规格；可反复使用；冲洗装置配套，供外科手术中灌注/吸引、冲洗在体或离体血管用。 | Ⅱ
153 | 03-12-02 | 导管皮下隧道工具 | 本产品由通条、手柄、闭孔器和外套管等部件组成。产品经灭菌，一次性使用。用于创建皮下隧道，帮助皮下放置植入泵的连接建立通道的导管。 | Ⅱ
154 | 03-13-00 | 球囊导引导管 | 可用于术中血管通路的建立，输送介入器材，还可同时临时封堵血管。可由导管、球囊、座等组成。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
155 | 03-13-01 | 造影导管 | 用于注射或输入对照介质和/或液体。通常由管体、导管尖端、不透射线标记和座等结构组成，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
156 | 03-13-02 | 导引导管 | 用于术中血管通路的建立，或用于注射或输入对照介质和/或液体和/或栓塞材料。通常由管体、导管尖端、不透射线标记和座等结构组成，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
157 | 03-13-04 | 导管消毒连接器(输液接头消毒帽、酒精保护帽) | 通常由接头、70%异丙醇或乙醇水溶液等组成。既用于物理屏障输液（注射）器具无针接头（包括外圆锥锁定接头及内圆锥锁定接头）；也用于对输液（注射）器具无针接头进行消毒。 | Ⅱ | 修订产品管理类别描述。
158 | 03-13-05 | 灌注导管 | 通常由导管管体、导管尖端、不透射线标记、导管座组成，导管远端可有侧孔。可带涂层，可带附件。用于将各种诊断和治疗溶剂递送至血管内。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
159 | 03-13-06 | 球囊扩张导管 | 用于动静脉血管狭窄或血管内支架（用于支撑血管管腔的可植入管状结构）、血管内假体、扩张治疗或人工动静脉透析瘘管的狭窄。可由球囊（特殊设计除外）、导管、导管座和射线可探测性标记组成，可带有附件，可涂有涂层。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。特殊设计如棘突球囊、切割球囊等除外。 | Ⅲ
160 | 03-13-06 | 锚定球囊扩张导管 | 锚定球囊扩张导管用于经皮冠状动脉成形术中，通过球囊的扩张，固定导引导管内的导引导丝，以实现其他介入装置的交换。主要由球囊、导管和管座组成。锚定球囊上有一个单腔的鲁尔接头，用于扩张和收缩球囊。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ | 新增
161 | 03-13-08 | 造影球囊 | 用于向血管内注入对照介心脏封堵器装载器质，进行血管内造影或测压。通常由球囊、不透射线标记、导管管体、座等结构组成，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
162 | 03-13-08 | 造影球囊导管 | 通常由球囊、不透射线标记、接头等结构组成。用于向血管内注入造影液，进行血管内造影。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
163 | 03-13-09 | 封堵球囊导管 | 通常由导管、球囊、接头组成。用于临时封堵外周血管或神经血管，亦可选择性地阻断或控制血流。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
164 | 03-13-10 | 血栓抽吸导管 | 通常由导管、接头、抽吸装置组成，导管有腔体用于血管内物质的吸出。用于抽吸血管内的血栓。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外；神经血管应用除外。 | Ⅲ
165 | 03-13-12 | 房间隔穿刺系统 | 房间隔穿刺系统通常包括房间隔穿刺针和房间隔穿刺鞘。其适用于行房间隔穿刺以建立两者通路。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。药械组合产品除外。 | Ⅲ |  | 新
166 | 03-13-12 | 穿刺针 | 由穿刺针管和管座组成，建立有助于血管内器械的经皮进入通路。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
167 | 03-13-13 | 导引套管 | 用于将导管或导丝插入，与穿刺针配合使用。通常由管体、接头组成。可带有附件，可涂有涂层。与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
168 | 03-13-14 | 导管鞘 | 与扩张器配合使用，用于将导丝、导管等医疗器械插入血管。通常由鞘管、管座和侧支等组件组成，可带有止血阀等附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
169 | 03-13-14 | 止血阀 | 用于经皮血管成形术中的辅助器械。可由止血阀外壳、阀、锁紧接头、侧管、三通旋塞等组件组成。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。 | Ⅱ | 修订产品管理类别描述。
170 | 03-13-15 | 扩张器 | 由扩张器管和管座组成。采用Seldinger术，建立有助于血管内器械的经皮进入通路，不具有血管内定位或建立血管内通路作用。Seldinger术多指数字减影血管造影术中的血管穿刺术。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
171 | 02-11-03 | 扩张器 | 由扩张器管和管座组成。采用经皮穿刺术，建立气道通路。产品基本原理、适用范围、性能和组成结构与已经上市产品相同。 | Ⅱ | 修订产品管理类别和分类编码描述。
豁免情况不包括：使用新型材料；包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等产品。
172 | 03-13-15 | 扩张鞘 | 扩张鞘由内鞘管和外鞘管组成。适用于经皮扩张心脏电极导线、留置导管和异物周围的组织。 | Ⅱ | 修订产品管理类别描述。
产品基本原理、适用范围、性能和组成结构与已经上市产品相同。
豁免情况不包括：使用新型材料；包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等产品。
173 | 03-13-16 | 导丝 | 用于引导其他器械插入血管，建立有助于血管内器械的经皮进入通路，或进行血管内定位，或建立血管内通路。可由芯丝、绕丝、安全丝、护套等组件组成，可带有附件，可涂有涂层，可包含延长导丝。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
174 | 03-13-17 | 球囊加压装置 | 通常由充气装置、压力表、连接管和接头组成。 用于心血管介入手术中，与球囊扩张导管连接使用，通过压力表显示的压力值，对球囊作精确充盈及收缩，从而达到扩张或收缩球囊的目的。 | Ⅱ
175 | 03-13-18 | 连接阀 | 用于连接管路，建立通道辅助器械进入人体，同时可减少血液流出。通常有两个或多个接口，可连接进入血管的导管或其他器械，可带有附件。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅱ | 修订产品管理类别描述。
176 | 03-13-19 | 腔静脉滤器回收装置 | 用于将可回收腔静脉滤器移出人体或转换可转换滤器。通常由与腔静脉滤器相匹配的回收圈套和回收鞘管等组成，可在显影设备下抓捕并移出可转换腔静脉滤器和可回收腔静脉滤器。可带有附件，可涂有涂层。与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
177 | 03-13-20 | 心脏封堵器装载器 | 用于收纳心脏封堵器，当与鞘管连接后，可将其中心脏封堵器推送至鞘管中。通常由导管、接头组成，导管较短，其长度可覆盖配用的心脏封堵器。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外。 | Ⅲ | 新增
178 | 03-13-21 | 心脏封堵器输送装置 | 通过股静脉/股动脉输送心脏封堵器。可由装载器、推送器、导引器、圈套器等组成，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ | 修订产品名称和产品描述。
179 | 03-13-22 | 血管内回收装置 | 用于经导管捕捉和取出血管内介入器械失效后脱落的异物，如支架、断裂的导管和导丝等。通常由抓取装置、杆、控制装置和鞘管等结构组成，具备异物抓取功能，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品、神经血管应用除外。 | Ⅲ
180 | 03-13-23 | 远端保护装置 | 通常由鞘管、导丝和滤网类结构等组成。鞘管一般采用高分子材料制成，导丝一般采用不锈钢材料制成，滤网结构一般采用镍钛合金制成，滤网上可覆膜。用于神经血管介入手术过程中容纳和移除栓塞物质。产品组成材料成熟，作用机理明确，与已获准境内注册产品具有等同性。新型结构设计、作用机理除外；药械组合产品除外。 | Ⅲ
181 | 03-13-24 | 环柄注射器 | 一般由外套、6%外圆锥接头、芯杆、芯杆手柄、卡圈、活塞、密封“O”型圈组成，预期用于微创伤介入治疗或诊断手术中手动推注造影液及药液。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
182 | 03-13-25 | 高压造影注射延长管 | 产品由导管,旋转鲁尔接头和O型圈/固定鲁尔接头,座组成。一次性使用产品，无菌供应。该产品在血管造影期间作为推注造影剂的管道。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅱ | 修订产品管理类别描述。
183 | 03-13-26 | 微导管 | 用于注射或输入对照介质和/或液体和/或栓塞材料，和/或适当的器械（如支架、弹簧圈）等。通常由管体和座等组件组成，可带有附件，可涂有涂层。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
184 | 03-14-02 | 一次性使用心表稳定板 | 一次性使用心表稳定板由稳定板（中式、左式、右式，塑料）、手臂（不锈钢）、锁定支架、锁定支架旋钮（ABS）、底座和手臂调节旋钮组成，可按材质、设计、技术参数等不同分为若干型号及规格；以无菌形式提供；供心脏不停跳搭桥手术时稳定局部心表，为冠脉吻合术提供条件。 | Ⅱ
185 | 03-14-02 | 一次性使用心脏固定器 | 用于心脏搭桥手术时，在心脏的目标血管处建立一个固定的手术区域。分为Stable-V1（吸引固定型）和Stable-M1（机械固定型）两种，通过固定头的压力和负压吸引力或仅通过固定头的压力固定手术区域。一般由调节手柄、固定卡座、万向支撑臂、吸盘、负压吸引管组成，也可由压脚叶（或吸脚固定座、吸脚）、负压连接管、三通阀、定向夹头组件、调节臂、固定座、调节手柄、旋柄、内穿拉紧钢丝组件组成。以无菌形式提供；产品不进入血液循环系统。也可称为心脏稳定器。 | Ⅱ
186 | 03-14-07 | 脑科平面定位尺 | 由冠状面、矢状面和中间板三部分组成。一次性使用无菌产品。用于脑外科手术中辅助测量、定位平面。 | Ⅱ
187 | 04 | 无源灭菌骨科手术器械 | 管理类别为Ⅰ类的无源骨科手术器械，以灭菌包装形式提供后管理类别升为Ⅱ类的，可以免于进行临床评价。 | Ⅱ
188 | 04 | 半月板缝合手术器械 | 该产品由半月板缝合针、导向管、半月板组锉组成，主体应采用符合ASTM F899、ASTM A269、ASTM A564标准规定的不锈钢材料制成。 | Ⅱ
189 | 04 | 无源内窥镜下骨科手术工具 | 无源内窥镜下暂时使用的骨科手术工具。一般由金属材料或高分子材料制成。可按材质、设计、技术参数、工作部位、预期用途等不同分为若干型号及规格。配合内窥镜使用，用于骨科手术、检查。 | Ⅱ
190 | 04-01-01 | 截骨用刀 | 通常由刀片和手柄组成，远端有坚硬、锋利、单刃配置的切割刀片，手柄位于其近端。一般采用不锈钢材料制成。无菌提供。用于切除、截断骨。 | Ⅱ |  | 新
191 | 04-01-02 | 关节镜用手术刀 | 通常由头部、杆部和手柄组成，头部为一刃口片，通过手柄操作传递、控制头部工作。头部一般采用不锈钢材料制成。手术中在内窥镜下操作，用于骨科微创手术中，对病变组织进行切除、剥离。 | Ⅱ
192 | 04-01-03 | 脊柱手术用铰刀 | 该类产品通常由不锈钢材料制成。用于脊柱后路手术中扩孔、铰孔或去除终板的软骨层。无菌提供。 | Ⅱ
193 | 04-01-03 | 椎间盘切除刀 | 椎间盘切除刀采用符合YY/T 0294.1或ASTM F899标准规定的不锈钢材料制成。用于切开椎间盘最外层的纤维环。无菌提供。 | Ⅱ
194 | 04-02-01 | 关节镜用手术剪 | 通常由头部、杆部或软性导管和手柄组成，头部为一对带刃口的叶片，通过手柄操作传递、控制头部工作。一般头部采用不锈钢材料制成。手术中在内窥镜下操作，用于剪切组织。 | Ⅱ
195 | 04-03-01 | 关节镜用手术钳 | 通常由头部、杆部或软性导管和手柄组成，头部为一对带钳喙的叶片，通过手柄操作传递、控制头部工作。头部一般采用不锈钢材料制成。手术中在内窥镜下操作，用于钳夹组织或器械。 | Ⅱ
196 | 04-03-02 | 脊柱侧弯矫正钳 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1127 咬骨钳。供脊柱手术时咬切骨组织和整修骨骼用。无菌提供。 | Ⅱ
197 | 04-03-03 | 椎骨咬骨钳 | 椎骨咬骨钳用适用的金属材料制成（如钳头采用40Cr13、32Cr13Mo或40Cr13Mo，钳柄、锶轴螺钉、撑簧及撑簧滑轮采用20Cr13、30Cr13或32Cr13Mo材料制成，撑簧螺钉采用12Cr18Ni9或06Cr19Ni10制成），可由钳头、钳柄、锶轴螺钉、撑簧螺钉、撑簧、撑簧滑轮等组成，可分为单关节、双关节、侧角等若干类，每类按钳头形式不同分直头、弯头和侧角头等若干种，也可按钳头的形状可分为圆形和尖形等若干种，按钳柄型式不同可分为有肩和无肩两种，按刃口的开档可分为正常开裆与打开档两种；每种按口阔及全长等尺寸不同分为若干规格；供骨科手术时咬切腐死骨和整修骨骼用（按设计不同可分别咬切颅骨、椎板或鼻中隔等）。产品性能指标采用下列参考标准中的适用部分，如：YY/T1127－2006咬骨钳。无菌提供。 | Ⅱ
198 | 04-03-03 | 颈椎咬骨钳 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1127 咬骨钳。供颈椎手术时咬切骨组织和整修骨骼用。无菌提供。 | Ⅱ
199 | 04-03-03 | 颈椎双关节咬骨钳 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1127 咬骨钳。供颈椎手术时咬切骨组织和整修骨骼用。无菌提供。 | Ⅱ
200 | 04-03-03 | 弯头平口棘突骨钳 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1127 咬骨钳。供骨科手术时咬切骨组织和整修骨骼用。 | Ⅱ
201 | 04-03-03 | 枪形咬骨钳 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T1127 咬骨钳。供骨科手术时咬切骨组织和整修骨骼用。 | Ⅱ
202 | 04-03-04 | 髓核钳 | 髓核钳由钳头、盖板、钳身、左钳柄、支撑弹簧组成，按型式不同分直头和弯头两种，每种按尺寸不同分为若干规格；供骨科手术中咬切髓核等软组织用。无菌提供。 | Ⅱ
203 | 04-03-05 | 脊柱后路撑开器 | 该类产品的工作原理是器械头端的宽度大于高度，使用时正向插入椎间隙，然后旋转器械90度，使宽度方向由水平变为立起，从而撑开椎间隙。通常由不锈钢材料制成。无菌提供。可分为直型和弯型，有多种尺寸。可重复使用。用于脊柱后路手术中，撑开椎间隙。 | Ⅱ
204 | 04-03-05 | 椎板撑开器 | 该类产品通常由不锈钢材料制成，在脊柱融合等手术中通过撑开椎板以撑开椎间隙。无菌提供。 | Ⅱ
205 | 04-04-01 | 脊柱手术用神经档钩 | 该类产品通常由不锈钢材料制成，在脊柱手术中用于挡住神经根，便于融合器植入。无菌提供。 | Ⅱ
206 | 04-04-01 | 脊柱手术用神经拉钩 | 该类产品通常由不锈钢材料制成，在脊柱手术中用于牵开神经根或软、硬膜囊。无菌提供。 | Ⅱ
207 | 04-05-01 | 关节镜用手术探针 | 通常由头部、杆部或软性导管和手柄组成，头部为针形，通过手柄操作传递、控制头部工作。头部一般采用不锈钢材料制成。手术中在内窥镜下操作，用于探查组织、取样。 | Ⅱ
208 | 04-05-02 | 无菌骨牵引针 | 通常由头部、针体和尾部组成，可分为螺纹型和光杆型两种型式。一般采用不锈钢材料或钛合金材料制成。无菌提供。 | Ⅱ
209 | 04-06-01 | 关节镜用手术刮匙 | 通常由头部、杆部或软性导管和手柄组成，头部为匙形，通过手柄操作传递、控制头部工作。头部一般采用不锈钢材料制成。手术中在内窥镜下操作，用于骨科微创手术中，对病变组织进行刮削。 | Ⅱ
210 | 04-06-02 | 脊柱手术用刮刀 | 该类产品通常由不锈钢材料制成。用于椎间融合等脊柱手术中将髓核、纤维环等刮除。无菌提供。 | Ⅱ
211 | 04-06-02 | 脊柱手术用刮匙 | 该类产品通常由不锈钢材料制成。用于脊柱后路手术中刮除已经捣碎的椎间盘、去除上下终板等。无菌提供。 | Ⅱ
212 | 04-06-02 | 椎体成形用刮匙器 | 该类产品通常由头部和柄部组成。在近端有手柄，远端为具锋利边缘的匙形凹尖，也可以是双端的。一般由不锈钢材料制成，可重复使用。用于刮除病灶、窦道内的瘢痕、肉芽组织，以及骨腔和潜在腔隙的死骨或病理组织等。无菌提供。 | Ⅱ
213 | 04-07-01 | 骨科微创介入术用骨锥 | 通常由导针、带手柄的空心钻和工作通道组成。金属部分一般采用不锈钢材料制成，手柄一般采用工程塑料制成。无菌提供。用于骨科微创介入手术中（经皮椎体成形、椎体后凸成形术等）钻孔，建立工作通道。 | Ⅱ
214 | 04-07-02 | 脊柱手术用手锥 | 该类产品头部通常由不锈钢材料制成。供脊柱手术时在骨骼上开孔。无菌提供。 | Ⅱ
215 | 04-08-01 | 脊柱后路手术用骨钻 | 该类产品通常由手柄和金属部分组成。手柄通常由塑料制成，金属部分通常由不锈钢制成。用于脊柱后路手术中钻出骨道。 | Ⅱ
216 | 04-08-01 | 环钻 | 通常为中空结构。一般采用符合YY/T 0294.1或ASTM F899标准规定的不锈钢材料制成。用于脊柱手术中切除骨或取骨用。 | Ⅱ
217 | 04-08-02 | 无菌磨头 | 通常由头部、刃部和刀柄组成。一般采用金刚石或不锈钢、碳化钨、金刚砂材料制成。无菌提供。用于骨组织开孔或打磨骨组织。 | Ⅱ
218 | 04-08-02 | 关节镜用手术钻头 | 通常由头部和柄部组成，头部有切割刃口。头部一般采用不锈钢、碳化钨、金刚砂材料制成，柄部一般采用不锈钢、钛或合成材料制成。手术中在内窥镜下操作，用于关节微创手术中，用于钻孔、攻螺纹。 | Ⅱ
219 | 04-09-01 | 椎间盘手术用环锯 | 该类产品为环形手术切割器械，有锯齿刃口，一般采用不锈钢材料制成。用于椎间盘手术时切除和修正腐骨、钻孔、减压或断钉取出等。 | Ⅱ
220 | 04-09-01 | 无菌骨锯 | 通常为扁平或线型的手术切割器械，有锯齿刃口。可带有附属手柄。一般采用不锈钢材料制成。无菌提供。不接触中枢神经系统。用于截锯骨骼。 | Ⅱ
221 | 04-10-01 | 椎骨骨凿 | 椎骨骨凿通常用不锈钢制成，手柄可有其他适用材料，通常为单刃口的手持手动一件式外科器械，远端有成型的锋利刃口，刃口形状可各异；可按材质、设计、技术参数等不同分为若干型号及规格；可使用手术槌或手术锤手动将力从器械的近端传至远端，供手术时凿切、修整、剥离颈椎、腰椎等椎骨并使椎骨成型。无菌提供。 | Ⅱ
222 | 04-10-01 | 颈椎骨凿 | 该类产品通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 91141骨科凿类通用技术条件。用于颈椎手术时骨组织的切削和修整。无菌提供。 | Ⅱ
223 | 04-10-01 | 椎板骨凿 | 该类产品通常由柄部和刀头组成，刀头是斜面锋利刃口。通常由不锈钢材料制成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 91141骨科凿类通用技术条件。用于脊柱手术时切削和修整椎板。无菌提供。 | Ⅱ
224 | 04-11-01 | 椎管锉刀 | 该类产品通常由不锈钢材料制成。供脊柱手术时铲除骨片及修正骨骼用。无菌提供。 | Ⅱ
225 | 04-11-01 | 脊柱手术用骨锉 | 该类产品通常由不锈钢材料制成，用于脊柱后路手术中终板、椎体边缘等骨面的处理。无菌提供。 | Ⅱ
226 | 04-11-01 | 关节镜用手术骨锉 | 通常由锉身和手柄组成。锉身一般采用不锈钢、金刚砂材料制成。手术中在内窥镜下操作，用于关节微创手术中，对组织进行磨削。 | Ⅱ
227 | 04-11-02 | 椎管铲刀 | 该类产品通常由不锈钢材料制成。供脊柱手术时铲除骨片及修正骨骼用。无菌提供。 | Ⅱ
228 | 04-12-01 | 骨科动力手术设备 | 通常由主机、软轴、电缆、手机和刀具（或脚踏开关）组成，或由手机、刀具、电池和控制系统等组成。用于对骨组织进行钻、切取、锯、磨、铣等。 | Ⅱ
229 | 04-13-02 | 骨科外固定支架 | 骨科外固定支架通常由针杆链接装置、杆、环、连接部件等组成，不包括骨针。一般采用铝合金、不锈钢、碳纤维、钛合金材料制成，可由多个支架、固定部件组合而成，可为单臂或环形，可有万向结构，可按材质、设计、技术参数、适用部位、预期用途等不同分为若干型号及规格；连接体内骨针使用，通过固定、加压或牵拉骨端，实现骨科畸形矫形、骨折复位等治疗目的。 | Ⅱ
230 | 04-14-02 | 一次性使用/无菌骨水泥填充器 | 一次性使用/无菌骨水泥填充器可由骨水泥胶枪（由金属制手柄、机构螺杆或类似结构组成的枪状）、可分离喷嘴套筒（套筒两端有螺纹等连接结构，可按装在骨水泥胶枪上，套筒可由金属或高分子材料制成）、骨水泥配制容器、骨水泥勺、搅拌棒、套筒座等组成。金属部件通常重复使用，非金属部件通常为一次性使用。无菌提供或一次性使用，供骨科手术中混匀骨水泥，并将骨水泥注入（输送到）骨髓腔需填充部位用。包括低粘骨水泥填充套件。 | Ⅱ
231 | 04-14-02 | 骨水泥填充套件 | 无菌骨水泥填充器、一次性使用骨水泥填充器可由骨水泥胶枪（由金属制手柄、机构螺杆或类似结构组成的枪状）、可分离喷嘴套筒（套筒两端有螺纹等连接结构，可按装在骨水泥胶枪上，套筒可由金属或高分子材料制成）、骨水泥配制容器、骨水泥勺、搅拌棒、套筒座等组成。金属部件通常重复使用，非金属部件通常为一次性使用。无菌提供或一次性使用，供骨科手术中混匀骨水泥，并将骨水泥注入（输送到）骨髓腔需填充部位用。包括低粘骨水泥填充套件。 | Ⅱ
232 | 04-14-02 | 骨水泥加压塞 | 该类产品通常由股骨端加压塞和髋臼加压塞组成，材料分别为聚氨酯和硅6810橡胶，用于髓腔或髋臼注入骨水泥后加压塑形。一次性使用，灭菌包装。 | Ⅱ
233 | 04-14-03 | 脊柱后路手术用植骨漏斗 | 该类产品通常由不锈钢材料制成，在脊柱后路融合术中用于将植骨导入植骨区域。无菌提供。 | Ⅱ
234 | 04-14-04 | 骨取样器 | 通常由套管、管芯和手柄组成。套管和管芯一般采用不锈钢材料制成。用于经皮椎体后凸成形术/经皮椎体成形术中骨组织取样。无菌提供。 | Ⅱ
235 | 04-14-05 | 脊柱手术用试模 | 该类产品通常由不锈钢材料制成，在脊柱后路融合等手术中用于评估椎间隙的高度，为椎间融合器的选择提供参考。无菌提供。 | Ⅱ
236 | 04-14-05 | 脊柱后路手术用测深器 | 该类产品通常由不锈钢材料制成，用于脊柱后路手术中切除髓核等椎间组织后腔体深度的测量。 | Ⅱ
237 | 04-15-01 | 扩髓器 | 该类产品分为两类，一类由柔性杆、钻头、导针和送针器组成；另一类由扩髓头、管型组装部件、过滤装置、密封头、锁定夹等组成。骨科手术时用于髓腔扩大。与有源器械配合使用。 | Ⅱ
238 | 04-16-01 | 膝关节用骨水泥定型模具 | 一般采用医用级硅胶制成，不包含植入体内的加固柄等组件。用于膝关节用骨水泥定型。无菌提供。 | Ⅱ
239 | 04-16-01 | 髋关节用骨水泥定型模具 | 由不锈钢制成的加固柄和医用级硅橡胶制成的骨水泥定型模具组成，一次性使用，无菌产品。用于塑造临时半髋关节置换假体，适用于因感染而需进行二期翻修的骨骼发育成熟的患者。 | Ⅲ
240 | 04-16-03 | 3D打印截骨导板 | 骨科手术配套工具，用于匹配患者解剖结构。采用高分子聚酰胺材料或金属材料，经3D打印制成。非无菌提供，经灭菌后一次性使用。预期用于特定病人的骨科手术中定位、导向和保护。 | Ⅱ |  | 新
241 | 04-17-04 | 椎体成形术辅助器械 | 金属材料器械的主体通常采用不锈钢材料制成；充盈装置(充压装置)采用聚碳酸酯和聚氨酯等高分子材料制成；骨水泥注入器可采用高分子材料或金属材料制成；可由穿刺套件、扩张套件、工作套管、骨水泥注入套件和充盈装置等组成，不含椎体扩张球囊导管。用于经皮椎体后凸成形术/经皮椎体成形术中，建立工作通道，填充骨水泥，恢复椎体解剖结构。 | Ⅱ
242 | 04-17-05 | 一次性使用纤维环缝合器 | 该类产品通常由缝合器壳体、缝合组件及传动组件组成。用于单纯椎间盘突出髓核摘除手术后的纤维环缝合。 | Ⅱ
243 | 04-17-06 | 椎体后缘处理器 | 该类产品通常由引导丝定位、扩张套管、高精度钻、工作套管等组成。用于处理后缘骨赘、硬性突出物，疏通神经根通道。 | Ⅱ
244 | 04-17-13 | 脊柱手术通道系统 | 脊柱手术配套手术工具，一般由可扩张通道、逐级撑开套筒、撑开通道及内芯和测深尺组成。采用医用高分子材料制成。无菌提供。用于脊柱手术通道的建立。 | Ⅱ
245 | 04-17-15 | 脊柱后路手术用植骨推骨器 | 该类产品通常由不锈钢材料制成，在脊柱后路融合术中用于将植入骨推入植骨区域并压实。 | Ⅱ
246 | 04-17-15 | 脊柱手术用持笼器 | 该类产品通常由不锈钢材料制成，在脊柱后路融合术中用于抓取融合器以完成融合器的植入或取出。 | Ⅱ
247 | 04-18-01 | 椎板剥离器 | 该类产品通常由不锈钢材料制成。用于脊柱手术时剥离椎板附着组织。 | Ⅱ
248 | 04-18-01 | 椎体前方剥离器 | 该类产品通常由不锈钢材料制成。用于脊柱手术时剥离或分离粘膜、骨膜等组织。 | Ⅱ
249 | 04-18-02 | 颅骨成形术材料形成模具 | 该类产品通常由材料形成用造型正负模具（主要由医用硅橡胶材料制成），头颅造型术的辅助工具组成。作为颅骨修补材料的成形模具使用。 | Ⅱ
250 | 04-18-02 | 婴儿颅骨矫形固定器 | 由外壳、填充材料/垫和固定装置组成。一般采用高分子材料制成。用于3-18个月之间的婴幼儿，通过给婴幼儿头部接触部位一定的压力，预防或治疗先天性或后天发生的非对称或斜头症及短头症的矫形器械。 | Ⅱ
251 | 05-04-01 | 射线束扫描测量系统-水模体系统 | 通常由水箱、控制单元、探测器、控制软件、电缆线等组成。用于测量射线束在水中的吸收剂量分布和在空气中的比释动能分布，测量结果用于对放射治疗计划系统的数据配置和修改，及放射治疗设备的质量控制。 | Ⅱ
252 | 05-04-01 | 射线束扫描测量系统-剂量测量矩阵 | 通常由探测器阵列、测量控制系统、控制软件等组成。 | Ⅱ
用于测量放射治疗设备所执行的放射治疗计划的剂量分布，用于与计划系统的数据进行比较，比较结果作为计划系统验证和修改的依据。
253 | 05-04-06 | 放射治疗激光定位系统 | 通常由多个固定或移动式激光灯等组成。 | Ⅱ
254 | 05-04-07 | 放射性粒籽植入防护器 | 通常由防护件、附件等组成，用于对操作者及其他周边人群的防护。 | Ⅱ
255 | 05-04-08 | 头颅肿瘤放射治疗定位装置 | 通常由带有头垫和前片的框架、牙垫、热塑性面膜、真空垫、患者控制单元、适配器、立体定向框架、重复定位检查工具等组成。用于立体定向放射治疗过程中，对头部进行固定、定位和重新定位。 | Ⅱ
256 | 06-01-02 | 泌尿X射线机 | 通常由X射线发生装置、图像显示系统、专用泌尿床组成。X射线透视摄影设备。用于泌尿科、妇科、胃肠道等X射线透视和摄影，获得影像供临床诊断用。产品不包括数字化体层摄影功能。 | Ⅱ |  | 新
257 | 06-01-03 | 乳腺X射线机 | 通常由机架、X射线发生装置、乳腺压迫器、影像接收装置组成。数字化产品还带有工作站和显示系统。一般采用钼或铑等材料制X射线管靶面。使用较低的管电压形成低能量的X射线进行摄影。用于对人体乳腺组织摄影，获得影像供临床诊断用。 | Ⅱ
258 | 06-01-04 | 牙科X射线机 | 通常由X射线发生装置及其支撑部件组成。配合口内影像接收器使用。用于对牙齿进行X射线摄影，获得影像供临床诊断用。 | Ⅱ
259 | 06-01-05 | 透视摄影X射线机 | 通常由X射线发生装置、图像显示系统、患者支撑装置组成，有的配有专用胃肠床、压迫器、点片装置、体层摄影装置等，是用于辅助胃肠诊断，兼有摄影和透视功能的X射线设备。使用时一般通过透视动态图像锁定感兴趣区，而后用较大剂量摄影得到该区域清晰的静态图像。用于常规、胃肠道透视摄影检查，可带有泌尿摄影功能，获得影像供临床诊断用。 | Ⅱ | 修订产品描述。
260 | 06-01-05 | 透视摄影X射线机 | 通常由X射线发生装置、图像显示系统、患者支撑装置组成，有的配有专用胃肠床、压迫器、点片装置等，是用于辅助胃肠检查，兼有摄影和透视功能的X射线设备。使用时一般通过透视动态图像锁定感兴趣区，而后用较大剂量摄影得到该区域清晰的静态图像。用于常规、胃肠道透视摄影检查，且带有外周血管部位的数字减影血管造影功能，获得影像供临床诊断用。产品不包括数字化体层摄影功能。 | Ⅲ | 修订产品描述和产品管理类别描述。
261 | 06-01-06 | 移动式C形臂X射线机 | 通常由移动式C形臂支架、X射线发生装置、影像增强器电视系统或数字平板探测器成像系统等组成。用于外科手术透视及摄影，获得影像供临床诊断用。 | Ⅱ
262 | 06-01-06 | 移动式C形臂X射线机 | 通常由移动式C形臂支架、X射线发生装置、影像增强器电视系统或数字平板探测器成像系统等组成。用于外科手术透视及摄影，且带有外周血管部位的数字减影血管造影功能，获得影像供临床诊断用。产品不包括数字化体层摄影功能。 | Ⅲ
263 | 06-01-07 | 摄影X射线机 | 通常由X射线发生装置和摄影X射线附属设备组成。数字化系统还带X射线探测器及其影像系统。利用从X射线管发射出的X射线穿过患者身体不同组织和器官时对射线衰减不同的原理，将穿过患者且携带足够信息的X射线投射到成像介质上所形成的影像，转化为可见的平面灰度影像的通用X射线设备。用于对患者的常规摄影，获得单幅影像供临床诊断用。 | Ⅱ
264 | 06-01-08 | 透视X射线机 | 通常由X射线发生装置、荧光屏或电视系统或动态探测器等影像接收装置组成，可能带有患者支撑装置等。利用人体不同组织和器官对射线衰减不同的原理，通过对X射线源的连续加载，在成像介质上转化为动态影像的通用X射线设备。用于对患者的常规透视，获得连续影像供临床诊断用。 | Ⅱ
265 | 06-01-09 | X射线骨密度仪 | 通常由X射线发生装置、探测器、信息分析和显示系统组成，还可能带有患者支撑装置。根据不同密度的骨骼和组织对X射线的吸收程度不同，将接收到的带有人体信息的数字信号输入计算机进行分析得出骨密度的结果。用于通过对人体的X射线衰减测量，评估患者骨骼及邻近组织的骨密度和矿物质含量，供临床诊断用。 | Ⅱ
266 | 06-01-10 | 车载X射线机 | 安装在可移动运输工具上的X射线机，有透视和/或摄影功能。用于机动条件下，在远离医院的现场开展X射线透视、摄影诊断检查。 | Ⅱ
267 | 06-01-11 | 携带式X射线机 | 通常由X射线源组件、影像接收装置等组成。在使用时或使用的间隔期间，可由一个人或几个人携着从一个地方移到另一个地方的X射线机。用于对四肢或其他衰减度较小的身体部位组织进行X射线成像，供临床诊断用。 | Ⅱ
268 | 06-03 | X射线发生、限束装置 | 用于X射线的产生、限束。 | Ⅱ
269 | 06-04-01 | X射线影像增强器 | 影像增强器是将携带患者信息的X射线信号转换为可见光图像的光电真空管。装配于诊断X射线机，用于将X射线图像转换成可见光图像。 | Ⅱ
270 | 06-04-01 | X射线影像增强器电视系统 | X射线影像增强器电视系统通常由X射线影像增强器、光学系统、摄像机信号处理系统及影像显示装置组成，将X射线图像转换成相应的可见光影像的系统。装配于诊断X射线机，用于将X射线图像转换成可见光图像。 | Ⅱ
271 | 06-04-02 | X射线探测器 | X射线探测器（包括平板探测器或光电耦合器（CCD）探测器等）采用特定的光电转换介质将穿过人体的X射线信号转化为数字信号。装配于或配合诊断X射线机，用于将X射线信号转化为数字信号。 | Ⅱ
272 | 06-04-02 | 数字化X射线成像系统 | 数字化X射线成像系统一般包括X射线探测器、图像传输，处理和显示系统。装配于或配合诊断X射线机，用于将X射线信号转化为数字信号。 | Ⅱ
273 | 06-04-03 | 影像板成像装置 | 数字化X射线成像系统一般包括X射线探测器、图像传输，处理和显示系统。装配于或配合诊断X射线机，用于将X射线信号转化为数字信号。 | Ⅱ
274 | 06-05 | X射线体层摄影装置 | X射线体层摄影装置能清楚地摄取与人体纵轴相平行的某一层或几层组织的影像，又使其他各体层的影像模糊不清的摄片装置；可按设计、技术参数、附加辅助功能等不同分为若干型号；与摄影平床配合，用于对人体进行体层摄影。 | Ⅱ
275 | 06-05-01 | X射线透视摄 | 通常由床体、点片装置、遥控操作装置和/或近台操作装置等组成。用于胃肠X射线检查，配合胃肠X射线机使用。 | Ⅱ
影床
276 | 06-05-02 | X射线导管床 | 通常为单臂支撑，床面板一般采用碳纤维等材料。在与 C臂、U臂等介入手术X射线设备组合进行多方向摄影时，视野不受妨碍。用于普通介入治疗和数字减影血管造影（DSA），配合血管造影X射线设备使用。 | Ⅱ
277 | 06-05-03 | X射线电动摄影平床 | 患者支撑装置。可电动平移、转动等。用于X射线摄影成像中对患者的支撑。 | Ⅱ
278 | 06-05-04 | X射线诊断设备附属悬吊、支撑装置 | 通常由基座、立柱、支撑臂或悬吊架及导轨等部件组成。可电动平移，转动等。用于X射线源组件、影像接收装置等部件的悬吊、支撑。 | Ⅱ
279 | 06-05-04 | 胸片架 | 胸片架是医用X射线诊断附属设备，由基座、立柱、支撑臂等部件构成，可电动平移，转动等，可按设计、技术参数、附加辅助功能等不同分为若干型号；供支撑X射监视器和胸片摄影点片装置，提供合适的人体摄片体位等用。 | Ⅱ
280 | 06-05-05 | X射线造影剂注射装置 | 通常由外套、活塞和活塞密封圈组成，附件包括连接管、吸药管等。与高压注射器配套使用。 | Ⅲ
281 | 06-05-05 | 一次性使用高压造影注射器及附件 | 由针筒、连接管、吸药管等组成。产品预期用途：为配套CT高压注射泵、MRI高压注射泵、DSA高压注射泵及超声检查和心血管造影时推注造影剂的器械，不接触人体。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品分类编码。
06-10-01
282 | 06-05-05 | 造影剂注射装置 | 通常由注射机头、控制部分等组成。通过向人体血管内注入造影剂获得清晰的血管影像。与高压注射器针筒及附件配套使用。用于X射线造影、CT成像时，造影剂、生理盐水等的注入。 | Ⅱ | 与原277项合并修订产品描述。
283 | 06-05-05 | 造影剂注射装置 | 通常由外套、活塞和活塞密封圈组成，附件包括连接管、吸药管等。与高压注射器配套使用。豁免情况不包括新材料、新作用机理、新结构、新功能的产品。 | Ⅲ
284 | 06-05-06 | 防散射滤线栅 | 通常由铅条、介质等组成。放置于影像接收面之前，以减少辐射到影像接收面上的散射辐射，从而改善X射线影像对比度的一种装置。配合X射线机使用，用于增加X射线影像的对比度。 | Ⅱ
285 | 06-05-07 | X射线摄影暗盒 | 承装X射线摄影胶片等的装置，带有滤线栅，按照不同应用分为不同尺寸。 | Ⅱ
286 | 06-06 | 医用射线防护喷剂 | 医用射线防护喷剂主要成分为超氧化物歧化酶及其稳定剂，以冻干粉形式存在于可加盖喷雾泵的棕色塑料瓶(A瓶)中，其他辅料(姜黄素、山梨醇、山梨酸钾)以水溶液形式存在于加盖的白色塑料瓶(B瓶)中，可按设计、技术参数、装量等不同分为若干型号与规格；供预防和减轻医用射线等理化因素产生的自由基对人体皮肤、粘膜组织造成的损伤。 | Ⅱ
287 | 06-06-01 | 医用射线防护用具 | 通常由薄厚均匀柔软的铅橡胶、铅橡塑、铅玻璃或其它含重金属元素的防护材料片组成，用于人体放射治疗时的防护。 | Ⅱ
288 | 06-07-01 | 超声脉冲回波成像设备 | 超声脉冲回波成像设备利用超声脉冲回波原理，完成人体器官组织成像的超声系统。通常由探头（线阵、凸阵、相控阵、机械扇扫）、超声波发射/接收、信号处理和图像显示等部分组成；可按机型、成像性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠对腹部器官、心脏、浅表组织器官等部位进行超声成像，不包括非常规的应用方式。仅包括基础的超声B模式成像，不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能（如复合成像、谐波成像）。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备等。 | Ⅱ
289 | 06-07-01 | 超声脉冲回波成像设备 | 超声脉冲回波成像设备利用超声脉冲回波原理，完成人体器官组织成像的超声系统。通常由探头（线阵、凸阵、相控阵、机械扇扫）、超声波发射/接收、信号处理和图像显示等部分组成；可按机型、成像性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠或术中对腹部器官、心脏、浅表组织器官等部位进行超声成像，不包括非常规的和接触循环血液的应用方式。仅包括基础的超声B模式成像，不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能（如复合成像、谐波成像）。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备等。 | Ⅲ
290 | 06-07-02 | 超声频谱多普勒诊断设备 | 超声频谱多普勒诊断设备利用超声多普勒原理提取血流的方向、速度等信息，经处理后以频谱－时间或血流－时间方式予以显示，又分为脉冲波多普勒和连续波多普勒。通常由探头（单元式、凸阵、线阵、相控阵、机械扇扫）、超声波发射/接收、信号处理和频谱显示等部分组成；可按机型、信号采集/处理性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠或术中对腹部器官、心脏、浅表组织器官等部位的血流进行检测，不包括非常规的应用方式。仅包括基础的脉冲波多普勒（PW）和连续波多普勒模式（CW），不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能。产品性能指标采用下列参考标准中的适用部分，如：YY 0767-2009 超声彩色血流成像系统等。 | Ⅲ
291 | 06-07-02 | 超声彩色血流成像设备 | 超声彩色血流成像设备在超声脉冲回波成像基础上采用多普勒和自相关技术对血流成像，并将彩色编码信息叠加在B模式灰阶图像上予以实时显示。通常由探头（凸阵、线阵、相控阵、机械扇扫）、超声波发射/接收电路、信号处理和图像显示等部分组成；可按机型、成像性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠或术中对腹部器官、心脏、浅表组织器官等部位的血流进行成像，不包括非常规的应用方式。仅包括基础的超声彩色血流成像模式，不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能（如超声血流向量成像）。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备、YY 0767-2009 超声彩色血流成像系统等。 | Ⅲ
292 | 06-07-02 | 超声频谱多普勒诊断设备 | 超声频谱多普勒诊断设备利用超声多普勒原理提取血流的方向、速度等信息，经处理后以频谱－时间或血流－时间方式予以显示，又分为脉冲波多普勒和连续波多普勒。通常由探头（单元式、凸阵、线阵、相控阵、机械扇扫）、超声波发射/接收、信号处理和频谱显示等部分组成；可按机型、信号采集/处理性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠对腹部器官、心脏、浅表组织器官等部位的血流进行检测，不包括非常规的应用方式。仅包括基础的脉冲波多普勒（PW）和连续波多普勒模式（CW），不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能。产品性能指标采用下列参考标准中的适用部分，如：YY 0767-2009 超声彩色血流成像系统等。 | Ⅱ
293 | 06-07-02 | 超声彩色血流成像设备 | 超声彩色血流成像设备在超声脉冲回波成像基础上采用多普勒和自相关技术对血流成像，并将彩色编码信息叠加在B模式灰阶图像上予以实时显示。通常由探头（凸阵、线阵、相控阵、机械扇扫）、超声波发射/接收电路、信号处理和图像显示等部分组成；可按机型、成像性质（模拟与数字）、探头配置、技术参数、附加辅助功能、预期用途等不同分为若干型号；经体表、经阴道、经直肠对腹部器官、心脏、浅表组织器官等部位的血流进行成像，不包括非常规的应用方式。仅包括基础的超声彩色血流成像模式，不包括在其基础上通过修改波束形成、图像前处理、图像后处理算法产生的成像模式/功能（如超声血流向量成像）。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备、YY 0767-2009 超声彩色血流成像系统等。 | Ⅱ
294 | 06-08-01 | 消毒型医用超声耦合剂 | 消毒型医用超声耦合剂可由消毒剂（如三氯羟基二苯醚）、丙二醇、丙三醇（甘油）、三乙醇胺、卡波姆（交联聚丙稀酸树脂）组成；无菌提供。可按配方、技术参数、适用部位、装量等不同分为若干型号及规格；供超声诊断或治疗操作中，充填或涂敷于皮肤-黏膜与探头（或治疗头）辐射面之间，用于透射声波的中介媒质，用于改善探头与患者皮肤/粘膜之间的超声耦合效果，并具有消毒皮肤、粘膜功能。 | Ⅱ
295 | 06-08-01 | 腔道用医用超声耦合剂 | 腔道用医用超声耦合剂由博克-DP（三氯羟基二苯醚）、丙二醇、丙三醇三乙醇胺、卡波姆（交联聚丙稀酸树脂）组成；可以以无菌形式提供；涂布与腔内超声探头头端，供提高腔内粘膜与探头的超声耦合用。 | Ⅱ
296 | 06-08-03 | 超声内窥镜专用水囊 | 超声内窥镜专用水囊用天然胶乳/硅胶和聚四氟乙烯等材料制成，由水囊和水囊外套管组成，水囊外套管由插入部、连接部构成；安装于超声内窥镜前端，注无菌生理盐水膨胀后作为超声传导介质；可按材质、容量、形状、使用部位等不同分为若干种；以无菌形式提供（接头可为非无菌）；供超声内窥镜检查时探头与被检查部位皮肤/粘膜间的耦合用。 | Ⅱ
297 | 06-08-04 | 超声机械扫描探头 | 超声机械扫描探头通常由超声换能器和/或超声换能器阵元组，声透镜、壳体、电缆线等组成，换能器阵元平面或凸阵排列，可以按结构、型式、技术参数和功能等不同分为若干型号，与超声设备配套，经完好的皮肤、腔道粘膜（如阴道、直肠，不包括经食道超声探头、上下消化道超声内窥镜的超声部分）实现包括生物信息采集、组织成像等预期功能。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备。 | Ⅱ
298 | 06-08-04 | 电子线阵探头 | 电子线阵探头通常由超声换能器和/或超声换能器阵元组，声透镜、壳体、电缆线等组成，换能器阵元平面排列，可以按结构、型式、技术参数和功能等不同分为若干型号，与超声设备配套，经完好的皮肤、腔道粘膜（如阴道、直肠，不包括经食道超声探头、上下消化道超声内窥镜的超声部分）实现包括生物信息采集、组织成像等预期功能。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备。 | Ⅱ
299 | 06-08-04 | 电子凸阵探头 | 电子凸阵探头通常由超声换能器和/或超声换能器阵元组，声透镜、壳体、电缆线等组成，换能器阵元凸阵排列，可以按结构、型式、技术参数和功能等不同分为若干型号，与超声设备配套，经完好的皮肤、腔道粘膜（如阴道、直肠，不包括经食道超声探头、上下消化道超声内窥镜的超声部分）实现包括生物信息采集、组织成像等预期功能。产品性能指标采用下列参考标准中的适用部分，如：GB 10152-2009 B型超声诊断设备。 | Ⅱ
300 | 06-08-05 | 超声穿刺架 | 产品配合腔内或体外超声探头使用，通常安装在超声探头上，可用于固定超声探头和穿刺针，并在超声的引导下将穿刺针引导到人体的目标位置。产品为无菌提供。穿刺架应为无源医疗器械，不应含有导航模块、自动穿刺等有源结构和/或功能。 | Ⅱ | 与原298合并修订产品名称及产品描述。
301 | 06-10-01 | 磁共振高压造影注射系统 | 通常由操作室组件和扫描室组件组成。扫描室组件由非铁磁材料制成，导线经过射频屏蔽。不包含配套使用的针筒。用于静脉注射磁共振造影剂和/或在人体血管内注射常用的冲刷液，以满足磁共振扫描诊断需要。 | Ⅱ
302 | 06-10-02 | 磁共振辅助刺激系统 | 通常由在控制室使用的计算机、操作员显示器、光纤介质转换器等部件和在磁共振扫描室内使用的患者界面显示器、电源、光缆、按钮响应单元、耳机等部件组成。配合磁共振成像系统使用。用于在大脑扫描时，通过播放文本、图片、幻灯片、声音、CD和DVD，向患者提供视觉、运动、听觉和语言任务刺激。 | Ⅱ
303 | 06-11-01 | 伽玛（γ）照相机 | 通常由机架、准直器、探测器、数据采集、图像处理工作站、系统软件、患者支撑装置等组成。用于获取单光子放射性核素在人体全身或部分器官组织中的分布情况，形成平面图像。 | Ⅱ
304 | 06-11-04 | 放射性核素扫描仪 | 通常由主机、计算机、彩色打印机、工作软件和隔离电源组成。用于人体甲状腺放射性核素扫描成像。 | Ⅱ
305 | 06-13-01 | 医用红外热像仪 | 通常由红外摄像机、处理系统、支架和显示屏组成。通过红外摄影标出人体热图像。 | Ⅱ
306 | 06-13-02 | 红外乳腺检查仪 | 红外乳腺检查仪由探头(卤素灯探头、激光探头)、摄像机、主机(包含显示器、打印机、计算机、软件)组成，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；供乳腺疾病的红外光检查用。产品性能指标采用下列参考标准中的适用部分，如： YY 0324-2008红外乳腺检查仪。 | Ⅱ
307 | 06-13-04 | 手术显微镜 | 手术显微镜由光学系统、机架、照明系统、电气装置四个基本部分组成，也可包括照相/摄像系统及图像管理系统，可按结构组成、技术参数、预期用途、附加辅助功能等不同分为若干型号；供（非眼科）外科显微手术放大、照明、图像管理等用，不含荧光模块及其他特殊光谱用途。参照标准 GB 11239.1-2005手术显微镜 第1部分。 | Ⅱ
308 | 06-13-05 | 微循环显微镜检查仪 | 微循环显微镜检查仪可由微循环显微镜、光源（LED光源/冷光源）、CCD（黑白/彩色）、显示器（监视器）等组成，可包括图像管理软件，可按结构组成、技术参数、预期用途、附加辅助功能等不同分为若干型号；供人体未梢微循环（如甲襞）检查用，以此评估人体微循环状况。产品性能指标采用下列参考标准中的适用部分，如：YY 0067-2007微循环显微镜。 | Ⅱ
309 | 06-14-01 | 膀胱内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于尿道、膀胱的观察成像。 | Ⅱ
310 | 06-14-01 | 直肠内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于肛门、直肠的观察成像。 | Ⅱ
311 | 06-14-01 | 纤维鼻内窥镜 | 通常由光缆连接器、含导丝的柔性管以及远端和近端光学透镜组成，可按设计、技术参数等不同分为若干种；柔性管以无菌形式提供；与适配光源配套，用于鼻和鼻窦结构的观察成像。 | Ⅱ
312 | 06-14-01 | 一次性使用光纤喉镜 | 一次性使用光纤喉镜由塑料制手柄和内含防雾CCD摄像头及光纤导线的窥片（塑料制）组成，手柄另一端可接显示器。手柄与窥片通常为可分离式，可按适用人群、材质、设计、技术参数、窥片尺寸等不同分为若干型号和规格；产品以无菌形式提供；供临床挑起患者会厌部暴露声门，指引医护人员准确进行气道插管供麻醉或急救用，也可用于口腔内诊察、治疗用。用于观察和检查患者的上呼吸道以及协助呼吸插管的插入。 | Ⅱ
313 | 06-14-01 | 鼻窦内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于在鼻窦观察成像。 | Ⅱ
314 | 06-14-01 | 喉内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于喉部观察成像。 | Ⅱ
315 | 06-14-01 | 纤维鼻咽喉内窥镜 | 软性纤维内窥镜，一般由光学成像系统和照明系统组成，光学成像系统由物镜、转像系统、目镜三部分组成。照明系统由混编排列的多束导光纤维构成，成像系统由规则排列的多束导光纤维构成。物镜将物体直接聚焦成像于光纤面上，光纤面阵上的每一像素（每一根导光纤维）分别接收对应位置像的光能，并将该光能传输至传像光纤的另一端发出，所有像素在像方端的全部光能重组形成物镜的聚焦像，再由目镜或摄像系统放大用于观察。用于在鼻咽喉的检查和手术中观察成像。 | Ⅱ
316 | 06-14-01 | 纤维上消化道内窥镜 | 软性纤维内窥镜，一般由光学成像系统和照明系统组成，光学成像系统由物镜、转像系统、目镜三部分组成。照明系统由混编排列的多束导光纤维构成，成像系统由规则排列的多束导光纤维构成。物镜将物体直接聚焦成像于光纤面上，光纤面阵上的每一像素（每一根导光纤维）分别接收对应位置像的光能，并将该光能传输至传像光纤的另一端发出，所有像素在像方端的全部光能重组形成物镜的聚焦像，再由目镜或摄像系统放大用于观察。用于对上消化道（不包括十二指肠）进行观察成像。 | Ⅱ
317 | 06-14-01 | 纤维下消化道内窥镜 | 软性纤维内窥镜，一般由光学成像系统和照明系统组成，光学成像系统由物镜、转像系统、目镜三部分组成。照明系统由混编排列的多束导光纤维构成，成像系统由规则排列的多束导光纤维构成。物镜将物体直接聚焦成像于光纤面上，光纤面阵上的每一像素（每一根导光纤维）分别接收对应位置像的光能，并将该光能传输至传像光纤的另一端发出，所有像素在像方端的全部光能重组形成物镜的聚焦像，再由目镜或摄像系统放大用于观察。用于对下消化道进行观察成像。 | Ⅱ
318 | 06-14-01 | 纤维支气管内窥镜 | 软性纤维内窥镜，一般由光学成像系统和照明系统组成，光学成像系统由物镜、转像系统、目镜三部分组成。照明系统由混编排列的多束导光纤维构成，成像系统由规则排列的多束导光纤维构成。物镜将物体直接聚焦成像于光纤面上，光纤面阵上的每一像素（每一根导光纤维）分别接收对应位置像的光能，并将该光能传输至传像光纤的另一端发出，所有像素在像方端的全部光能重组形成物镜的聚焦像，再由目镜或摄像系统放大用于观察。用于对气管和支气管进行观察成像。 | Ⅱ
319 | 06-14-01 | 腹腔内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜系统、转像系统、目镜系统三大系统组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于腹腔的观察成像。 | Ⅱ
320 | 06-14-01 | 关节内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于关节部位的观察成像。 | Ⅱ
321 | 06-14-01 | 硬性光学胸腔内窥镜 | 硬性光学胸腔内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于胸腔部位的观察成像。 | Ⅱ
322 | 06-14-01 | 输尿管内窥镜 | 一般为半硬性光学内窥镜。通常由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统由混编排列的多束导光纤维构成，通过外部光源传递光线，为内窥镜影像提供照明。成像系统由规则排列的多束导光纤维构成，被观察物经物镜通过导像纤维传输到目镜，再由目镜或摄像系统放大用于观察。外管为医用不锈钢。可含有工作通道，用于兼容手术器械和吸引灌流。该产品插入尿道，通过膀胱穿入输尿管，用于输尿管的观察成像。 | Ⅱ
323 | 06-14-02 | 前列腺电切内窥镜 | 一般由硬性光学内窥镜、鞘套、操作器、手术电极组成。在内窥镜观察下，利用高频电流对前列腺组织进行切除。 | Ⅲ
324 | 06-14-03 | 电子腹腔内窥镜 | 产品由物镜端、镜体、光导接头和视频接头等构成。头端部的图像传感器（通常为CCD或CMOS）将接收到的光学信号转换为电信号，通过图像处理装置处理后在显示器上观察。产品用于腹腔手术中的观察成像。产品不含内置光源。 | Ⅱ |  | 新
325 | 06-14-03 | 电子上消化道内窥镜 | 电子上消化道内窥镜为软性电子内窥镜，一般由头端部、弯曲部、插入部及电气和光源连接的部分组成。头端部的光电转换器件将接收到的光学信号转换为电信号，通过摄像系统在显示器上观察。通过视频监视器提供影像供上消化道（不包含十二指肠）的观察、诊断、摄影用。 | Ⅱ
326 | 06-14-03 | 电子鼻咽喉内窥镜 | 软性电子内窥镜，一般由头端部、弯曲部、插入部、操作部及电气和光源连接的部分组成，可含有工作通道。头端部的光电转换器件将接收到的光学信号转换为电信号，通过摄像系统在显示器上观察。通过视频监视器提供影像，供鼻咽喉等上呼吸道的观察、诊断和治疗提供图像。 | Ⅱ
327 | 06-14-03 | 电子下消化道内窥镜 | 软性电子内窥镜，一般由头端部、弯曲部、插入部、操作部及电气和光源连接部组成，一般含有工作通道。头端部的光电转换器件将接收到的光学信号转换为电信号，通过图像处理系统在显示器上观察。通过视频监视器提供影像供下消化道(不含小肠)的观察、诊断和摄影用。 | Ⅱ
328 | 06-14-03 | 电子膀胱肾盂镜 | 软性电子内窥镜，一般由头端部、弯曲部、插入部、操作部及电气和光源连接部组成，可含有工作通道。头端部的光电转换器件将接收到的光学信号转换为电信号，通过图像处理系统在显示器上观察。通过视频监视器提供影像供尿道、膀胱、肾盂的观察、诊断、摄影和治疗用。 | Ⅱ
329 | 06-15-01 | 内窥镜用冷光源 | 通常由冷光源主机和电源线组成，可含导光束，灯泡类型有LED灯、氙灯和卤素灯，冷光源主机通过导光束与内窥镜连接，为内窥镜的检查或手术提供照明。不含特殊光谱。产品性能指标采用下列参考标准中的适用部分，如：YY 1081-2011医用内窥镜 内窥镜功能供给装置 冷光源。 | Ⅱ
330 | 06-15-01 | 频闪光源 | 频闪光源一般由主机、光源、麦克风、脚踏开关组成，由网电源或电池供电，通常配合喉镜和监视器共同使用。可按设计、技术参数、辅助功能不同分为若干型号。为喉部检查提供频闪光和普通光。 | Ⅱ
331 | 06-15-02 | 内窥镜摄像系统 | 由主机、摄像头、适配器和电缆线组成，与光学内窥镜及监视器配合使用，将内窥镜采集的光学信号转化为电子信号，并传输至监视器进行成像。 | Ⅱ
332 | 06-15-03 | 电子内窥镜图像处理器 | 该产品一般由主机和附件组成，附件包括电源线、接口电缆、数据键盘等，与电子内窥镜及显示器配合使用，用于将电子内窥镜采集到的图像进行处理后输送至显示器。 | Ⅱ
333 | 06-15-04 | 气腹机 | 由主机、气体过滤器和连接管路组成，用于腹腔镜微创手术时向腹腔内注入二氧化碳气体，建立并维持气腹，提供手术操作和视野的空间。 | Ⅱ
334 | 06-15-05 | 内窥镜用冲洗吸引器 | 由主机、连接线、脚踏开关、冲洗管路和吸引管路组成；该产品用于内窥镜手术中的冲洗和/或吸引。 | Ⅱ
335 | 06-15-06 | 子宫加压膨宫器 | 子宫加压膨宫器由加压器主机、管路及充气针组成，需有超压报警并自动泄气功能；可按设计、技术参数、辅助功能、预期用途等不同分为若干型号；供妇科内窥镜检查和手术过程中，对子宫进行加压膨胀以形成可视空间用。 | Ⅱ
336 | 06-16-05 | 内窥镜先端帽 | 一次性使用内窥镜用先端帽用硅橡胶等制成，可按材质、技术参数、适用内窥镜等不同分为若干种；无菌状态提供；安装于内镜先端部，以保持适当的内镜视野。 | Ⅱ
337 | 06-16-06 | 内窥镜用活检袋 | 产品可由输送装置、张开装置、结扎绳、纳物袋等部件构成。为一次性使用，无菌产品。用于内窥镜手术中将组织和/或活检样本从手术区域取出。 | Ⅱ | 与原336合并修订产品名称及产品描述。
338 | 06-16-07 | 一次性使用内窥镜套管 | 通常由开口端和套体组成。一般采用聚乙烯等高分子材料制成。无菌提供。配合内窥镜使用，用于手术或检查时使内窥镜与患者隔离，预防交叉感染。 | Ⅱ
339 | 06-16-07 | 内窥镜咬口 | 手术或检查时患者开口的辅助器械，通常采用聚乙烯等高分子材料制成。一次性使用，无菌提供。用于经口腔手术或检查时维持患者的开口状态，防止非预期咬合保护器械损坏。 | Ⅱ
340 | 06-18-01 | 图像显示处理工作站 | 通常由专用诊断显示装置、存储处理系统、软件等组成。配合医学影像设备，用于显示、处理、传输和存储数字诊断图像。 | Ⅱ
341 | 06-18-02 | 医用胶片数字化扫描仪 | 通常由胶片上板、胶片下板、胶片传动机构、光电耦合器（CCD）、数字信号处理器等组成。用于医用胶片的扫描，输出BMP、DICOM等格式的数字化图像。 | Ⅱ
342 | 07 | 腹内压监测套件 | 包括腹腔压力监测和尿液收集两部分装置。腹腔压力监测装置由连接器、连接器帽、管路、管路调节器、导夹、和过滤装置组成。尿液精准收集装置由引入管、测量室、收集袋、排出阀及附件（悬挂带/挂绳/挂钩）组成。使用时连接到导尿管的排泄端口，用于测量膀胱内静水压力，从而监测腹腔内压力，并精准计量、收集尿液。 | Ⅱ
343 | 07-01-01 | 压舌板 | 通常由木质或其他材料制成。无菌提供。用于检查时压低舌部。 | Ⅱ
344 | 07-01-02 | 电子听诊器 | 通常由拾音器、信号处理模块和耳机组成。对收集的声音进行（频率）非线性放大。用于收集和放大从心脏、肺部、动脉、静脉和其他内脏器官处发出的声音。 | Ⅱ
345 | 07-02-01 | 气体测定设备 | 临床用于连续测量呼吸气体中的氧气、二氧化碳等气体的浓度、分压、流速或容量等参数。 | Ⅱ | 修订产品名称及产品描述。
346 | 07-02-03 | 肺功能测试设备 | 通常由主机、流速传感器和鼻夹组成。用于呼吸内科、胸科、职业病防止机构、医院体检等，测量肺活量、最大通气量及用药前后激发试验。 | Ⅱ
347 | 07-02-04 | 呼气流量计 | 通常由壳体、滑标、簧片和咬嘴组成。用于测量哮喘患者、慢性阻塞肺病患者最大呼气流量，测量患者呼出气受限制的程度，可监测哮喘病情。 | Ⅱ
348 | 07-02-05 | 呼吸压力计 | 通常由主机、接嘴、呼气压力测试阀、吸气压力测试阀、细菌过滤器等组成。用于测量口腔最大吸气压力、口腔最大呼气压力和鼻腔吸气压力。 | Ⅱ
349 | 07-02-07 | 呼气分析仪 | 通常由主机、一氧化氮或一氧化碳等气体传感器、呼吸过滤装置或气体收集装置等附件组成。用于检测呼出气体中的一氧化氮或一氧化碳等的浓度。 | Ⅱ |  | 新
350 | 07-03-01 | 运动负荷试验测试系统 | 通常由主机、附件、运动单元组成。主机部分通常包括信号输入部分，放大回路、控制电路、显示部分、记录部分和电源部分；附件组成通常包括电极、电缆；运动单元由提供不同强度的设置单元、指示单元、运动部件组成。运动单元可提供不同负荷运动，主机及附件部分可监测受试者在运动过程中的心电信号，对信号进行处理、实时显示。用于实时检测患者运动状态下的心电图变化，供临床诊断。 | Ⅱ
351 | 07-03-01 | 心电图机 | 心脏在机械收缩之前，首先产生电激动，心肌激动所产生的微小电流可经过身体组织传导到体表，使体表不同部位产生不同的电位，对这种心脏动作电位的可见记录就是心电图。心电图机利用体表放置的通过患者电缆联接到设备的两个电极，按照心脏激动的时间顺序，依次记录并经过处理后输出或显示体表两点间的电位差。心电图机一般为台式或手提式，由主机、患者电缆和电极组成，电极也可分为可重复使用和一次性使用两种形式。记录方式可采用热笔式或热线阵记录方式等。有些产品具有信号输入或信号输出端口，可按技术参数和功能不同分为若干型号，如可按同步测量心电导联数不同分为单道、双道、多道等，也可将采集的心电信号输入计算机进行处理（如高频信号处理等），但不包括辅助分析和诊断功能。产品性能指标采用下列参考标准中的适用部分，如：YY 1139－2013心电诊断设备；单道和多道心电图机，YY 0782-2010医用电气设备 第2-51部分：记录和分析型单道和多道心电图机安全和基本性能专用要求等。 | Ⅱ
352 | 07-03-01 | 动态心电记录仪 | 动态心电记录仪通常由移动式记录器和导联线组成，也可包括数据管理软件和附件(导联线、充电器、USB线等)，可按导联数、技术参数、预期用途、附加辅助功能等不同分为若干型号；供人体24小时动态心电监测用。不包括自动分析诊断功能。产品性能指标采用下列参考标准中的适用部分，如：YY 0885-2013 医用电气设备 第2部分：动态心电图系统安全和基本性能专用要求。 | Ⅱ
353 | 07-03-01 | 胎儿心电图机 | 胎儿心电图机主要由主机、导联线、电极组成，可包括数据管理软件、打印设施及附件；可按导联数、技术参数、预期用途、附加辅助功能等不同分为若干型号；供胎儿心电检查、监测用，但不包括辅助分析和诊断功能。 | Ⅱ
354 | 07-03-03 | 汞柱式血压计 | 汞柱式血压计可有台式、立式和挂式等型式，基本参数：测量范围：0—40kPa（0—300mmHg），儿童血压计0—20kPa（0—150mmHg）；采用双刻度标尺，最小分度值为0.5 kPa（2 mmHg）；示值允许误差为±0.5kPa（±3.75 mmHg）；袖带可根据气囊尺寸不同分为多个规格；小儿用参数可调整；供测量人体血压用。产品性能指标采用下列参考标准中的适用部分，如： GB 3053-1993 血压计和血压表。 | Ⅱ
355 | 07-03-03 | 机械弹性元件式血压表 | 机械弹性元件式血压表采用双刻度，成人用测量范围0—40kPa（0—300mmHg），最小分度值分别为0.5kPa和2mmHg，允许误差为±0.5kPa（±3.75mmHg），袖带可根据气囊尺寸不同分为多个规格；小儿用参数可调整，可按型式、技术参数及适用对象不同分为若干型号；供测量人体血压用。产品性能指标采用下列参考标准中的适用部分，如： GB 3053-1993 血压计和血压表。 | Ⅱ
356 | 07-03-03 | 无创自动测量血压计 | 无创自动测量血压计主要由主机、充气管、血压袖带和电源组成。 | Ⅱ
该产品采用示波法或类似的其他方法，无创地测量人体动脉血压。
该产品应符合YY 0670-2008无创自动测量血压计，制造商应提供关于血压测量整体有效性的临床验证报告。
357 | 07-03-04 | 额温计 | 通常由红外温度传感器、探头套、显示单元、供电电路、测量电路组成。采用红外感温方法测量温度显示或者数据输出。应提供临床准确性的临床验证资料。 | Ⅱ
358 | 07-03-04 | 体温测量设备（无源） | 通常由玻璃管、感温泡、汞或其他感温液体和刻度尺标组成。采用汞或其他液体的热胀冷缩原理测量温度。用于临床测量患者体温。通常放置于人体的口腔、腋下、肛门部位测量。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
359 | 07-03-04 | 医用电子体温计 | 医用电子体温计为间歇接触式监控患者体温的电子装置，可由塑胶外壳、电路板、温度测量部件、显示屏、电源等组成；可按设计、技术参数、附加辅助功能、预期用途等不同分为若干型号，供测量人体体温或女性监测排卵周期等用，不包括预测模式，也不包括医用红外体温计。产品性能指标采用下列参考标准中的适用部分，如：GB/T 21416-2008 医用电子体温计。 | Ⅱ
360 | 07-03-04 | 耳腔式医用红外体温计 | 耳腔式医用红外体温计可由塑胶外壳、电路板、温度测量部件、显示屏、电源、隔离膜等组成；可按设计、技术参数、附加辅助功能、预期用途等不同分为若干型号；通过热辐射显示被测人体耳腔体温。需提供临床准确度与临床重复性报告。产品性能指标采用下列参考标准中的适用部分，如：GB/T 21417.1-2008医用红外体温计 第1部分：耳腔式。 | Ⅱ
361 | 07-03-05 | 医用脉搏血氧监测仪 | 医用脉搏血氧监测仪通常由主机（包括控制、数据处理/显示模块）和与之相联的血氧探头组成，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；通过建立光辐射与人体组织血氧饱合度间关系，供临床无创估算监测人体动脉血氧饱合度和脉率用。需提供血氧准确度人体评估报告。产品性能指标采用下列参考标准中的适用部分，如：YY 0784-2010医用电气设备——医用脉搏血氧仪设备基本安全和主要性能专用要求。 | Ⅱ
362 | 07-03-06 | 脑电图机 | 脑电图机可由放大器输入盒、计算机、显示器、专用软件、电极组成，也可包括闪光灯、闪光控制器、打印机等，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号，供人体脑电生理信号检测、处理、显示和储存等用。产品性能指标采用下列参考标准中的适用部分，如：GB9706.26-2005 医用电气设备第2-26部分：脑电图机安全专用要求。 | Ⅱ
363 | 07-03-06 | 脑电地形图仪 | 脑电地形图仪主要由主机、显示器、放大器、打印机、键盘、鼠标等组成，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；供人体脑电生理信号数据定量处理、分析功能，不包括自动诊断部分。 | Ⅱ
364 | 07-03-06 | 肌电图机 | 肌电图机主要由主机、刺激电极组成，也可包括计算机、专用数据管理软件及打印机，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；供人体肌电信号及神经传导数据采集、处理、分析用，不包括自动诊断部分。产品性能指标采用下列参考标准中的适用部分，如：YY 0896-2013医用电气设备第2部分：肌电及诱发反应设备安全专用要求。 | Ⅱ
365 | 07-03-09 | 经皮黄疸测试仪 | 经皮黄疸测试仪为手持式，由光学测试单元、显示单元、电源/电池及充电附件等组成，可按工作原理、设计、技术参数、适用部位、适用对象、辅助功能等不同分为若干型号；通过特定波长光反射，主要用于新生儿经皮黄疸测量用。 | Ⅱ
366 | 07-03-11 | 鼻阻力测量仪 | 通常由主机、取压管和流量传感器组成。用于通过检测鼻腔气体流动参数，分析鼻腔气道阻力。 | Ⅱ
367 | 07-04-01 | 多参数监护仪 | 多参数监护仪主要由主机和附件组成，附件可以包括心电导联电缆、血压袖带、血氧探头、体温探头、呼吸气体测量组件等，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；供患者的心电、无创血压、脉搏、血氧饱和度、体温、呼吸、呼吸气体监测用。需提供无创血压系统整体有效性和血氧准确度人体评估报告。产品性能指标采用下列参考标准中的适用部分，如： YY 1079-2008心电监护仪、YY 0670-2008无创自动测量血压计、YY 0784-2010医用电气设备——医用脉搏血氧仪设备基本安全和主要性能专用要求、YY 0785-2010临床体温计——连续测量的电子体温计性能要求、YY 0667-2008 医用电气设备——自动循环无创血压监护设备的安全和基本性能专用要求等。 | Ⅱ | 修订产品描述。
368 | 07-04-01 | 母亲/胎儿多参数监护仪 | 母亲/胎儿多参数监护仪主要由主机和附件组成，附件可以包括心电导联电缆、血压袖带、血氧探头、体温探头、超声探头、宫缩压力探头、标记手柄等，可按设计、型式、技术参数、附加辅助功能等不同分为若干型号；供母亲的心电、无创血压、脉搏、血氧饱和度、体温、呼吸、宫缩压力以及胎儿胎心率、胎动监测用。需提供无创血压系统整体有效性和血氧准确度人体评估报告。产品性能指标采用下列参考标准中的适用部分，如： YY 1079-2008心电监护仪、YY 0670-2008无创自动测量血压计，YY 0784-2010医用电气设备——医用脉搏血氧仪设备基本安全和主要性能专用要求，YY 0785-2010临床体温计——连续测量的电子体温计性能要求等。 | Ⅱ
369 | 07-05-01 | 听力计 | 通常由电源、声卡、功率放大器、控制器、电声换能器件、操作软件和患者应答器组成。用于测定个体对各种频率感受性大小的仪器，通过与正常听觉相比，就可确定被试的听力损失情况。 | Ⅱ
370 | 07-05-02 | 电声门图仪 | 通常由主机、声门图电极和音频输出线组成。声带振动时，声带接触阻抗变化引起调制电流变化，形成电声门图。用于检测声门组织阻抗变化和声带接触面积的变化，反映声带振动每一周期中声门闭合阶段的特点以及声带振动时每个周期的运动轨迹。 | Ⅱ
371 | 07-05-04 | 中耳分析仪 | 由主机、探头、耳机等组成，用于对中耳的声阻抗和静态压进行测试，适用于中耳病变的诊断。 | Ⅱ
372 | 07-06-03 | 伽玛射线探测仪 | 通常由固体闪烁探测器和电子学元件组成。用于查找体内放射性活度分布，可在手术中使用。 | Ⅱ |  | 新
373 | 07-07-01 | 超声多普勒血流分析仪 | 超声多普勒血流分析仪利用超声多普勒频移原理，用来探查、测量非胎儿的血流的运动信息。通常由探头（一般采用单元探头）、超声波发射/接收电路、信号处理和显示等部分组成；可按机型、成像性质（模拟与数字）、探头配置、技术参数、附加辅助功能、软件处理、预期用途等不同分为若干型号；主要用于经颅、颈部和外周血管的血流测量，不可在手术中使用，不包括分析诊断功能。参考或适用标准：YY 0593-2015 超声经颅多普勒血流分析仪等。 | Ⅱ
374 | 07-07-01 | 超声多普勒脐带血流分析仪 | 超声多普勒脐带血流分析仪由主机和超声多普勒探头组成，可按设计、技术参数、辅助功能等不同分为若干型号及功能；采用多普勒超声技术，对围产期胎儿脐带血流进行检测/监护用。 | Ⅲ
375 | 07-07-02 | 超声骨密度仪 | 超声骨密度仪利用超声波传导速度的差异和振幅的衰减来反映人体骨矿含量、骨结构以及骨强度的情况，通过测量跟骨、胫骨和/或桡骨的超声速度（SOS）和/或宽带超声衰减（BUA），反映骨骼密度状况。适用范围不包括骨质疏松、骨损失等相关疾病的诊断结论。主要包括超声波发射电路、接收电路、信号处理显示部分和记录部分等组成；可按设计、技术参数、附加辅助功能等不同分为若干型号；根据临床数据库人群分布，适用于相应年龄段人群；经体表用于人体骨密度的测量。产品性能指标采用下列参考标准中的适用部分，如：YY 0774-2010超声骨密度仪等。 | Ⅱ
376 | 07-08-03 | 多参数中央监护系统 | 多参数中央监护系统分有线和无线两类。有线监护系统是由中央主机和床边多参数监护仪组成，无线监护系统由中央主机、发射机和床边多参数监护仪组成；可按设计、技术参数、辅助功能等不同分为若干型号；供医疗机构对患者多种生理参数进行集中监护、管理等。需提供无创血压系统整体有效性和血氧准确度人体评估报告。产品性能指标采用下列参考标准中的适用部分，如： YY 1079-2008心电监护仪、YY 0670-2008无创自动测量血压计，YY 0784-2010医用电气设备——医用脉搏血氧仪设备基本安全和主要性能专用要求，YY 0785-2010临床体温计——连续测量的电子体温计性能要求等。 | Ⅱ
377 | 07-09-01 | 电子尿量计量仪 | 电子尿量计量仪由主机（含显示终端）、配套尿袋、托架等组成（不含导尿管）；仪器可按技术参数、附加功能等不同分为若干型号；通过测量尿袋内尿液的重量和体积，计算尿量相关参数。 | Ⅱ
378 | 07-09-03 | 睡眠呼吸监测设备 | 睡眠呼吸监测仪通常由记录仪、脑电电极、眼动电极、肌电电极、胸/腹呼吸探头、体位传感器、鼻气流管、脉搏血氧探头和心电电极等组成。记录睡眠时的生理参数进行分析。用于记录睡眠时各种生理参数，对睡眠障碍、睡眠呼吸紊乱和睡眠呼吸暂停、低通气综合征疾病进行分析、诊断。 | Ⅱ
379 | 07-09-04 | 平衡测试系统 | 通常由测量平台、辅助支架、平衡测试软件等组成。患者站立或坐在测试平台上，对患者平衡能力进行评估。用于测试人体平衡能力。 | Ⅱ
380 | 07-09-08 | 酸碱度检测仪 | 通常由带PH值传感器的导管、数据记录装置组成。用于监测患者胃和食道的pH值等生理参数，对胃食道反流疾病进行检查、诊断。 | Ⅱ
381 | 07-10-03 | 一次性使用心电电极 | 一次性使用心电电极是心电图采集设备附件，由传感元件和电解质组成，可带或不带连接导线。电极可由基衬材料、导电膏、电极扣等组成，基衬材料可用透气纸、水刺布、无纺布、发泡纸、棉布或PE等加涂医用压敏胶制成，并可按其形状不同分为圆形、椭圆型、方形等，并按其尺寸大小不同分为若干规格；配合仪器，供心电检测、监测用。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0196－2005一次性使用心电电极。 | Ⅱ
382 | 07-10-03 | 一次性使用无创脑电电极 | 一次性使用无创脑电电极由电极片（导电塑料及金属镀层）、导线和设备接插件组成，可一端与脑电诊断/监护设备连接，一端与人体头部皮肤连接，供采集、传输脑电生物信号用。单个病人短时间内可以重复使用。 | Ⅱ
383 | 07-10-03 | 体表参考电极 | 一般由电极（粘性贴片、传感器）和连接线组成，按其形状可分为圆形、椭圆型、方形等。非灭菌产品，一次性使用。在心脏电生理检查中使用，应与相应电生理标测设备或导管配合使用，用于传送导管的位置。产品性能指标采下列参考标准中的适用部分，如：YY/T0196－2005一次性使用心电电极。 | Ⅱ
384 | 07-10-04 | 无创医用传感器 | 包括温度、压力、重力、角度等医用传感器。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0781-2010血压传感器等。有特别要求的按特别要求处理。 | Ⅱ
385 | 07-10-05 | 医用导电膏 | 医用导电膏通常为水基质高分子凝胶，内含生物相容性较好且稳定的电解质，用于心电、脑电、肌电等检查和电疗电极耦合用。 | Ⅱ | 修订产品描述。
386 | 08-02-02 | 麻醉穿刺针 | 通常由衬芯座、针座、针管、保护套、衬芯组成。针管一般采用不锈钢材料制成；衬芯一般采用不锈钢或塑料材料制成。用于对人体进行穿刺，注射药物。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
387 | 08-02-02 | 一次性使用麻醉穿刺针 | 性能、结构、组成、用途等属于YY 0321.2 一次性使用麻醉用针可完全涵盖的产品。产品一般由衬芯座、针座、针管和衬芯等组成，在麻醉时进行穿刺、注射药物。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
388 | 08-02-02 | 一次性使用麻醉用过滤器 | 性能、结构、组成、用途等属于YY 0321.3一次性使用麻醉用过滤器可完全涵盖的产品。一次性使用麻醉用过滤器一般由上盖、滤膜、底座、内圆锥接头保护套（可无）、外圆锥接头保护套（可无）组成，分为药液过滤器和空气过滤器。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
389 | 08-03-05 | 人工复苏器（简易呼吸器） | 通常由进气阀、压缩单元（如气囊）和患者阀组成，一般配有储气袋、呼吸面罩等附件。是一种通过操作者按压设备上压缩单元（如气囊），从而实现向患者肺部通气的复苏装置。用于供电供气不完备场合和紧急情况下对突发呼吸困难或呼吸衰竭的患者实施人工呼吸急救时提供肺通气。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
390 | 08-04-01 | 医用分子筛制氧系统 | 通常由空气压缩系统、气源净化系统、空气罐、医用分子筛吸附分离系统、成品气罐、控制系统、监测和报警系统等组成。一种利用分子筛变压吸附原理，通过分离大气中的氮气，从空气中富集氧气，用于生产富氧空气（93%氧）的气源系统。用于生产富氧空气（93%氧），经医用气体管道系统向其他用氧医疗器械提供气源，并按其临床适用范围向患者供氧。 | Ⅱ | 修订产品描述。
391 | 08-04-02 | 小型分子筛制氧机 | 通常由空气压缩泵、医用分子筛吸附分离系统、氧罐、输出流量控制显示装置、氧浓度监测装置、计时装置和报警系统组成。一种利用分子筛变压吸附原理，通过分离大气中的氮气，从空气中富集氧气，用于生产富氧空气（93%氧）的设备。用于生产富氧空气（93%氧），按其临床适用范围向患者供氧。 | Ⅱ | 修订产品描述。
392 | 08-04-03 | 医用膜分离制氧系统 | 利用膜分离技术原理，从空气中富集氧气，获得含氧量90%以上的富氧空气。通常由空气压缩与预处理设备、医用膜分离制氧主机、控制与监测仪器仪表、管道附件等组成。在医疗机构内建立并使用，经医用气体管道系统向其他用氧医疗器械提供气源，并按其临床适用范围向患者供氧。 | Ⅱ
393 | 08-04-04 | 小型医用膜分离制氧机 | 利用膜分离技术原理，从空气中富集氧气，获得含氧量90%以上的富氧空气。通常由空气压缩与预处理设备、医用膜分离制氧主机、控制与监测仪器仪表、管道附件等组成。可由制氧机主机和附件（湿化器、吸氧管/面罩等）组成，具备氧浓度过低的提示功能，可按单位时间制氧量、附加组件、功能等不同分为不同型号，供需要按照处方规定进行高浓度辅助供氧的患者使用。 | Ⅱ
394 | 08-04-05 | 氧气发生器 | 通常由筒体和上盖组成，可配有吸氧管等附件。以水为原料，利用水电解制氧剂使制氧剂发生化学反应，从而产生医用氧气，并向患者提供临床呼吸用氧。 | Ⅱ
395 | 08-05-01 | 麻醉蒸发器 | 麻醉蒸发器有机械调节机械控制式、机械调节电子控制式和电子调节电子控制式等不同实现方式，可配用的麻醉剂有七氟醚、异氟醚、地氟醚等。该产品一般由浓度控制装置、加药装置（包括罐充瓶适配器等）和麻醉剂液面指示器等组成。通过浓度控制装置在输出口形成含有一定百分比浓度的麻醉蒸气气流。麻醉蒸发器是麻醉机的组成部件，用于提供浓度可控的麻醉剂蒸气。 | Ⅲ
396 | 08-05-02 | 医用呼吸道湿化器 | 医用呼吸道湿化器通常与呼吸机或氧气吸入器配套使用，一般由温度控制系统、湿化室和加热装置组成，或仅由进气口、贮水瓶、湿化室和出气口组成。是一种用于提高输送给患者的医用气体湿度水平的设备或装置。用于湿化输送给患者的呼吸气体。 | Ⅱ
397 | 08-05-03 | 呼吸系统过滤器 | 呼吸系统过滤器可由上盖、过滤介质、下盖、护帽等四个部分组成，过滤器壳体可ABS等材料制成，过滤介质可用聚丙烯复合材料等制成，可按过滤介质、外形、尺寸等不同分为若干型号和规格， 以无菌形式提供，与麻醉呼吸设备和肺功能仪相配套，供降低患者吸入或呼出颗粒性物质（包括微生物）的数量用。产品性能指标采用下列参考标准中的适用部分，如： YY/T 0753 麻醉和呼吸用呼吸系统过滤器。 | Ⅱ
398 | 08-05-04 | 一次性使用热湿交换器/过滤器 | 产品可由外壳、滤芯、热湿交换介质和接口等组成。一次性使用，清洁包装或者无菌包装。本产品适用于建立人工气道的患者,对其呼吸的气体进行热湿交换和/或过滤微粒。产品推荐标准 YY/T 0735 麻醉和呼吸设备用于加湿人体呼吸气体的热湿交换器。 | Ⅱ
399 | 08-05-05 | 呼吸管路辅助器械 | 通常用于实现气道产品间的连接，或辅助插入气道等的附件。用于气道连接、取样、导入等功能的附件。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
400 | 08-05-05 | 插管导入器 | 插管导入器可由导管、加硬套管、接头等部件组成。产品经灭菌，一次性使用。插管导入器应用于当患者的声门显露不完全时,辅助将气管插管插入患者气管内。 | Ⅱ
401 | 08-05-05 | 肺叶通气功能检查仪专用导管及接头 | 通常产品包含导管一个，连接装置一个和操作手册一本。无菌装置，与肺叶通气功能检查仪连接，通过支气管镜的工作管道使用。产品用于在由支气管镜检医生在支气管镜检查室内为患有慢性阻塞性肺部疾病（COPD）和肺气肿的成年患者进行诊断性支气管镜检查的过程中使用。 | Ⅱ | 新增
402 | 08-05-06 | 麻醉咽喉镜 | 通常由手柄、窥视片、内部电源和照明用光源组成，可带有视频显示功能。是一种气管插管时使用的辅助器械。 | Ⅱ
403 | 08-05-06 | 麻醉视频喉镜 | 麻醉视频喉镜由手柄、摄像头和窥视片组成，手柄另一端可接显示器。窥片可用金属制成，重复使用，也可用高分子材料制成，一次性使用；手柄与窥片通常为可分离式，可按适用人群、材质、设计、技术参数、窥片尺寸等不同分为若干型号和规格；供临床挑起患者会厌部暴露声门，指引医护人员准确进行气道插管供麻醉或急救用，也可用于口腔内诊察、治疗。 | Ⅱ
404 | 08-05-06 | 一次性使用麻醉咽喉镜 | 一次性使用麻醉咽喉镜可由窥视片和镜筒组成，窥视片可分为直接照明和光纤照明两种；窥视片采用聚氯乙烯材料制成，镜筒应采用高密度聚乙烯树脂制成；可按使用对象、型式、尺寸等不同分为若干型号与规格；窥视片以无菌形式提供，配合麻醉咽喉手柄，供临床喉部检查、急救及麻醉手术时协助插管用。产品性能指标采用下列参考标准中的适用部分，如：YY 0499-2004 麻醉喉镜通用技术条件。 | Ⅱ
405 | 08-05-06 | 一次性使用麻醉窥视片 | 一般由医用高分子材料（如：PC）制成；产品由锁扣、套管和压舌板三部分组成。供临床挑起患者会厌部暴露声门，指引医护人员准确进行气道插管供麻醉或急救用，也可用于口腔内诊察、治疗。 | Ⅱ
406 | 08-05-07 | 网式雾化器 | 网式雾化器可由主机（控制部分）、雾化装置、盛液容器、喷嘴/面罩等组成，可按结构、组成、技术参数、附加辅助功能等不同分为若干型号，通过压电陶瓷振动带动雾化装置运动，将药液从孔穴挤出，雾化药物供吸入治疗用。产品性能指标采用下列参考标准中的适用部分，如：YY 0109-2003超声雾化器。 | Ⅱ |  | 新
407 | 08-05-07 | 雾化设备/雾化装置 | 通过超声波、自带的电动泵、外接气源等方式进行雾化。是一种用于把液体转化为气雾剂的设备或装置。 | Ⅱ
408 | 08-05-07 | 压缩式雾化器 | 压缩式雾化器由往复活塞微型压缩机、连接软管、汽雾手动开关和喷嘴、软管、面罩组成；可按技术参数、附加功能等不同分为若干型号；通过压缩气体产生的气流雾化药物并将其输送到呼吸道，供呼吸道雾化药物吸入治疗用。 | Ⅱ
409 | 08-05-07 | 医用超声雾化器 | 医用超声雾化器可由主机（含雾化及控制部分）及雾化输送附件组成，可按结构、组成、技术参数、附加辅助功能等不同分为若干型号，雾化药物供治疗用。产品性能指标采用下列参考标准中的适用部分，如：YY 0109-2003超声雾化器。 | Ⅱ
410 | 08-05-07 | 一次性使用气流雾化器 | 在呼吸装置系统中，一端接入氧气，用氧气的压力将放置于雾化器中的药剂雾化，产生的雾化气体通过雾化器的另一端吸入病人口中。 | Ⅱ
411 | 08-05-07 | 加药喷雾瓶 | 产品一般可采用医用聚氯乙烯、硅橡胶、聚丙烯、聚醋酸酯等材料制成。分为面罩式和口含式，可由加药喷雾瓶、药杯、微量吸入器等部件组成。用于雾化药物并将其给入患者口腔或鼻腔中。 | Ⅱ
412 | 08-05-08 | 麻醉储气囊 | 通常由呼吸袋和连接件组成，是麻醉机的麻醉呼吸系统中储存气体的弹性容器。 | Ⅱ
413 | 08-05-09 | 一次性使用麻醉废气吸附器 | 一次性使用麻醉气体吸附器主体结构主要由封装盖、分气头、滤板、上罩、下座构成，采用符合GB/T 21460.2-2008的塑料制成；吸附介质应采用符合GB/T 7701.1-2008表3规定的活性炭（不包括粒度指标），粒度指标应符合表4的规定。活性炭填充量不小于80g,平均吸附率≥90%。产品以非灭菌方式提供。供临床吸附麻醉管路排出的残留麻醉气体用。 | Ⅱ
414 | 08-05-10 | 麻醉气体净化传递和收集系统 | 通常由连接口、储气罐等组成，用于传输和收集麻醉废气并进行集中净化处理的装置。用于输送呼出和/或排除的多余麻醉气体至适当排放处。 | Ⅱ
415 | 08-06-01 | 一次性使用麻醉导管及接头 | 性能、结构、组成、用途等采用下列参考标准中的适用部分，如YY/T0321.1-2022。产品一般由导管和导管接头组成，与一次性使用麻醉用针配合，供临床输送麻醉剂用。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
416 | 08-06-02 | 呼吸管路 | 通常为 “人”字形或“一”字形结构的波纹管，部分管路可以做轴向伸缩，人字形结构的管路由吸气支路和呼气支路组成，一般由塑料或硅橡胶材料制成的。具有加热功能的呼吸管路还包括加热丝和电源适配器。一次性或重复使用。或通常为“一”字形结构的波纹管或软管，也可包含咬嘴，一般由塑料或硅橡胶材料制成。无菌提供。 | Ⅱ
417 | 08-06-02 | 麻醉机和呼吸机用呼吸管路 | 麻醉机和呼吸机用呼吸管路采用聚氯乙烯等适用材料制成，分为麻醉机管路和呼吸机管路两类，麻醉机管路由Y型接头、波纹管组成，呼吸机管路由Y型接头、波纹管、旋转弯管组成，两类管路可组合，可选配储气囊、面罩，可按材质、设计、组合、使用对象（成人/儿童）、技术参数、选配件等不同分为若干型号和规格；以无菌形式提供；用于麻醉机、呼吸机接口与供氧面罩的管路连接。产品性能指标采用下列参考标准中的适用部分，如：YY 0461-2003麻醉机和呼吸机用呼吸管路。 | Ⅱ
418 | 08-06-03 | 气管内插管/气管套管 | 常见的插管头部有一个或两个套囊，套囊充起后可以起到固定插管和密封气道的作用，也可以不带套囊。插管管身通常由高分子材料制成，管身内埋有钢丝线圈，以提高径向强度和轴向柔软度。部分插管管身采用抗激光材料或复层，以抗激光照射。为经鼻/口或经皮插入病人气管的插管。一端通过呼吸管路与麻醉呼吸机连接，以维持病人呼吸。无菌提供，一次性使用。或通常由底板、内套管、外套管和管芯组成。 | Ⅱ
419 | 08-06-03 | 一次性使用气管插管 | 一次性使用气管插管通常由气管插管管体、充气管、套囊、接头和充气阀组成；其中气管插管的管体、充气管、套囊可采用聚乙烯等塑料、硅橡胶等适宜材料制成，接头、充气阀可采用聚乙烯等材料制成；气管插管可分为普通有/无套囊插管和加强型有/无套囊插管等不同型号，每型可按公称内径、套囊等不同分成若干规格；产品应以无菌形式提供；供临床麻醉或急救时建立人工气道用。产品性能指标采用下列参考标准中的适用部分，如：YY 0337-2002 气管插管。不包括支气管插管。 | Ⅱ
420 | 08-06-03 | 一次性使用气管切开插管 | 一次性使用气管切开插管由气管切开插管管体（病人端有圆滑尖部、斜面和固定球囊）、标准外圆锥接头（机器端）、一体式指示球囊/阀组合和固定翼组成；其中，气管切开插管的管体、充气管、套囊可采用聚乙烯塑料、橡胶等适宜材料制成，接头、充气阀可采用聚乙烯等材料制成；气管切开插管可分为普通有/无套囊切开插管和加强型有/无套囊切开插管等不同型号，每型可按适用人群、公称内径、套囊容量等不同分成若干规格；产品应以无菌形式提供；供临床急救气管节切开，建立人工气道用。产品性能指标采用下列参考标准中的适用部分，如：YY 0338-2002 气管切开插管。 | Ⅱ
421 | 08-06-03 | 支气管插管 | 产品可由接头、管身、气管腔/套囊/指示球囊、支气管腔/套囊/指示球囊、充气管/阀、导丝等部件组成。产品经灭菌，一次性使用。用于主干支气管插管，允许选择性的进行充气/吸气、吸痰和对任一肺进行支气管镜检查。 | Ⅱ
422 | 08-06-03 | 支气管双腔插管 | 支气管双腔插管可由带探针的支气管双腔插管和接头等部件组成。产品经灭菌，一次性使用。支气管双腔插管放置于支气管主干内，可用于选择性的充气或放气、吸痰、单侧肺通气，或用于支气管镜检查时使用。 | Ⅱ
423 | 08-06-03 | 气道交换导管 | 气道交换导管可由导管、接头等部件组成。本产品经灭菌，一次性使用产品。用于在需要时进行气管插管的交换。 | Ⅱ
424 | 08-06-05 | 喉罩 | 通常由套囊、充气管、喉罩插管、机器端、接头、指示球囊等组成，插到喉部后通过充起套囊堵塞口腔和食道，同时又能使患者气管保持畅通。一般无菌提供，一次性使用。 | Ⅱ
425 | 08-06-05 | 一次性使用医用喉罩 | 一次性使用医用喉罩可用硅橡胶等制成，由套囊、充气管、喉罩插管、机器端、接头、指示球囊等组成，喉罩插管可为金属加强型，可按材质、设计、技术参数、辅助功能等不同分为若干型号及规格；供全麻或心肺人工复苏建立短期人工气道用。 | Ⅱ
426 | 08-06-06 | 口咽/鼻咽通气道 | 一次性使用口/鼻咽通气道用高分子材料制成，为带有凸缘末端的呼吸管，可按材质、设计、插入部位管径/长度等不同分为若干型号及规格；可以无菌形式提供，供建立口/鼻咽通气道，防止舌后坠引起气道阻塞等用。 | Ⅱ
427 | 08-06-07 | 支气管堵塞器 | 通常由导管、导管座、球囊、球囊充盈接头、多路气道转换接头等器件组成。球囊供插入支气管插管充起后可以临时封堵支气管。用于需要单肺通气的手术中，插入患者的支气管内，达到阻断左肺气道或右肺气道的目的。 | Ⅱ
428 | 08-06-08 | 鼻氧管 | 通常由进氧接口、氧气软管、调节环、鼻塞（或面罩）等组成。鼻氧管与输氧系统连接，供患者吸入氧气使用。一次性使用。用于吸氧时氧源与吸氧者之间的氧气直接输送或湿化后输送。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
429 | 08-06-09 | 呼吸道用吸引导管 | 呼吸道用吸引导管通常由有端孔的导管和带塞接头组成；导管可采用聚氯乙烯、天然胶乳、硅橡胶、橡胶材料制成，带塞接头可相应采用聚氯乙烯、硅橡胶或MD型聚氯乙烯制成；导管按材质、管径等不同分为若干型号与规格，产品应无菌形式提供。可与真空负压系统或设备连接，供呼吸道吸引痰液等用。不包括没有端孔的多腔呼吸道吸引导管。产品性能指标采用下列参考标准中的适用部分，如：YY 0339-2009 呼吸道用吸引导管。又名一次性使用吸痰管。 | Ⅱ
430 | 08-06-10 | 呼吸面罩 | 通常由头带、前额软垫、面罩架、接口盖、松紧带和硅胶密封罩组成，直接与患者接触经患者口、鼻腔通气。 | Ⅱ
431 | 08-06-11 | 持续正压通气用面罩、口罩、鼻罩 | 通常由头带、前额软垫、面罩架、接口盖、松紧带和硅胶密封罩组成，直接与患者接触经患者口、鼻腔通气。 | Ⅱ
432 | 08-06-11 | 持续正压通气用面罩、口罩、鼻罩 | 通常由鼻罩、口罩或口鼻罩主体，固定头带、气路接口、排气口、防窒息阀等组成；直接与患者面部接触，经鼻腔和/或口腔通气的界面连接装置。用于慢性呼吸功能不全、改善通气和睡眠治疗等无创通气支持。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
433 | 08-06-13 | 麻醉面罩 | 通常由接口、气囊、罩体组成。组成面罩的材料有塑料等。一次性使用或重复使用均可。 | Ⅱ
434 | 08-06-13 | 一次性使用麻醉面罩 | 一次性使用麻醉面罩可采用硅橡胶等高分子材料制成，可由面杯、气囊、气门芯组件（气门芯阀体、阀芯、密封圈、弹簧和底盖）和固定圈组成，可按材质、设计、技术参数、预期用途、辅助功能等不同分为若干型号及规格；供临床麻醉呼吸用。 | Ⅱ
435 | 08-06-14 | 输氧面罩 | 通常由面罩和连接管等组成。采用医用高分子材料制成。用于对缺氧患者进行输氧，作为氧气进入患者体内的通道。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ | 新增
436 | 08-07-01 | 医用空气压缩机 | 与呼吸机等配套使用，用于制取医用压缩空气，为呼吸机或类似呼吸通气系统提供压缩空气源。 | Ⅱ | 新增
437 | 08-07-02 | 医用气体混合器 | 通常由主机（气体压力平衡处理模块、气体比例分配控制模块、气体差压或缺压报警模块）、出气口、氧气进气口、空气进气口、氧气浓度调节装置等组成。 | Ⅱ | 修订产品名称和产品描述。
用于对输入的医用氧和空气按照设定浓度进行混合，对空氧混合气体的氧浓度和流量调节和控制。
438 | 08-07-03 | 医用供氧器 | 医用供氧器由流量计、压力表、减压器、推车等，供医疗急救或医疗现场供氧用。 | Ⅱ
439 | 08-07-03 | 便携式医用供氧器 | 通常由气雾阀、吸氧面罩或鼻氧管等部分组成（不含氧）。用于轻度缺氧的预防。 | Ⅱ | 修订产品描述。
440 | 08-07-04 | 医用空气压缩机组 | 医用空气压缩机组用于医用气体导管系统，可由气罐、连接装置和电控箱组成；产生压缩气体供诊疗用。 | Ⅱ | 修订产品分类编码、产品名称和产品描述。
441 | 08-07-04 | 医用气体管道系统－供氧系统 | 医用气体管道系统－供氧系统通常包含一个中心供氧站（可以是氧气浓缩器、医用氧站或两者的结合）、一套监测和报警系统和具有若干终端的管道分配系统组成，可按中心供氧站类型等不同分为若干型号，系统具欠压等声、光报警等功能；供医疗机构集中供氧用。主要产品性能指标采用下列参考标准中的适用部分，如：YY 0801-2010医用气体管道系统终端，YY/T 0799-2010 医用气体低压软管组件，YY 1468-2016 用于医用气体管道系统的氧气浓缩器供气系统，YY/T 0187 医用中心供氧系统通用技术条件。 | Ⅱ
442 | 08-07-05 | 医用气体汇流排 | 通常由供电装置、气体阀门、流量控制阀、气体偏差控制器等组成，不含气瓶。 | Ⅱ
443 | 08-07-06 | 医用气体报警系统 | 医用气体报警系统可由多种多点气体传感器械及显示传输系统构成的监测点通过有线连接方式集中传输至中央气体管理系统，由中央管理系统对医用气体监测点进行集中监测，超出预警值可自动报警；可用于对医院供气系统的各种气体如氧气、麻醉气体等进行监测。 | Ⅱ
444 | 09 | 乳头内陷吸牵器 | 乳头内陷吸牵器由筒形管、内芯等组成，利用负压原理，将乳头吸牵到位；供治疗乳头内陷用。 | Ⅱ
445 | 09 | 微波热疗仪硅胶治疗导管 | 微波热疗仪硅胶治疗导管采用硅胶制成，通常由阀门、腔头、多腔管和球囊等组成，不带电极；可按材质、设计、技术参数、预期用途等不同分为若干型号及规格；以无菌形式提供；配合药物和微波治疗仪分别用于药物灌注和微波治疗。 | Ⅱ
446 | 09-01-03 | 中频电疗产品 | 产品由主机（信号产生及控制装置）、电极、导线及其他附属部件。应用频率为1kHz-100kHz的交流电（包括正弦波、脉冲波和调制波等）进行治疗、康复的方法。预期用途应体现临床适应证和治疗作用。如：该产品对肩周炎具有消炎和镇痛作用等。产品性能指标、临床治疗作用及适应证应在《中频电疗产品注册技术指导原则》范围内。 | Ⅱ
447 | 09-01-05 | 神经和肌肉刺激器用电极 | 神经和肌肉刺激器用电极由电极片（导电塑料及金属镀层）、导线和设备接插件组成，可按材质、技术参数、尺寸等不同分为若干型号及规格；可一端与神经和肌肉刺激器连接，一端与人体完好皮肤连接；产品可以无菌形提供，单个病人一定时间内可以重复使用；供向人体无创传输神经和肌肉刺激器刺激电信号用。产品性能指标采用下列参考标准中的适用部分，如： YY 0868-2011神经和肌肉刺激器用电极。 | Ⅱ
448 | 09-02-01 | 电脑恒温电蜡疗仪 | 通常由主机、熔腊装置、温度控制装置、温度检测装置、腊等组成。利用加热熔解的石蜡、蜂蜡作为导热体，将热能传至机体达到治疗作用的设备。 | Ⅱ
449 | 09-02-01 | 升温毯 | 升温毯一般由主机、毯子和体温传感器组成。在医疗机构临床使用环境下，通过控制毯子温度，具有对人体进行体外物理升温功能，达到辅助调节人体温度目的的设备。可按设计、技术参数、辅助功能等不同分为若干型号。 | Ⅱ
450 | 09-02-01 | 热敷贴（袋） | 通常含有发热材料，并封装于医用无纺布或其他医用材料内。不具有温度保护装置，使用时直接贴敷于患部，以传导的方式将热量传递于患处，用于促进局部血液循环、辅助消炎、消肿和止痛。 | Ⅱ
451 | 09-02-03 | 医用控温毯 | 通常由温度调节装置、温控电路、控制机构及应用部分组成。在医疗机构临床使用环境下，通过控制应用部分的温度实现对人体进行体外物理升温和/或降温功能，以达到辅助调节人体温度的目的。 | Ⅱ
452 | 09-02-03 | 降温毯 | 降温毯一般由主机、毯子和体温传感器组成。在医疗机构临床使用环境下，通过控制毯子温度，具有对人体进行体外物理降温功能，达到辅助调节人体温度目的的设备。可按设计、技术参数、辅助功能等不同分为若干型号。 | Ⅱ
453 | 09-03-06 | 新生儿蓝光治疗仪 | 新生儿蓝光治疗仪由蓝光灯，控制盒，支撑杆，底座及其配件组成。无其他混杂光源。用于治疗新生儿黄疸。YY 0669-2008 医用电气设备第2部分：婴儿光治疗设备安全专用要求适用于本产品。 | Ⅱ
454 | 09-03-08 | 导光鼻塞 | 与光治疗设备配合使用，其组成与原理取决于光治疗设备的型式和功能。主要用于连接或包裹导光原件的末端并插入鼻腔中，便于固定及避免交叉感染。 | Ⅱ
455 | 09-04-01 | 电动吸鼻器 | 电动吸鼻器由吸头、主机、电池组成。使用时，吸头接触儿童鼻孔外缘，由主机产生吸力。利用负压原理吸除儿童鼻腔分泌物。 | Ⅱ
456 | 09-04-01 | 振动叩击排痰机 | 振动叩击排痰机主要由主机、传动软轴和多种叩击头组成，叩击头转速、叩头频率、强度可调、叩击模式、工作时间可调，可有多路输出，可按设计、技术参数、适用对象等不同分为若干型号。通过振动叩击,改善患者肺部血液循环状况、协助排出呼吸道分泌物。 | Ⅱ
457 | 09-04-02 | 医用弹力袜 | 产品符合YY/T 0851防血栓袜或YY/T 08531医用静脉曲张压缩袜标准要求。用于促进静脉血液流速，预防深静脉血栓形成，缓解静脉曲张症状。 | Ⅱ
458 | 09-04-03 | 电动颈腰椎牵引装置 | 电动颈腰椎牵引装置由牵引动力系统、控制系统及附属部件构成，可按设计型式、技术参数、适用部位、预期用途等不同分为若干型号，可供颈椎适应症的牵引治疗。不包括快牵装置。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0697-2016 电动颈腰椎牵引治疗设备。 | Ⅱ
459 | 09-04-03 | 手动牵引床 | 通常由产生和调节机械力的牵引主机和传输力的绳索构成，也包括承载患者的床（椅）和配套的患者固定带等附件。牵引主机是手动结构，患者固定带绑在患者的枕、颌部、胸部、髋部或四肢等部位，通过皮肤摩擦力将牵引力传递至患者，可提供水平的颈椎、腰椎牵引，或垂直的颈椎牵引。 | Ⅱ
460 | 09-04-04 | 牵引器具 | 通过充气气囊或表面呈弧形的器具等装置固定在颈部、腰部，使得颈椎、腰椎能够保持一定的角度并对周围肌肉产生拉伸作用，用于放松脊椎周围肌肉，缓解椎间压力。 | Ⅱ
461 | 09-08-01 | 医用空气加压舱 | 一般由舱体(壳体、舱门、观察窗、安全阀等)、供排气系统、供排氧系统、空调系统、电气系统、水喷淋系统、控制台组成。加压介质为空气，最高工作压力不大于0.3MPa的氧舱。根据舱内治疗人数不同分为单人氧舱和多人氧舱。将病人暴露于高压氧环境中，使其呼吸高压氧，达到治疗疾病的目的。供需要按照医师处方进行辅助供氧治疗的成人患者使用，用于缺血、缺氧性疾病的辅助治疗。不包括软体舱、异型舱、婴幼儿氧舱。 | Ⅲ
462 | 09-08-01 | 医用加压氧舱 | 一般由舱体、供排氧系统、空调系统和控制系统组成。加压介质为医用氧气，最高工作压力不大于0.2MPa，进仓人数为1人。旨在供需要按照处方规定进行高浓度辅助供氧的成人患者使用，用于缺血、缺氧性疾病的辅助治疗。将病人暴露于高压氧环境中，使其呼吸高压氧，达到治疗疾病的目的。供需要按照医师处方进行辅助供氧的治疗的成人患者使用，用于缺血、缺氧性疾病的辅助治疗。不包括软体舱、异型舱、婴幼儿氧舱。 | Ⅲ
463 | 09-08-04 | 医用浸浴治疗机 | 医用浸浴治疗机由主机、病人浴床、供/排水系统、温控单元等组成，也可包括病人转运吊架及冲其他辅助浸浴功能，可按设计、技术参数、辅助功能等不同分为若干型号。用于烫伤病人浸浴处理。 | Ⅱ
464 | 09-08-05 | 肠道水疗机 | 通常由主机、温度控制装置、压力（流量）控制系统、液箱、蠕动泵、注液管、排液管等组成。治疗时将液体灌注到肠道内，同时通过排液管将液体引流到体外。在医疗机构用于对结肠的清洗。 | Ⅱ |  | 新
465 | 10 | 血路连接器 | 血路连接器（简称接头），属于心肺转流系统，供体外循环手术心肺转流系统中血路通道输送血液、观察和连接使用。豁免情况不包括使用了新材料、具有新作用机理、新功能的产品。 | Ⅲ
466 | 10-01 | 医用真空冷冻干燥机 | 医用真空冷冻干燥机通过将血浆等被冷冻干燥物质冻结成固态，而后使其中的水分从固态升华成气态，以除去水分而保存血浆的方法。主要由由制冷系统、真空系统、加热系统、控制系统等组成。 | Ⅱ
467 | 10-01-06 | 血液融化设备 | 通常由加热水箱、解冻槽、循环管路、进水管、排水管、控制箱组成，临床用于血浆或血液的加热、解冻。 | Ⅱ
468 | 10-01-06 | 冰冻血浆解冻箱 | 冰冻血浆解冻箱由加热水箱、解冻槽、循环管路、进水管、排水管、电器控制箱等组成，可具有报警等功能，可按结构组成、最大化浆量、解冻时间等技术参数、附加功能等不同分为若干型号，用于对临床血浆进行加热、解冻。 | Ⅱ
469 | 10-02-03 | 一次性动静脉穿刺针/器（内瘘针） | 通常由穿刺针管、软管等组成。一般由奥氏体不锈钢材料、聚氯乙烯等材料制成。无菌提供，一次性使用。配合血液成分采集机（如离心式、旋转膜式）或血液透析机等使用，用于从人体静脉或动脉采集血液，并将处理后的血液或血液成分回输给人体。豁免情况不包括采用新材料、新作用机理、新功能的产品。 | Ⅲ
470 | 10-02-04 | 传统型一次性使用输血器 带针 | 性能、结构、组成、用途等参考GB 8369.1《一次性使用输血器 第1部分：重力输血式》 或GB 8369.2 《一次性使用输血器 第2部分：压力输血设备用》 、GB 18671《一次性使用静脉输液针》可完全涵盖且原材料符合GB15593等相关标准要求的传统型一次性使用输液器，用于静脉输血。可带输液贴（位于单包装内）。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
471 | 10-02-07 | 一次性使用冰冻红细胞洗涤机用管路 | 通常由主管路（含硅橡胶泵管、大滴壶、连接管）、回血管和废液收集部分（含废液管路和废液袋）组成。无菌提供，一次性使用。用于将红细胞与保护液实施洗涤、分离时使用的管路。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
472 | 10-03-05 | 血液透析水处理设备 | 通常由罐式过滤器、活性炭过滤器、软化器、精密过滤器、反渗透装置、动力装置、消毒装置、监测装置和输送管道组成，用于制备血液透析用水。 | Ⅱ
473 | 10-03-05 | 血液透析器复用机 | 通常由控制系统、监测系统和水路系统组成，与专用消毒液联合使用，用于可重复使用透析器的冲洗、清洁、测试、专用消毒液灌注。 | Ⅱ
474 | 10-03-05 | 电动透析椅 | 通常由座位、靠背、搁脚板、滑动式脚踏板、靠枕、扶手、可锁定的脚轮、推手柄、控制器组成，用于电动调整背垫、坐垫、脚垫的位置，方便患者透析治疗。 | Ⅱ
475 | 10-03-05 | 血液透析用血流监测系统 | 通常由主机（电子流量计）、流量/稀释度感应器等组成。 | Ⅲ
476 | 10-03-07 | 腹透液袋加温仪 | 通常由加热板、电源和电缆连接器组成，用于腹透液袋使用前的加温。 | Ⅱ
477 | 10-04 | 腹膜透析附件 | 包含硅胶导管延长管、导管接头、固定器架台和圆盘固定器。导管延长管和导管接头为一次性使用无菌产品。用于辅助腹膜透析治疗。基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅱ
478 | 10-04 | 腹膜透析外接管 | 由管体、接头、锁管接头和保护帽等组成，不包含碘液保护帽等Ⅲ类医疗器械组件。无菌提供。在腹膜透析时与腹膜透析管外置部分连接使用。基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅱ
479 | 10-04-01 | 连续性肾脏替代治疗用管路 | 通常由动脉管路、静脉管路、置换液管路等组成，组件不包括血液透析滤过器、透析器等。产品无菌提供，一次性使用。配合连续性肾脏替代专用透析器或血液透析滤过器、连续性肾脏替代专用透析设备使用，用于连续性肾脏替代治疗中，承担血液通路的功能”。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ |  | 新
480 | 10-04-01 | 一次性使用血液透析管路 | 通常由动脉管路、静脉管路、置换液管和其他必要的配件组成，不包含任何血管通路器械（如透析导管、动静脉内瘘穿刺针等）和血液透析器、血液透析滤过器和血液灌流器等。无菌提供，一次性使用。与血液透析器、血液透析滤过器、血液灌流器配套使用，用于血液透析、血液透析滤过、血液灌流串联血液透析治疗中，承担血液通路的功能。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
481 | 10-04-01 | 一次性使用血液加温仪用管路 | 通常由管路、导管、接头、保护帽等组成。在CRRT治疗中配合加温仪使用，即时加温血液。 | Ⅲ
482 | 10-04-01 | 血液透析浓缩物 | 通常由 A 剂和 B 剂组成。Ａ剂通常由氯化钠、氯化钾、氯化钙、氯化镁、醋酸钠/醋酸/冰醋酸组成,B剂由碳酸氢钠组成。新成分、新配方（供应状态、混合比例、最终溶质浓度）、新功能除外。 | Ⅲ | 新增 | 新
483 | 10-04-03 | 透析液滤过器 | 通常由中空纤维、密封剂、外壳、外壳盖和垫圈组成。利用空心纤维膜的作用，用于清除透析液中的内毒素、细菌与不溶性微粒。 | Ⅲ
484 | 10-04-03 | 透析机消毒液 | 通常由柠檬酸或冰醋酸和水等组成。用于透析机的清洗和消毒。原料建议符合药典要求。豁免情况不包括使用新配方、新作用机理、新功能的产品。 | Ⅲ
485 | 10-04-03 | 一次性使用补液管路/置换液管 | 通常由管路、接头、保护套和夹具等组成。无菌提供，一次性使用。用于血液透析滤过、血液滤过时作为补充置换液的管路。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
486 | 10-04-04 | 一次性使用腹膜透析管探针 | 一般采用不锈钢材料制成。无菌提供，一次性使用。用于促进导入急性和慢性腹膜透析导管。 | Ⅱ
487 | 10-04-04 | 一次性使用腹膜透析接头 | 通常由连接头和螺旋锁盖组成。无菌提供，一次性使用。用于腹膜透析导管与外接延长管或腹膜透析外接短管的连接。 | Ⅱ
488 | 10-04-04 | 腹膜透析外接短管 | 通常由尖端保护帽、开关、套筒、管路、腹透管连接端口、拉环帽和腹透液连接端口组成。无菌提供，一次性使用。用于与腹膜透析患者端管路（或者钛接头）以及腹膜透析液端管路进行无菌连接及分离。 | Ⅱ
489 | 10-04-04 | 腹膜透析管 | 通常由管路、连接端口、保护帽等组成。一般采用高分子材料制成。无菌提供，一次性使用。用于对肾功能衰竭患者进行腹膜透析建立治疗通路。 | Ⅱ
490 | 10-04-04 | 碘液保护帽 | 通常由微型盖、浸润聚维酮碘溶液的海绵、外包装等部件组成。无菌提供，一次性使用。用于保护腹透液袋的外凸接口与外接管路的连接处。 | Ⅱ
491 | 10-04-04 | 一次性使用腹膜透析引流器 | 通常由防护帽、连接接口、三通灌注管、引流管和废液收集袋组成。无菌提供，一次性使用。 | Ⅱ
492 | 10-04-04 | 一次性使用腹膜透析机管路 | 通常由卡匣、管组架、接头装配集合、管路、浇铸端口、夹子（例如O型夹）、Y型连接器、拉环末端保护帽、内拉环帽、拉环帽等组成。无菌提供，一次性使用。 | Ⅱ
493 | 10-05-02 | 体外循环连续血气监测系统 | 通常由主机、静脉探头、动脉探头、传感器、样本池等组成。通过各传感器把测量的数据传输到血气监测系统主机，以实现对体外循环过程中各参数的监测。应提供血气监测准确性的临床验证资料。 | Ⅲ
494 | 10-05-03 | 心肺转流设备-热交换设备 | 通常由水循环系统、制冷系统、加温系统组成，与心肺流转设备联合使用，用于心脏外科手术体外循环的温度调节。 | Ⅱ
495 | 10-05-05 | 加温仪 | 通常由控制装置、加热装置、监测装置组成，用于输血、输液、血液透析和体外循环中血路或液路的即时加温，保持患者体温。 | Ⅱ
496 | 10-06-02 | 静脉贮血器 | 通常由上盖、血库、滤网、出血嘴等组成。无菌提供，一次性使用。在临床心脏直视手术体外循环时，配合膜式氧合器使用，用于贮存、祛泡、过滤体外循环血液。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ |  | 新
497 | 10-06-03 | 心肺转流系统血液过滤器 | 通常由整体部件外壳、滤芯、底座、三通接头、应急旁路管、配用部件三通开关、测压管组件、排气管等组成。无菌提供，一次性使用。产品参考强制性行业标准YY 0580规定。用于体外循环心脏直视手术中，清除血液中的微小固体及气体栓子。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
498 | 10-06-05 | 一次性使用心脏停跳液灌注器 | 通常由气泡捕获器、变温器、连接管和接头组成。无菌提供，一次性使用。用于在体外循环心脏直视手术中，将心肌停跳液灌入冠状动脉，进行心肌的保护。产品参考强制性行业标准YY 0485。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
499 | 10-06-06 | 一次性使用心脏停跳液灌注管 | 灌注管通常由管身、接头、不锈钢针芯（可不含）、管夹等部件构成。无菌提供，一次性使用。适用于体外循环手术中向心脏灌注心脏停跳液，或者连接患者动脉与血液停跳液灌注装置使用。 | Ⅲ
500 | 10-06-06 | 动脉插管、静脉插管 | 通常由吸引头、手柄、接头、插管等部件构成。无菌提供，一次性使用。配套心肺转流系统，在体外循环施行心脏直视手术时引流或灌注血液时使用。产品参考强制性行业标准YY 0948规定。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
501 | 10-06-06 | 左心引流管、右心吸引管 | 由吸引头、管体、接头三部分结构组成。产品参考YY 1271规定。供心血管手术中用于左心脏排气，吸引减压或减轻左心负荷，吸引心脏术野内血液等液体使用。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
502 | 10-06-06 | 心肺转流用管路 体外循环管路 | 通常由管路、接头、泵管等部件构成,符合YY 1048的规定。无菌提供，一次性使用。用于在体外循环或不完全心肺功能支持手术中，提供血液转流通路。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
503 | 10-06-07 | 离心泵泵头 | 通常由外壳、叶片、轴、轴衬和磁铁组成。无菌提供，一次性使用。配合离心泵使用，用于在心脏手术中的体外循环或循环辅助。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅲ
504 | 13-01-01 | 金属锁定接骨板 | 由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11等标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（例如钝化、电解抛光、阳极氧化等）制成，不包含全新结构设计的产品，不包括3D打印等创新工艺。适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。 | 新
505 | 13-01-01 | 金属锁定接骨螺钉 | 产品由符合GB 4234.1、ISO 5832-1标准规定的不锈钢和符合GB/T 13810、ISO 5832-2、ISO 5832-3、ISO 5832-11标准规定的纯钛、Ti6Al4V、Ti6Al4VELI、Ti6Al7Nb钛合金材料，通过常规的机械加工、热处理及表面处理工艺（例如钝化、电解抛光、阳极氧化等）制成，不包含全新结构设计的产品，不包括3D打印等创新工艺。适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。 | 新
506 | 13-01-01 | 金属接骨螺钉(非锁定) | 该类产品结构参见 YY 0018，产品由符合GB 4234.1、ISO 5832-1标准规定的不锈钢和符合GB/T 13810、ISO 5832-2、ISO 5832-3、ISO 5832-11标准规定的纯钛、Ti6Al4V、Ti6Al4VELI、Ti6Al7Nb钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。可单独使用，适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。
507 | 13-01-01 | 空心接骨螺钉(非锁定) | 该类产品由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。可单独使用，适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。
508 | 13-01-01 | 解剖型金属接骨板(非锁定) | 该类产品结构参见YY 0017，由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢干骺端骨折内固定。 | Ⅲ | 修订产品描述。
509 | 13-01-01 | 直型金属接骨板（非锁定） | 该类产品结构参见YY 0017，由符合GB 4234.1、GB/T 13810、 ISO 5832-1、ISO 5832-2、ISO 5832-3标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢长骨骨干骨折内固定。 | Ⅲ | 修订产品描述。
510 | 13-01-01 | 角度型金属接骨板（非锁定） | 该类产品结构参见YY 0017，由符合GB 4234.1、GB/T 13810、 ISO 5832-1、ISO 5832-2、ISO 5832-3标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢长骨骨干骨折内固定。 | Ⅲ | 修订产品描述。
511 | 13-01-01 | U型钉 | 该类产品结构参见ISO 882，由符合GB 4234.1、GB/T 13810、 ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。
512 | 13-01-01 | 骨栓 | 该类产品由螺杆和螺母组成，由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢骨折内固定。 | Ⅲ | 修订产品描述。
513 | 13-01-01 | 金属股骨颈固定钉 | 该类产品结构参见YY 0346，由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于股骨颈骨折内固定。 | Ⅲ | 修订产品描述。
514 | 13-01-01 | 外固定架配合用固定钉 | 该类产品由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。与外固定支架配合，适用于四肢骨折复位时部分植入人体做固定。 | Ⅲ | 修订产品描述。
515 | 13-01-01 | 髋臼螺钉 | 该类产品由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11、ISO 5832-12标准规定的不锈钢、纯钛、钛合金、锻造钴铬钼材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于生物型髋臼假体内固定。 | Ⅲ | 修订产品描述。
516 | 13-01-01 | 骨缺损填充块固定螺钉 | 该类产品由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11、ISO 5832-12标准规定的不锈钢、纯钛、钛合金、锻造钴铬钼材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于髋关节、膝关节金属骨缺损填充块的固定。 | Ⅲ | 修订产品描述。
517 | 13-01-01 | 足踝锁定接骨板系统 | 该类产品由锁定接骨板和锁定接骨螺钉组成，锁定接骨板由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11标准规定的不锈钢、纯钛、钛合金材料，锁定接骨螺钉由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3、ISO 5832-11标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于足踝骨折内固定。 | Ⅲ | 修订产品描述。
518 | 13-01-01 | 手腕锁定接骨板系统 | 该类产品由锁定接骨板和锁定接骨螺钉组成。锁定接骨板由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11标准规定的不锈钢、纯钛、钛合金材料，锁定接骨螺钉由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3、ISO 5832-11标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于手腕部骨折内固定 | Ⅲ | 修订产品描述。
519 | 13-01-01 | 肋骨接骨板系统 | 该类产品由肋骨接骨板和螺钉组成。由符合GB/T 13810、ISO 5832-2、ISO 5832-3、ISO 5832-11标准规定的纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于肋骨骨折内固定。 | Ⅲ
520 | 13-01-04 | 金属带锁髓内钉 | 由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3、ISO 5832-11等标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺制成（不包括3D打印等创新工艺）。通常由主钉和/或附件组成，钉孔内可含由符合GB/T 19701标准要求的超高分子量聚乙烯或符合YY/T 0660标准要求的PEEK等在该类产品有成熟临床应用史材料制成的配件。适用于四肢骨折髓腔内固定。 | Ⅲ | 修订产品描述。
521 | 13-01-04 | 金属髓内针 | 该类产品结构参见YY 0019，不包含带锁髓内钉。由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工工艺、热处理及表面处理工艺（不包括3D打印等新技术和新工艺）制成；适用于四肢骨折内固定。植入骨髓腔内由螺钉等连接固定 | Ⅲ | 修订产品描述。 | 新
522 | 13-01-05 | 柔性金属丝 | 该类产品结构参见YY/T 0816，由符合YY 0605.9、ISO 5832-1、ISO 5832-2标准规定的不锈钢、纯钛、高氮不锈钢材料，通过常规的机械加工工艺、热处理工艺（不包括3D打印等新技术和新工艺）制成。适用于骨固定或骨接合。 | Ⅲ
523 | 13-01-05 | 金属缆线和缆索 | 该类产品结构参见YY/T 0812，可含有与其配合使用的附件（如锁定针、束缚器、锁扣等）。由符合GB/T 13810、YY 0605.9、ISO 5832-1、ISO 5832-2标准规定的不锈钢、纯钛、高氮不锈钢、钛合金材料，符合ASTMF1314标准规定的22Cr-13Ni-5Mn不锈钢材料、符合ISO5832-5/YY 0605.5标准规定的锻造钴铬钨镍合金材料，通过常规的机械加工、热处理工艺（不包括3D打印等创新工艺）制成。适用于四肢骨折捆扎内固定。 | Ⅲ | 修订产品描述。
524 | 13-01-06 | 金属骨针 | 该类产品包括克氏针和斯氏针，结构参见YY 0345。由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，符合ISO 5832-5标准规定的锻造钴铬钨镍材料通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。适用于四肢骨折复位时部分植入人体做牵引、配合外固定支架进行四肢骨折固定或单独植入用于四肢骨折内固定。 | Ⅲ | 修订产品描述。
525 | 13-02-01 | 不可吸收带线锚钉 | 该类产品由锚钉和缝线组成，可带缝合针。锚钉由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3、ISO 5832-11、YY/T 0660标准规定的不锈钢、钛合金、聚醚醚酮材料，通过常规的机械加工、注塑、热处理及表面处理工艺（不包括3D打印等创新工艺）制成。缝线可由聚乙烯、聚丙烯、聚酯、聚酰胺等不可吸收材料制成，性能指标参考YY0167标准规定的适用部分。缝合针参考YY/T 0043标准规定的适用部分。该产品可预装插入器，插入器与人体接触部分通常由符合YY/T 0294、ASTM F899或YY/T 0726规定的不锈钢材料制成。适用于骨与软组织的连接固定。 | Ⅲ | 修订产品描述。
526 | 13-02-01 | 不可吸收韧带固定螺钉 | 该类产品由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-3标准规定的不锈钢、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）或符合YY/T 0660的聚醚醚酮（PEEK）材料制成。适用于韧带的修复和重建。 | Ⅲ | 修订产品描述。
527 | 13-02-01 | 带袢钛板 | 产品由钛板、线圈和/或缝线组成。钛板应由符合GB/T 13810、ISO 5832-3标准规定的Ti6Al4V、Ti6Al4VELI钛合金制成，线圈和缝线材质为高分子材料，如聚对苯二甲酸乙二酯，聚丙烯、聚乙烯，聚酯材料。适用于骨科重建术中韧带或肌腱与骨的固定。 | Ⅲ
528 | 13-02-01 | 半月板缝合系统 | 半月板缝合系统一般由植入物和辅助工具组成，植入物包括固定锚（棒）和缝线，辅助工具包括插入器等，不植入体内。固定锚（棒）采用符合YY/T 0660标准要求的聚醚醚酮(PEEK)材料制成，缝线是由超高分子量聚乙烯（或为主）材料制成，性能指标符合 YY0167 标准规定的适用部分。插入器由手柄和不锈钢材料的针杆制成，插入器与人体接触部分通常由符合 YY/T 0294、ASTM F899 或 YY/T 0726 规定的不锈钢材料制成。适用于内窥镜下进行半月板缝合手术。 | Ⅲ | 新增 | 新
529 | 13-03-01 | 椎板固定板系统 | 该类产品由固定板和螺钉组成，固定板由符合GB/T 13810 、ISO 5832-2 、ISO 5832-3标准规定的纯钛、钛合金材料，螺钉由符合GB/T 13810 、ISO 5832-3标准规定的钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D 打印等创新工艺）制成。适用于脊柱椎管扩大减压术后的椎板成形。 | Ⅲ
530 | 13-05-04 | 填充棒/填充块 | 产品由符合ISO 13782标准规定的纯钽制成（不包括3D打印等创新工艺制备产品）。产品单独使用，用于四肢骨填充。 | Ⅲ
531 | 13-06-01 | 单/多部件预制颅骨成形术板及紧固件 | 通常由板及紧固件组成，一般采用纯钛、钛合金、钴铬合金等金属材料以及不可吸收高分子材料制成，其中板通过紧固装置（如螺钉、金属丝或其他组件）固定就位。适用于颅骨缺损的修补。 | Ⅲ | 新增
532 | 13-06-03 | 头钉 | 产品通常由主体、基座、O形环、橡胶帽组件组合而成，其中仅主体与人体接触。主体应使用GB 4234.1、ISO 5832-1、ASTM F899标准规定的不锈钢或ISO 5832-3、GB/T 13810标准规定的Ti6Al4V、Ti6Al4VELI钛合金材料制成，产品与颅骨固定器配合使用，在头骨需要重度修复的外科手术过程中将患者的头部与颈部固定在特定的位置。该产品为非植入物，仅在术中使用，与人体短期接触，术后即刻拆除。 | Ⅱ | 修订产品描述。
533 | 13-06-03 | 颅颌面接骨板系统 | 由符合GB 4234.1、GB/T 13810、ISO 5832-1、ISO 5832-2、ISO 5832-3标准规定的不锈钢、纯钛、钛合金材料，通过常规的机械加工、热处理及表面处理工艺（不包括3D打印等创新工艺）制成，不包含全新结构设计的产品。适用于颅颌面骨折内固定。修补、覆盖、填充骨缺损或孔洞者除外。 | Ⅲ |  | 新
534 | 13-06-05 | 动脉瘤夹 | 一般采用钛合金等金属材料制成。用于使颅内动脉瘤颈或囊闭塞的装置。用于颅内动脉瘤瘤颈或瘤体的永久性闭塞治疗，也可用于动静脉畸形小血管或动静脉瘘的永久性闭塞治疗。产品与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
535 | 13-09-04 | 外科疝修补补片 | 腹腔外置疝补片：植入体内的补片（如平片、网塞等），由聚丙烯、聚对苯二甲酸乙二醇酯或聚偏二氟乙烯材料制成，可带有部分可吸收材料，用于腹腔外对疝与腹壁缺损进行修补；腹腔内置疝修补补片：腹壁侧由聚丙烯、聚对苯二甲酸乙二醇酯或聚偏二氟乙烯材料制成，腹腔侧由可降低粘连的材料制成，用于腹腔内对疝与腹壁缺损进行修补。上述两类补片与已获准境内注册产品具有等同性，组成材料成熟,可降低粘连的材料在上市同类产品中已应用。新型结构设计、新型作用机理、药械组合产品除外。不包括最终全部吸收的产品及以自体组织重建实现修复的产品。 | Ⅲ
536 | 13-09-06 | 胆道/胰管/输尿管支架 | 支架由不可降解/吸收的高分子材料（如聚乙烯、聚氨酯等）制成管状结构，可带有不透射线标记等组件，可涂有涂层，可包含输送器组件。用于在内窥镜下植入腔道，扩张植入部位的狭窄，进行引流。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
537 | 13-09-06 | 非血管自扩张金属支架 | 支架由具有形状记忆性能的金属（如镍钛合金等）制成，可带有不透射线标记，可有/无覆膜，可包含输送器组件。用于扩张食道、胆道、气道、肠道等非血管部位的狭窄。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
538 | 13-09-10 | 组织扩张器 | 一般由壳体、导管、注射座（或起相同作用的组件）和连接器（如果有）等组成。壳体的主要材料为硅橡胶。作用原理为可填充盐水以增大自身体积对组织进行扩张。用于获取自体皮肤组织以解决皮肤供区不足，手术中使用或植入6个月内取出。与已获准境内注册产品具有等同性。新型结构设计、药械组合产品除外。 | Ⅲ
539 | 13-04-01 | 中置器 | 该类产品由PMMA、符合GB/T19701.2标准规定的超高分子量聚乙烯材料制造。适用于骨水泥型节髋关节置换手术中保持股骨柄远端中置，阻隔骨水泥并确保骨水泥在假体周围的均匀分布。 | Ⅲ | 修订分类编码。
540 | 13-04-01 | 远端塞 | 该类产品由PMMA、符合GB/T19701.2标准规定的超高分子量聚乙烯材料制造。适用于骨水泥型髋关节置换术中，阻隔骨水泥。 | Ⅲ | 修订分类编码。
541 | 13-04-01 | 髋臼杯用孔塞 | 产品由符合GB/T 13810、ISO 5832-3标准Ti6Al4V钛合金制成。配合髋臼杯使用，用于髋臼杯产品孔的填塞。 | Ⅲ | 修订分类编码。
542 | 13-04-01 | 髓腔塞 | 产品由符合GB/T 19701.2、ISO 5834-2标准规定的超高分子量聚乙烯材料制造。用于骨水泥型假体植入手术，作用为防止骨水泥溢入髓腔限制区域，适用部位为股骨、胫骨和肱骨的髓腔。 | Ⅲ | 修订分类编码。
13-04-02
13-04-03
13-04-04
543 | 13-11-03 | 胸骨结扎带 | 产品由符合GB/T 13810、ISO 5832-3、ISO 5832-11、ISO 5832-1标准规定的Ti6Al4V、Ti6Al4VELI、Ti6Al7Nb钛合金材料、不锈钢材料制成，穿引弯钩也可由符合ASTM F899标准规定的材料制成，产品通常由结扎带和穿引弯钩两部分组成，可含锁扣，其中穿引弯钩非植入人体，仅用于辅助结扎带穿引安装，安装后废弃。用于开胸术中的胸骨固定。不包含全新结构设计的产品。 | Ⅲ | 修订产品描述。 | 新
544 | 14 | 医用口罩（非外科用） | 医用口罩（非外科用）通常用无纺布或无纺布复合材料制成，可为二层或三层结构，可有可塑性鼻夹，口罩带可为弹性或非弹性，具有过滤颗粒物和细菌等特性，供临床各类人员在非有创操作过程中佩带，覆盖住使用者的口、鼻及下颌，为防止病原体微生物、颗粒物等的直接透过提供一定的物理屏障。豁免情况不包括具有抗病毒等新功能的产品。 | Ⅱ | 修订产品描述。
545 | 14 | 自锁注射器 | 一般由外套、圆锥接头、芯杆、卡圈、活塞、密封“O”型圈组成，预期用于介入治疗、造影手术或放射性手术前检查球囊或抽吸用。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
546 | 14-01 | 一次性使用生物蛋白胶配制器 | 本产品可由配药器、双腔推液器、支架、三通、喷头及喷针等部件组成。本产品经灭菌，一次性使用。用于配制、混合蛋白胶组份后喷至人体创口表面。 | Ⅱ
547 | 14-01-01 | 注射泵 | 注射泵产品是通过一个或多个单一动作的注射器或类似容器来控制注入患者体内液体流量的设备（例如通过推动推杆清空筒内溶液），输液速度由操作者设定，并由设备指示单位时间内的流量；一般由泵外壳、电机驱动系统、输入系统、存储系统、控制系统、显示系统、传感监测系统和报警系统组成，产品结构组成中一般不包含注射器和输注管路，可按设计、技术参数、预期用途、附加辅助功能等不同分为若干型号；与特定注输器具配套，用于动、静脉长时间恒定给药速度和精确给药量输液。产品性能指标采用下列参考标准中的适用部分，如： GB 9706.27-2005医用电气设备 第2-24部分：输液泵和输液控制器安全专用要求。 | Ⅱ
548 | 14-01-02 | 一次性使用无菌注射器 | 性能、结构、组成、用途等参考GB 15810《一次性使用无菌注射器》可完全涵盖的产品，一般由注射器外套、活塞、密封圈、芯杆、按手、锥头组成，是供抽吸液体或在注入液体后立即注射用的手动注射器。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
549 | 14-01-02 | 一次性使用无菌注射器 带针 | 性能、结构、组成、用途等参考GB 15810一次性使用无菌注射器、GB 15811一次性使用无菌注射针可完全涵盖的产品，为已发布豁免目录中一次性使用无菌注射器配合已发布豁免目录中一次性使用无菌注射针。一般由注射器外套、活塞、密封圈、芯杆、按手、锥头针座、连结部、针管、护套组成，用于人体皮内、皮下、肌肉、静脉注射或抽取药液。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
550 | 14-01-02 | 一次性使用低阻力注射器 带针 | 性能、结构、组成、用途等参考YY/T0909《一次性使用低阻力注射器》可完全涵盖的产品。产品一般由外套、活塞帽、活塞、芯杆及注射针（可不带）组成，无菌供应，供抽吸液体或在注入液体后立即注射用。 | Ⅲ
豁免情况不包括新材料、新作用机理、新功能的产品。
551 | 14-01-02 | 一次性使用无菌胰岛素注射器 | 性能、结构、组成、用途等参考YY 0497《一次性使用无菌胰岛素注射器》可完全涵盖的产品，主要由外套、芯杆、橡胶活塞、注射针管、护套/端帽组成。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
552 | 14-01-02 | 一次性泵用注射器 | 一般由注射器外套、活塞、密封圈、芯杆、按手、锥头组成，可带针或不带注射针。供抽吸液体或在注入液体后立即注射，与注射泵配合使用。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
553 | 14-01-04 | 笔式胰岛素注射器 | 笔式胰岛素注射器由剂量调节钮、笔身 、释放按钮、引导螺杆、笔芯架、笔帽组成，不包括药物和注射针头，所有部件不与药液接触；须与特定的胰岛素和注射针或标准注射针配套，供胰岛素皮下注射用。 | Ⅱ
554 | 14-01-04 | 笔式注射器 | 通常由笔帽、笔芯架、螺旋杆、笔身、剂量调节栓和注射按钮组成；有源笔式注射器通常还包含具有辅助功能的其它电子组件。与笔芯和/或针配合使用，通过压力使药液穿透皮肤或黏膜表面，输送入体内。用于药液（如胰岛素）的注射。 | Ⅱ
555 | 14-01-05 | 全玻璃注射器 | 全玻璃注射器采用硅硼铝玻璃材料制成，由锥头、外套和芯子组成，可分为中头式和偏头式两种，每种按公称容量不同分为若干规格；装上注射针头，供人体进行皮下、肌肉、静脉注射药液或抽取液体等用。也可经清洗灭菌后可重复使用。产品性能指标采用下列参考标准中的适用部分，如：YY 1001.1《玻璃注射器 第1部分：全玻璃注射器》。 | Ⅱ
556 | 14-01-05 | 蓝芯全玻璃注射器 | 蓝芯全玻璃注射器采用硅硼铝玻璃材料制成，由锥头、外套和芯子（蓝色实心玻璃）组成，通常为中头式，按公称容量不同分为若干规格；装上注射针头，可供人体进行生化试验、皮下试验。注射疫苗、口腔麻醉用，也可供注射其他药液用。产品性能指标采用下列参考标准中的适用部分，如：YY 1001.2《玻璃注射器 第2部分：蓝芯全玻璃注射器》。 | Ⅱ
557 | 14-01-06 | 金属注射针 | 金属注射针通常用铜基镀铬（针座）和高等级不锈钢材料（针管）制成，一般由针座和针管构成，与注射器配套，供人体用于皮下、皮内、肌肉、口腔等部位注射药液、疫苗、麻醉剂或静脉输液、输血用，可重复使用。产品性能指标采用下列参考标准中的适用部分，如： YY/T 0282《注射针》。 | Ⅱ
558 | 14-01-06 | 一次性使用无菌注射针 | 性能、结构、组成、用途等参考GB 15811《一次性使用无菌注射针》可完全涵盖的产品，由针座、连结部、针管、护套组成，用于人体皮内、皮下、肌肉、静脉注射或抽取药液。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
559 | 14-01-06 | 一次性使用注射笔用针头 | 由密封贴膜、针管、针座、外护套、内护套组成。产品无菌、无热原。与注射笔配合使用向人体皮下注射药物。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
560 | 14-01-06 | 注射针（Ⅱ） | 通常由针管、针座和护套组成。针管和针座一般采用不锈钢材料制成。非无菌提供，可重复使用。用于人体皮内、皮下、消化道黏膜下、肌肉注射或抽取液体。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅱ
561 | 14-01-06 | 注射针（Ⅲ） | 通常由针管、针座和护套组成，可带有自毁装置。针管一般采用不锈钢材料制成，针座一般采用高分子材料制成。无菌提供。用于人体皮内、皮下、消化道黏膜下、肌肉、静脉等注射或抽取液体。豁免情况不包括新材料、新作用机理、新结构、新功能的产品。 | Ⅲ
562 | 14-01-07 | 注射器辅助推动装置 | 一般采用金属材料、高分子材料等制成。不接触注射药液。配合注射器等使用，用于对注射器进行辅助推注。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅱ
563 | 14-01-08 | 穿刺器械 | 通常由穿刺针、穿刺器、保护套组成。 | Ⅱ
564 | 14-01-08 | 金属胸腔穿刺针 | 金属胸腔穿刺针可用铜基镀铬（针座等）和高等级不锈钢材料（针管）制成，可由针座（有标准鲁尔接头）、三通、定位装置和有刻度针管构成，通常针管的公称外径为1.2mm或1.6mm，针管长度可为65mm—65mm或更长；与注射器配套，供胸腔穿刺，建立通道，抽取胸腔积液等用，可重复使用。 | Ⅱ
565 | 14-01-08 | 环甲膜穿刺针 | 环甲膜穿刺针可采用不锈钢等金属材料制成，由带止位结构的套管和匹配的锥形穿刺针组成,可按材质、设计、技术参数等不同分为若干型号及规格；供环甲膜穿刺，建立人工气道，缓解上呼吸道阻塞用。 | Ⅱ
566 | 14-01-08 | 一次性使用腹腔穿刺针 | 组成仅限于由针管、针座、连接管组成腹腔穿刺针，用于对人体作腹腔穿刺用。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
567 | 14-01-08 | 穿刺针、穿刺器（用于腰椎、血管、脑室穿刺） | 通常由穿刺针、穿刺器、保护套组成。用于对腰椎、血管、脑室进行穿刺，以采集人体样本、注射药物与气体等或作为其他器械进入体内的通道。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
568 | 14-01-08 | 金属腰椎穿刺针 | 金属腰椎穿刺针可用铜基镀铬（针座、衬芯座）和高等级不锈钢材料（针管、衬芯）制成，可由衬芯座、定位销、针座（有标准鲁尔接头）、衬芯、针管构成，通常针管的公称外径为0.4mm—1.6mm，针管长度为25mm—200mm,针管管壁可分为正常壁和薄壁两种；与注射器配套，供人体蛛网膜下腔阻滞（腰椎麻醉）穿刺，注射药液或抽取脑脊液用，可重复使用。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1148－2009 腰椎穿刺针。 | Ⅲ
569 | 14-01-08 | 硬膜外腔穿刺指示球囊 | 通常由腔体、单向阀、指示膜、O型环等组成。在硬膜外麻醉穿刺进程中为穿刺针针尖抵达硬膜外腔提供一个明确的、可视化的信号指示。 | Ⅱ
570 | 14-01-09 | 活检针 | 通常由针座、芯针、内针管、内外针定位鞘、外针管、保护套、组织采集装置组成。针管一般采用不锈钢材料制成。 | Ⅱ
571 | 14-01-09 | 一次性使用活检针 | 组成仅限于包括穿刺针和针芯的活检针，材料符合GB 18457的要求，通过直接进行人体组织器官穿刺以获取病理组织，不包括与内窥镜系统、负压吸引装置等配合使用或需超声、X线等影像设备引导下操作的活检针。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
572 | 14-01-09 | 活检针 | 由针管、针芯、手柄等部件组成。用于获取患者骨髓活检，或在B超或X-线监视下，经皮穿刺进行实质性脏器或肿瘤的细胞学活检或其他软组织活检。一次性使用。产品基本原理、适用范围、性能和组成结构与已经上市产品相同。 | Ⅱ
豁免情况不包括：使用新型材料；包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等产品。
573 | 14-01-09 | 同轴活检针 | 由带针座的不锈钢穿刺针（即外套管）和带针座的针芯，以及塑料环（用于套在不锈钢上以标记穿刺深度）组成，获取软组织活检标本，不用于骨活检。产品基本原理、适用范围、性能和组成结构与已经上市产品相同。 | Ⅱ
豁免情况不包括：使用新型材料；包含高分子、药物、生物制品等特殊涂层的产品，以及具有特殊结构功能、适用范围等产品。
574 | 14-01-10 | 活检枪 | 通常由弹射、释放、制动部分构成。非无菌提供，可重复使用。与活检针装配好后用于从人体组织获取标本进行活检。 | Ⅱ
575 | 14-02-01 | 输液泵 | 输液泵主要由控制系统、电机驱动单元、蠕动挤压机构、检测装置、报警装置、输入及显示装置、壳体及其支撑结构、软件组件组成，可按工作原理、设计、技术参数、附加辅助功能等不同分为若干型号；与输液器配套，不接触输注液体，供医院以可调节的方式为患者静脉输注药液等使用。产品性能指标采用下列参考标准中的适用部分，如： GB 9706.27-2005 医用电气设备 第2-24部分：输液泵和输液控制器安全专用要求。 | Ⅱ
576 | 14-02-00 | 便携式电动输液泵用储液器 | 通常由储液器、管路、截流夹、内圆锥锁定接头及保护帽等组成。不含药液。无菌提供，一次性使用。用于与特定的镇痛泵、化疗泵等便携式电动输液泵配合使用，向患者输注药液。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订分类编码、产品名称、产品描述和产品管理类别描述。
577 | 14-02-02 | 输液信息采集系统 | 输液信息采集系统由带有通讯接口和电源的移动支架、集成软件组成；为输液泵/注射泵供电，并通过通讯接口与指定型号的输液泵/注射泵进行数据通讯；采集输液泵/注射泵数据，通过有线/无线网络传输到中央工作站，并提示报警信息。 | Ⅱ/Ⅲ | 修订产品管理类别描述。
578 | 14-02-03 | 输液辅助电子设备 | 通常与输液器配合使用，能使输液过程实现流量控制、加温、报警等功能的电子仪器。 | Ⅱ
579 | 14-02-04 | 一次性使用输注泵 | 通常由弹力储药囊（不含药）、加药装置、延长管和流速控制器组成。以机械弹性为动能，为泵体提供动力。无菌提供，一次性使用。用于患者自控调节注入体内（静脉、皮下、硬膜外腔）的药液流量。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ | 新增 | 新
580 | 14-02-05 | 一次性使用避光输液器 带针/一次性使用避光输液器 | 性能、结构、组成、用途等参考GB 8368 《一次性使用输液器重力输液式》、YY 0286.3-2017 《专用输液器 第3部分：一次性使用避光输液器》、GB 18671《一次性使用静脉输液针》可完全涵盖的传统型一次性使用输液器，与输液容器、输液针配用，用于静脉输液。可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
581 | 14-02-05 | 传统型一次性使用输液器 | 性能、结构、组成、用途等参考GB 8368《一次性使用输液器 重力输液式》可完全涵盖的传统型一次性使用输液器，由瓶塞穿刺器及保护套、带空气过滤器和塞子的进气口、液体通道、滴管、滴斗、药液过滤器、管路、流量调节器、注射件、外圆锥接头及保护套组成，与输液容器、输液针配用，用于静脉输液。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
582 | 14-02-05 | 传统型一次性使用输液器 带针 | 性能、结构、组成、用途等符合GB 8368 《一次性使用输液器重力输液式》、GB 18671《一次性使用静脉输液针》的传统型一次性使用输液器，与输液容器、输液针配用，用于静脉输液。可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
583 | 14-02-05 | 一次性使用精密过滤输液器 带针 | 性能、结构、组成、用途等符合YY0286.1《专用输液器 第1部分：一次性使用微孔过滤输液器》、GB 18671《一次性使用静脉输液针》的传统型一次性使用输液器，与输液容器、输液针配用，用于静脉输液。可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
584 | 14-02-05 | 一次性使用吊瓶式和袋式输液器 带针 | 性能、结构、组成、用途等符合YY0286.5《专用输液器 第5部分：一次性使用吊瓶式和袋式输液器》、GB 18671《一次性使用静脉输液针》的传统型一次性使用输液器，用于静脉输液。可带输液贴（位于单包装内），瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
585 | 14-02-05 | 一次性使用滴定管式输液器 带针 | 性能、结构、组成、用途等符合YY0286.2《专用输液器 第2部分：一次性使用滴定管式输液器 重力输液式》、GB 18671《一次性使用静脉输液针》的传统型一次性使用输液器，用于静脉输液。可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
586 | 14-02-05 | 一次性使用流量设定微调式输液器 带针 | 性能、结构、组成、用途等符合YY0286.6《专用输液器 第6部分：专用输液器 第6部分：一次性使用流量设定微调式输液器》、GB 18671《一次性使用静脉输液针》的传统型一次性使用输液器，可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
587 | 14-02-05 | 一次性使用压力输液器 | 性能、结构、组成、用途等符合YY/T 0286.4《专用输液器 第4部分：一次性使用压力输液设备用输液器》的传统型一次性使用输液器，与输液容器、输液针配用，用于静脉输液，可带输液贴，瓶塞穿刺器可为双头或多头。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
588 | 14-02-05 | 输液器 | 通常由鲁尔圆锥接头、管路、滴斗、流量调节器、瓶塞穿刺器、药液过滤器等组成，部分输液器带有空气过滤器的进气气件、药液注射件。管路一般由聚氯乙烯或其他材料制成。其设计能使其在重力或压力的作用下，将输液容器中的药液通过静脉穿刺器械向静脉内输液。无菌提供，一次性使用。用于静脉输注药液。豁免情况不包括新材料、新作用机理、新结构、新功能的产品。 | Ⅲ
589 | 14-02-05 | 一次性泵用输液器 | 通常由鲁尔圆锥接头、管路、滴斗、流量调节器、瓶塞穿刺器、药液过滤器等组成。部分输液器带有空气过滤器的进气气件、药液注射件。管路一般由聚氯乙烯或其他材料制成。与输液泵配合使用，将输液容器中的药液通过静脉穿刺器械向静脉内输液。无菌提供，用于静脉输注药液。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
590 | 14-02-06 | 一次性使用静脉输液针 | 性能、结构、组成、用途等符合GB 18671《一次性使用静脉输液针》的产品，一般由针管、针柄、软管、连接座和保护帽组成，与输液器、输血器配套使用，用于建立外周静脉通路。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
591 | 14-02-06 | 一次性使用植入式给药装置专用针 | 性能、结构、组成、用途等符合YY 0881《一次性使用植入式给药装置专用针》的产品，植入式给药装置专用针分为注射用针和输液用针两种型式，与植入式给药装置和输注装置配合使用。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
592 | 14-02-07 | 一次性使用静脉留置针 | 性能、结构、组成、用途等符合YY/T 1282《一次性使用静脉留置针》的产品。一般由导管组件和针管组件两部分组成，导管组件是由导管、导管座、延长管（如果有）和任何一体的接头组成的装配件，针管组件是由针管和针座和/或其他一体的附件组成的装配件。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
593 | 14-02-07 | 动脉留置针 | 动脉留置针由流量控制阀、针管把手、导管座、导管、保护套、硅胶管、排气接头、针管、不锈钢球、导引套管及导管套管组成。以无菌形式提供；产品用于动脉压监测及连续动脉血气的监测。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
594 | 14-02-07 | 一次性使用胸腔穿刺针 | 组成仅限于由针管、针座、连接管组成胸腔穿刺针，用于对人体作胸腔穿刺用。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
595 | 14-02-08 | 一次性使用压力延长管 | 一次性使用压力延长管主要由内、外圆锥鲁尔接头、延长管和/或三通阀组成，使用时与进入人体血管的导管相连。预期用途为在介入手术中用于联接管路和设备，提供药液的输注通路和/或用于有创血压监测。原材料应符合GB/T 15593等相关标准要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
596 | 14-02-08 | 输液连接管路 | 通常由输液管路、至少一个外圆锥鲁尔接头和一个内圆锥鲁尔接头组成。管路一般由聚氯乙烯或其他材料制成。无菌提供，一次性使用。通过鲁尔接头与其他输液器械连接，实现两个器械间的管路连接后输注药液。豁免情况不包括新材料、新作用机理、新结构、新功能的产品。 | Ⅲ
597 | 14-02-08 | 输液连接管路 | 通常由输液管路、外圆锥鲁尔接头和内圆锥鲁尔接头组成。一次性使用。通常由内、外圆锥鲁尔接头、延长管和/或三通阀组成，使用时与进入人体血管的导管相连或用于联接设备，提供药液的输注通路和/或用于有创血压监测。包括常压和压力下输注。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
06-05-05
598 | 14-02-09 | 输液用连接件及附件 | 无针连接件：产品一般由接口、接头和管路（可选）组成。为非穿刺式的输液连接件，用于与血管内留置导管配合使用，通过它向血管内输注药液或抽取血液用。三通（阀）：一般由外壳、保护套、阀门、旋转手柄、标志塞、螺帽组成。本产品与输液器具连接使用，用于串接在输液系统上使其增加了一项或多项特定的功能，如液路开关、防回流、过滤等。肝素帽：性能、结构、组成、用途参考YY 0581 输液用肝素帽完全涵盖的以穿刺为主要使用形式的产品，一般由外壳、胶帽组成，与外周套针导管等配合使用，通过穿刺可向血管内注入药液。连通板：产品一般由接头、连接主板、开关（阀门）、软管组成，作为连接通路，建立多通道，提供药液或造影剂注射使用。过滤器：通常由过滤膜、硅胶阀瓣、鲁尔座、鲁尔接头及保护帽组成, 用于与温液仪及温液管路配套连接使用，重力式输注时，排除加热液体时产生的微孔气泡，不包括滤除微粒或细菌功能。产品性能指标采用下列标准中的适用部分，如：GB8368《 一次性使用输液器 重力输液式》；YY 0585.3《压力输液设备用一次性使用液路及附件 第3部分：过滤器》”。以上产品的豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
599 | 14-02-09 | 输液用连接件及附件 | 通常至少带有一鲁尔圆锥接头，一般采用高分子材料制成。无菌提供，一次性使用。用于串接在输液系统上使其增加了一项或多项功能，如液路开关、防回流、过滤等。常见如：无针连接件、三通（阀）、肝素帽、连通板、过滤器等。以上产品的豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
600 | 14-02-11 | 静脉营养输液袋、泵用输液袋 | 通常由袋体、药液加入口和输液器插口组成，一般由高分子材料制成。无菌提供，一次性使用。用于医疗机构中，使用前充入营养液或药液，再与输液器和静脉内器械（如中心静脉导管）连接向体内输注。豁免情况不包括新材料、新作用机理、新结构、新功能的产品。 | Ⅲ
601 | 14-02-11 | 一次性使用静脉营养输液袋 | 性能、结构、组成、用途等参考YY 0611《一次性使用静脉营养输液袋》，通常由贮液袋、药液加入口和输液器插口组成，部分输液袋带有输液管路。一般由高分子材料制成，无菌提供，一次性使用。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
602 | 14-02-11 | 一次性泵用输液袋 | 通常由袋体、药液加入口和输液器插口组成，一般由高分子材料制成。用于医疗机构中，使用前充入药液，再与输注泵等连接向体内输注。无菌提供，一次性使用。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
603 | 14-02-12 | 药液用转移、配药器具 | 通常由瓶塞穿刺器、连接管路等组成，可带有过滤器件。当它与两个容器相连后可以实现不同药剂间的转移、溶药、配药或多联输液（多瓶输液）。产品以无菌形式提供。包括：一次性使用配药用注射器：可由护帽、配药针、外套、芯杆和活塞组成，可按公称容量不同分为若干规格，可配侧孔或斜面配药针，供抽取或配制药液用。一次性使用配药针：由针座、针管和护帽组成，可按型式不同分为侧孔和斜面等，与一次性使用配药用注射器配套，供抽取和配制药液用。药液转移器：转移器可以有一个外罩，也可以另接一个空气过滤器或药液过滤器；穿刺器可以是不锈钢，也可以是塑料制，其他组件可由聚氯乙烯等塑料制成。 | Ⅱ
604 | 14-03-01 | 肠内营养泵 | 肠内营养泵主要由主机、电源适配器、内置电池等部分组成，可按工作原理、设计、技术参数、附加辅助功能等不同分为若干型号；与肠内营养泵用输注管路配套使用，不接触输注液体，供医院以可调节的方式为患者胃肠道输注营养液等用，仅适用于肠内营养输注。产品性能指标采用下列参考标准中的适用部分，如：GB 9706.1-2007医用电气设备 第一部分：安全通用要求。 | Ⅱ
605 | 14-03-03 | 一次性使用胰岛素泵用皮下输液器 | 通常由皮下穿刺针、固定装置、管路、连接件等组成。无菌提供，一次性使用。用于与特定的胰岛素泵、胰岛素泵用储液器配合使用，向糖尿病患者持续皮下输注胰岛素。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
606 | 14-03-04 | 一次性使用胰岛素泵用储液器 | 通常由芯杆、活塞、外套等组成。不含胰岛素。无菌提供，一次性使用。用于与特定的胰岛素泵、胰岛素泵用输液器配合使用，向糖尿病患者持续皮下输注胰岛素。豁免情况不包括新材料、新作用机理、新功能的产品。 | Ⅲ
607 | 14-03-05 | 肠营养器 | 通常由瓶塞穿刺器、管路、滴斗和肠营养导管连接接头组成，管路一般由聚氯乙烯或其他高分子材料制成，能与肠营养容器和肠营养管连接。部分肠营养器还带有泵管，使之可在肠营养泵的作用下向肠内输注营养液；部分肠营养器与肠营养袋连为一体用于向胃肠内输送营养液。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
608 | 14-03-06 | 肠营养袋 | 通常由营养制剂加入口、袋体和肠营养器穿刺接口组成。袋体一般由聚氯乙烯或其他高分子材料制造。不含营养物质。用于通过连接鼻饲管或胃管向患者肠胃输送营养。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
609 | 14-04-01 | 电子气压止血带 | 由止血带主机、气管和传感器等组成，用于对病人四肢处手术中暂时阻断肢体的血供，为手术提供一个无血的手术视野。 | Ⅱ
610 | 14-04-02 | 无源止血器 | 通常由压迫装置和固定装置组成。压迫装置可分为压迫垫、压迫气囊等，固定装置类型分为螺母、固定板、固定带、粘合贴等。无菌提供。 | Ⅱ
611 | 14-04-03 | 无源止血带 | 通常分为捆扎型、卡扣型、松紧带粘贴型，加压压迫血管以达到止血目的。用于临床静脉抽血、静脉穿刺的止血扎结及事故现场、战场救治四肢出血时进行止血处理。或用于四肢外科手术驱赶四肢中血液。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
612 | 14-04-03 | 股动脉气动压迫装置 | 股动脉气动压迫装置由无菌气动压力圆顶拱架，装有压力表的充气管路，充气源，绑带等附件组成，可按设计、技术参数、辅助功能等不同分为若干种；供血管插管术后压迫股动脉或静脉止血或股动脉假性动脉瘤超声导引的压迫修复用。 | Ⅱ
613 | 14-05-01 | 经皮肠营养导管 | 通常为由硅橡胶或聚氨酯等材料制造的导管，头端有固定装置。球囊的作用是将导管经皮插入胃或空肠内后，充起球囊，起到固定导管的作用。在体内滞留时间大于等于30天。无菌提供，一次性使用。用于为不能经鼻肠营养的患者输送营养物质。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
614 | 14-05-02 | 经鼻肠营养导管 | 通常为由聚氯乙烯或其他材料制成的管路和不能与静脉输液器的鲁尔外圆锥接头相连接的接头组成。可经鼻插入患者胃或十二指肠内，以通过它给入肠营养液或药液。无菌提供，一次性使用。用于经鼻向胃肠道引入营养液等。部分可实现冲洗等其他辅助功能。 | Ⅱ
615 | 14-05-03 | 导尿管 | 一般采用高分子材料制成。部分头端固定有球囊。可将头端插入膀胱，并向体外导尿。部分管身上涂有润滑涂层（不含药物、银盐或抗菌成分涂层），浸湿后便于插入，减轻插管痛苦。无菌提供，一次性使用。用于将病人膀胱中的尿液经尿道向体外导出并导入到集尿容器中。14-05-03项下所有Ⅱ类产品 | Ⅱ
616 | 14-05-04 | 直肠管（肛门管） | 通常由导管和接头组成。一般采用医用聚氯乙烯塑料制成。导管头端分有孔和无孔两种，两侧有一个或多个侧孔。无菌提供。用于供肠道清洁（冲洗、排空或灌注）用或排气用。 | Ⅱ
617 | 14-05-05 | 输尿管支架 | 通常为由硅橡胶、聚氨酯或其他聚合物制造的管状结构，单端或双端有环状弯曲。放置于肾脏与膀胱之间。通过对人体输尿管进行支撑和引流，用于治疗输尿管堵塞和狭窄。豁免情况不包括使用新型材料、新作用机理、新功能的产品。在体内滞留时间小于30天。 | Ⅱ
618 | 14-05-05 | 输尿管支架 | 通常为由硅橡胶、聚氨酯或其他聚合物制造的管状结构，单端或双端有环状弯曲。放置于肾脏与膀胱之间。通过对人体输尿管进行支撑和引流，用于治疗输尿管堵塞和狭窄。豁免情况不包括使用新型材料、新作用机理、新功能的产品。在体内滞留时间大于等于30天。 | Ⅲ
619 | 14-05-06 | 一次性使用胆道引流管 | 一次性使用胆道引流管可采用硅胶等适用材料制成，可由螺旋形引流端、管身和接头等组成；可按材质、设计及技术参数等不同分为若干型号及规格；产品以无菌形式提供；供内窥镜胆道结石手术或胆囊摘除术后胆汁引流用。 | Ⅱ
620 | 14-05-06 | 一次性使用体表引流球/管 | 一次性使用体表引流球/管可采用硅橡胶、乳胶等适用材料制成，可设计成多种形状，管身有侧孔，可接负压球，可按材质、设计、管径、适用部位等不同分为若干型号及规格；以无菌形式提供；供体表创伤或切口引流用。 | Ⅱ
621 | 14-05-06 | 一次性使用输尿管引流管 | 一次性使用输尿管引流管可采用聚氨酯等适用材料制成，通常为直型，管身有侧孔；可按管径、长度等不同分为若干规格，以无菌形式提供；供输尿管支撑、引流用。 | Ⅱ
622 | 14-05-06 | 鼻胰/胆引流管 | 通常由引流管、连接管和鼻转换管等部件组成。由高分子材料制成。产品经灭菌，一次性使用。鼻胰/胆引流管被设计用于经内窥镜直视插入胰腺管/胆道，引流胰液/胆汁。 | Ⅱ
623 | 14-00-00 | 一次性使用胸腔引流管及穿刺套件 | 产品以医用硅橡胶/聚氨酯等为原材料，用于胸腔积液的引流。穿刺套件可由扩张器、穿刺针、导丝、导管鞘及附件等组成，用以建立皮下人工通道，辅助引流管的放置。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅱ | 修订分类编码、产品名称和产品描述。
624 | 14-00-00 | 一次性使用腹腔引流管及穿刺套件 | 产品以医用硅橡胶/聚氨酯等为原材料，用于腹腔积液的引流。穿刺套件可由扩张器、穿刺针、导丝、导管鞘及附件等组成，用以建立皮下人工通道，辅助引流管的放置。豁免情况不包括使用新型材料、新作用机理、新功能的产品。 | Ⅱ | 修订分类编码、产品名称和产品描述。
625 | 16-07-12 | 泪道管/泪道引流管 | 通常由硅胶管、硅胶矛、扩张器和穿孔塞组成，泪道管采用硅橡胶材料制成，配备专用手术牵引钩或导丝等附件。用于泪道阻塞探通术后、泪囊炎鼻腔泪囊吻合术后、泪小管断裂吻合术后的泪道支撑与植入治疗。与已获准境内注册产品具有等同性。产品组成材料成熟，新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ | 修订分类编码。
626 | 14-05-07 | 输尿管扩张导管 | 产品可由球囊、导管、导管座等部件组成。产品经灭菌，一次性使用。用于输尿管狭窄的经腔扩张，或在输尿管镜检查或结石操作之前进行输尿管扩张。 | Ⅱ
627 | 14-05-07 | 鼻窦球囊导管 | 通常由球囊导管、插入探针及冲洗管等部件组成。本产品经灭菌,一次性使用。用于在诊断和治疗程序中,扩张窦口和鼻窦腔内空间的器械;也可用于在目标窦内进行冲洗,从而完成治疗程序和方便诊断程序。 | Ⅱ
628 | 14-05-07 | 经皮经肝胆管引流术穿刺套件 | 用于经皮经肝胆管引流术，可由扩张器、穿刺针、导丝、导管鞘及附件等组成，用以建立皮下人工通道，辅助引流管的放置。 | Ⅱ | 新增 | 新
629 | 14-05-08 | 造影导管 | 通常由导管和导管座组成。无菌提供，一次性使用。在导引器械的配合下导管插入体内的某个部位（非血管组织），用于向靶向部位输入造影剂。 | Ⅱ
630 | 14-05-09 | 测压导管 | 通常由导管、导管座等组成，常在导引器械的配合下使导管插入体内的某个部位，供测量生理压力。无菌提供，一次性使用 用于测量靶向部位（非血管组织）生理压力。 | Ⅱ
631 | 14-06-01 | 颅脑外引流收集装置 | 通常由体外引流管路、滴瓶和收集容器等附件组成。用于脑脊液脑室穿刺外引流，收集装置还应有颅内压刻度尺（脑室外引流以外的颅脑穿刺外引流收集装置可不包括颅内压刻度尺，可根据实际需要选配）。无菌提供，一次性使用。如能测量颅内压应提供颅内压测量准确性的临床验证资料。豁免情况不包括插入病人脑室的装置、向脑室给药的装置及药物成分。 | Ⅱ | 修订产品描述。
632 | 14-06-02 | 胸腔引流装置 | 通常至少由胸腔引流接管和积液腔组成，可分为水封式、干封阀式等，部分带有吸引接口，以实现吸引引流。不包括带有自体血回输的引流装置。从生产工艺上可分为注塑和吹塑两种。不包括插入病人胸腔的胸腔引流导管。无菌提供，一次性使用。与胸腔引流导管配套，用于气胸、胸腔积液及手术后需要进行闭式引流的患者。 | Ⅱ
633 | 14-06-03 | 负压引流器及组件 | 通常由器身、弹簧、调节器组、连接接头、连接管、止流夹、护帽、引流袋、负压泵或手动负压源（负压球）组成。无菌提供。不包括插入体内的引流导管。用于临床负压引流时，与插入体内的引流导管相连接，起到充当负压传导介质和/或引导、收集引流液的作用。 | Ⅱ
634 | 14-06-03 | 一次性使用负压引流器 | 通常由器身、弹簧、调节器组、连接接头、连接管、止流夹、护帽、引流袋、负压泵或手动负压源（负压球）组成。无菌提供。不包括插入体内的引流导管。用于临床负压引流时，与插入体内的引流导管相连接，起到充当负压传导介质和/或引导、收集引流液的作用。 | Ⅱ
635 | 14-06-04 | 真空负压机 | 通常由真空泵、管路、接头等组成。通过真空泵抽吸，使系统管路产生负压，达到持续或间歇进行创面引流或者提供负压环境辅助伤口闭合的目的。用于去除腔隙或创面分泌物和坏死组织。 | Ⅱ
636 | 14-06-06 | 负压引流封闭膜 | 通常为聚氨酯膜上均匀涂有合成胶粘物质组成封闭膜的粘贴面，粘贴面上覆盖有一个保护层。对创面进行负压引流时，将负压引流海绵置于创面上，贴上负压引流封闭膜，在海绵的支撑作用和封闭膜的封闭作用下形成封闭的负压引流腔体，在吸引导管与外接的吸引源的作用下实现对创面的引流。无菌提供，一次性使用。 | Ⅱ | 新增 | 新
637 | 14-06-07 | 医用电动吸引器械 | 通常由负压指示器、过滤器、收集容器和防溢流装置组成，以网电源或蓄电池驱动。负压可由旋片泵、活塞泵、膜片泵和电磁泵提供，利用负压抽出人体中的气体、液体和/或固体。可在医疗场所、家庭、野外或运输途中使用。可吸取人体呼吸道中的分泌物；胃中、胸腔中的气体、液体；体表创面、伤口的渗出液；羊水；手术中、手术后的血水（冲洗液）、渗出液、废液等。不包括仅用于妇科流产、口腔科吸引、眼科吸引或内窥镜手术的吸引器械。 | Ⅱ
638 | 14-06-07 | 医用负压吸引器 | 医用负压吸引器主要由负压系统、压力显示、主控电路、储液容器组成，通常不包括吸引管，可负压产生原理、设计、技术参数、预期用途、附加辅助功能等不同分为若干型号；供医院和家庭为病人作吸痰等普通负压吸引用，不适用于流产吸引用，不包括以负压或压力源为动力的吸引设备。产品性能指标采用下列参考标准中的适用部分，如：YY 0636.1-2008 医用吸引设备 第1部分：电动吸引设备 安全要求。 | Ⅱ
639 | 14-06-08 | 以负压源或压力源为动力吸引器械 | 通常由负压指示器、过滤器、收集容器和防溢流装置组成，可含有电气控制元件。与负压源连接使用，由中心吸引系统负压管道或压缩气体驱动，也可由自带气瓶驱动。与负压管道连接还需要配备负压调节器。利用负压抽出人体中的气体、液体和/或固体。可在医疗场所、家庭、野外或运输途中使用。可吸取人体呼吸道中的分泌物；胃中、胸腔中的气体、液体；体表创面、伤口的渗出液；羊水；手术中、手术后的血水（冲洗液）、渗出液、废液等。不包括仅用于妇科流产、口腔科吸引、眼科吸引或内窥镜手术的吸引器械。 | Ⅱ
640 | 14-06-09 | 医用人工驱动吸引器械 | 通常由负压指示器、过滤器、收集容器和防溢流装置组成。包括脚踏、手动或两者并用，也可以与电气设备组合使用，利用负压抽出人体中的气体、液体和/或固体。可在医疗场所、家庭、野外或运输途中使用。可吸取人体呼吸道中的分泌物；胃中、胸腔中的气体、液体；体表创面、伤口的渗出液；羊水；手术中、手术后的血水（冲洗液）、渗出液、废液等。不包括仅用于妇科流产、口腔科吸引、眼科吸引或内窥镜手术的吸引器械。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅱ
641 | 14-06-10 | 医用真空负压机组 | 一般由真空泵、真空罐、管路、阀门、控制柜和真空仪表组成，用于医用气体管路系统，提供负压吸引源。 | Ⅱ | 与原632合并修订产品名称及产品描述。
642 | 14-06-11 | 体外引流、吸引管 | 通常由软管和连接件组成。能在引流导管与引流装置之间连接，使之组成密闭的引流系统。不直接接触人体。无菌提供，一次性使用。 与适宜设备配套后，用于手术中、手术后的血水、废液等引流、吸引使用。 | Ⅱ
643 | 14-06-12 | 无菌引流袋（容器）、无菌集尿袋（容器） | 通常为袋式的收集容器。通过体外管路与引流导管连接，形成密闭的引流系统。无菌提供，一次性使用。 | Ⅱ
644 | 14-06-12 | 无菌引流袋/尿袋 | 本产品可由收集袋、输入管、排放阀、抗返流片和引流管接头等部件组成。通过体外管路与引流导管连接，形成密闭的引流系统。可有抗返流功能，废液不会返流至人体。原材料一般为医用PVC塑料和聚丙烯等高分子材料。灭菌产品，一次性使用。用于患者治疗过程中收集分泌物和体液。 | Ⅱ
645 | 14-06-12 | 引流袋（容器）/收集袋（容器）、粪便管理器械 | 通常为袋式的收集容器。通过体外管路与引流导管连接，形成密闭的引流系统。无菌提供，一次性使用。用于医院临床科室及手术中或手术后患者一次性引流体液（血液、胃液等）、分泌物（痰液、冲洗液等）以及人体排泄物的收集。 | Ⅱ
646 | 14-06-13 | 无菌球囊扩充压力泵 | 通常由推注系统、压力表、连接管路等组成。无菌提供。 | Ⅱ
647 | 14-07-01 | 无菌冲洗器 | 根据不同的预期用途有不同的结构。使用前装入冲洗液，或与相关的冲洗设备或器具连接成冲洗系统，可向患者冲洗部位进行冲洗。无菌提供。 | Ⅱ
648 | 14-07-01 | 电动洗胃机 | 电动洗胃机主要由压力泵、控制单元、冲吸转换装置、过滤瓶、软管、冲洗液和污水容器等组成，可按压力泵、控制方式、技术参数、附加辅助功能等不同分为若干型号。供医疗单位作抢救服毒、食物中毒及手术前洗胃用。产品性能指标采用下列参考标准中的适用部分，如：YY 1105-2008 电动洗胃机。 | Ⅱ
649 | 14-07-01 | 鼻腔冲洗器 | 鼻腔冲洗器采用塑料制成，由瓶体、瓶盖和出水管（可带流量控制阀）组成，不包括冲洗液；可按材质、设计、技术参数（如容量）等不同分为若干型号及规格；可以无菌形式提供；供鼻腔冲洗用。 | Ⅱ
650 | 14-07-01 | 冲洗升温套件 | 用于手术室内冲洗液的加热。 | Ⅱ
651 | 14-07-01 | 电动冲洗器 | 通常由主机、加温模块、电源模块组成。不含冲洗液。用于冲洗自然腔道（不包括阴道专用）、术中术后冲洗组织，也可对冲洗液加温。 | Ⅱ
652 | 14-07-02 | 无菌灌肠器 | 通常由软袋、夹子、接管、接头等部分组成，一般采用玻璃、高分子等材料制成。软袋供装入灌肠液，接头可与直肠导管连接。不含内容物。无菌提供。 | Ⅱ
653 | 14-07-03 | 无菌给药器 | 通常由吸嘴、腔体和吸入器连接口组成。采用高分子和金属材料制成。也可包括气雾剂药物吸入给药的装置。不含药物。 | Ⅱ
654 | 14-09-01 | 外科织造布类敷料 | 对应《医疗器械分类目录》中14-09-01中的Ⅱ类产品（除止血纱布）。产品性能指标符合相关标准（含修改单）中的适用部分。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
655 | 14-09-01 | 可显影纱布 | 带有X射线可探测钡线（片）的脱脂棉纱布，一次性使用无菌产品。作为外科手术用敷料，具有X射线可探测功能。 | Ⅱ
656 | 14-09-02 | 外科非织造布类敷料 | 对应《医疗器械分类目录》中14-09-02中的Ⅱ类产品。通常为由非织造敷布经过加工制成的敷料，供临床护创、吸湿用。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
657 | 14-09-03 | 外科海绵敷料 | 通常为由高分子材料加工成的海绵状敷料。无菌提供，一次性使用。用于吸收手术过程中的体内渗出液、手术过程中承托器官、组织等。还用于腔道（如鼻腔）的填塞压迫止血。 | Ⅱ
658 | 14-10-01 | 创面敷贴 | 对应《医疗器械分类目录》中14-10-01中举例的创面敷贴。通常由涂胶基材、吸收性敷垫和可剥离的保护层组成。其中吸收性敷垫一般采用棉纤维、无纺布等可吸收渗出液的材料制成。吸收性敷垫可单独使用，用绷带或胶带等进行固定。所含成分不可被人体吸收。无菌提供，一次性使用。用于体表非慢性创面的护理。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品，如新材料、新作用机理、新功能的产品。 | Ⅱ
659 | 14-10-01 | 固定贴 | 对应《医疗器械分类目录》中14-10-01中的Ⅱ类产品。由背衬和粘胶剂等组成，一次性使用无菌产品。用于对穿刺器械（如导管）的穿刺部位的护理并固定穿刺器械或用于敷料的次级固定。豁免情况不包括：新型结构设计、新型作用机理、新预期用途、含活性成分等。 | Ⅱ
660 | 14-10-02 | 创口贴 | 对应《医疗器械分类目录》中14-10-02中的Ⅱ类产品。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
661 | 14-10-04 | 凝胶敷料 | 通常为成胶物质与水组成的定形或无定形凝胶敷料，可含有缓冲盐。所含成分不可被人体吸收，不具有药理学作用。无菌提供。用于吸收创面渗出液或向创面排出水分，用于手术后缝合创面等非慢性创面的覆盖。 | Ⅱ
662 | 14-10-05 | 水胶体敷料 | 对应《医疗器械分类目录》中14-10-05中举例的Ⅱ类水胶体敷料、水胶体敷贴、水胶体贴。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
663 | 14-10-05 | 水胶体敷料 | 对应《医疗器械分类目录》中14-10-05中举例Ⅲ类水胶体敷料或水胶体敷贴。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
664 | 14-10-06 | 藻酸盐敷料 | 由藻酸盐纤维组成，预期用途仅限于覆盖体表非慢性创面、吸收创面渗液的Ⅱ类藻酸盐敷料。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
665 | 14-10-06 | 亲水性纤维敷料 | 主要由羧甲基纤维素钠或乙基磺酸盐纤维素组成，预期用途限于覆盖体表非慢性创面、吸收渗液的Ⅱ类亲水性纤维敷料。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
666 | 14-10-06 | 藻酸盐敷料 | 由藻酸盐纤维组成，预期用途仅限于覆盖创面、吸收创面渗液的Ⅲ类藻酸盐敷料和藻酸盐填充条。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
667 | 14-10-06 | 亲水性纤维敷料 | 主要由羧甲基纤维素钠或乙基磺酸盐纤维素组成，预期用途限于覆盖创面、吸收渗液的Ⅲ类亲水性纤维敷料。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
668 | 14-10-07 | 聚氨酯泡沫敷料 | 主要由聚氨酯泡沫组成，预期用途仅限用于覆盖体表非慢性创面、吸收创面渗液的Ⅱ类聚氨酯泡沫敷料。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
669 | 14-10-07 | 聚氨酯泡沫敷料 | 主要由聚氨酯泡沫组成，预期用途仅限用于覆盖创面、吸收创面渗液的Ⅲ类聚氨酯泡沫敷料。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
670 | 14-10-08 | 液体、膏状敷料 | 通常为溶液或软膏（不包括凝胶），所含成分不具有药理学作用，不可被人体吸收。无菌提供。通过在创面表面形成保护层，起物理屏障作用。用于小创口、擦伤、切割伤等非慢性创面及周围皮肤的护理。 | Ⅱ
671 | 14-10-09 | 凡士林/石蜡纱布（油纱） | 由医用脱脂纱布、凡士林、石蜡油制成，预期用途仅限用于创面引流、覆盖、填塞的凡士林/石蜡纱布或纱布条。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
672 | 14-10-09 | 水胶体油纱 | 主要由凡士林、水胶体颗粒（羧甲基纤维素钠）和网状聚酯纤维组成。用于创面引流、覆盖、填塞。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅲ
673 | 14-11-01 | 急救止血绷带、带敷贴绷带 | 对应《医疗器械分类目录》中14-11-01中举例的Ⅱ类绷带。通常为纺织加工而成的卷状、管状、三角形的材料。部分具有弹力或自粘等特性。用于体表非慢性创面的护理，或用于对创面敷料或肢体提供束缚力，以起到包扎、固定作用。新型结构设计、新型作用机理、新预期用途、含活性成分等除外。 | Ⅱ
674 | 14-12 | 造口袋（含底盘） | 本产品包括肠造口袋（开口袋或闭口袋）和尿路造口袋。肠造口袋主要由外袋、底盘和过滤片组成。尿路造口袋主要由多腔式外袋(带抗返流结构)、底盘和引流阀组成。以无菌形式提供，一次性使用。用于处理造口排泄物。 | Ⅱ
675 | 14-13 | 一次性使用无菌洁净服 | 一次性使用无菌洁净服一般由具有一定防护特性的无纺布或覆膜无纺布复合材料制成，可由衣领、衣身、衣袖组成，也可由上衣、裤子组成，袖口、脚踝口为弹性收口，腰部收口可采用弹性收口、拉绳收口或搭扣；可分为褂式、连体式、分体式等；无菌提供，供进入手术洁净区人员穿着以降低因人员携带感染源造成的洁净区空气微生物污染，防止病人术后创面感染。产品性能指标采用下列参考标准中的适用部分，如： YY/T 0506 病人、医护人员和器械用手术单、手术衣和洁净服。 | Ⅱ
676 | 14-13-01 | 一次性使用无菌手术单 | 一次性使用无菌手术单一般由具有一定防护特性的无纺布或覆膜无纺布复合材料制成，以无菌形式提供，供覆盖在患者身体表面，降低患者皮肤等非手术部位感染源向手术部位移行，防止病人术后创面感染，也可因设计而具有收集手术污物等功能。产品性能指标采用下列参考标准中的适用部分，如： YY/T 0506 病人、医护人员和器械用手术单、手术衣和洁净服。 | Ⅱ
677 | 14-13-01 | 手术洞巾 | 手术洞巾可用具有一定防护特性的无纺布或覆膜无纺布复合材料制成，可按设计型式和尺寸不同分为若干型号和规格，通常以无菌形式提供，供覆盖在患者手术皮肤创面周围，降低患者手术创口面皮肤的感染源向手术创面部位移行，防止病人术后创面感染。 | Ⅱ
678 | 14-13-01 | 无菌保护套 | 通常由开口端和套体组成，一般采用聚乙烯等高分子材料制成，无菌提供。用于覆盖外科手术室器械台、操作台、显示屏等上，或用于隔离器械（如胃镜）与患者，避免手术中的医生接触上述部位后，再接触手术中的病人伤口部位造成感染。 | Ⅱ
679 | 14-13-02 | 一次性使用无菌手术膜 | 一次性使用无菌手术膜一般由天然胶乳（可不涂医用压敏胶）或聚氨酯、聚乙烯基材上涂敷低致敏性压敏胶，上置离型纸制成，以无菌形式提供；供简化术前护皮操作，预防接触性和移行性手术创面感染。产品性能指标采用下列参考标准中的适用部分，如：YY 0852 一次性使用无菌手术膜。 | Ⅱ
680 | 14-13-03 | 一次性使用灭菌外科手套 | 一次性使用灭菌外科手套采用天然橡胶胶乳或合成橡胶等材料制成，以无菌形式提供，供临床外科操作时防护使用。产品性能指标采用下列参考标准中的适用部分，如：GB 7543 一次性使用灭菌橡胶外科手套。 | Ⅱ
681 | 14-13-04 | 一次性使用医用外科口罩 | 一次性使用医用外科口罩通常用无纺布或无纺布复合材料制成，可为三层结构，有可塑性鼻夹，口罩带可为弹性或非弹性，具有较小的气流阻力、具有合成血液阻隔、过滤颗粒物和细菌、阻燃等特性；供临床医务人员在有创操作过程中佩带，覆盖住使用者的口、鼻及下颌，为防止病原体微生物、体液、颗粒物等的直接透过提供物理屏障。产品性能指标采用下列参考标准中的适用部分，如：YY 0469 医用外科口罩。豁免情况不包括具有抗病毒等新功能的产品。 | Ⅱ | 修订产品描述。
682 | 14-13-05 | 可重复使用手术衣 | 通常为基材和阻水层组成的手术室服装。基材一般由非织造布或纺织布制造，阻水层为阻水性的材料。非无菌提供可重复使用。手术衣按关键区域的屏障能力分为标准型和高性能型两种。 | Ⅱ
683 | 14-13-05 | 一次性使用手术衣帽 | 一次性使用手术衣帽一般由具有一定防护特性的无纺布或覆膜无纺布复合材料制成；其中手术衣可由衣领、衣身、衣袖组成，也可由上衣、裤子组成，袖口、脚踝口为弹性收口，腰部收口可采用弹性收口、拉绳收口或搭扣；前襟、肘部等可设计成加强防护，以无菌形式提供，供手术人员穿着，降低感染源向病人手术创面传播以防止术后创面感染；具有阻止液体透过的手术衣，也可减小血液或体液中携带的感染源向手术人员传播的风险。产品性能指标采用下列参考标准中的适用部分，如： YY/T 0506 病人、医护人员和器械用手术单、手术衣和洁净服。 | Ⅱ
684 | 14-14-01 | 医用防护口罩 | 医用防护口罩用具有较小的气流阻力、合成血液阻隔、特定强度、过滤效率、表面抗湿性及阻燃性能的适宜材料制成，为自吸式，由口罩体（含鼻夹）和口罩带组成，可按过滤效率不同分为3级，供医疗工作环境下，过滤空气中的颗粒物、阻隔飞沫、血液、体液、分泌物等用。产品性能指标采用下列参考标准中的适用部分，如：GB 19083 医用防护口罩技术要求。豁免情况不包括具有抗病毒等新功能的产品。 | Ⅱ
685 | 14-14-02 | 医用一次性防护服 | 医用一次性防护服用具有特定液体阻隔功能、特定强度、过滤效率、阻燃性能、抗静电和静电衰减功能的适宜材料制成，由连帽上衣、裤子组成，袖口、脚踝口为弹性收口，帽子面部收口及腰部收口可采用弹性收口、拉绳收口或搭扣；可分为连身式或分体式两类，供临床医务人员在工作时接触到的具有潜在感染性的患者血液、体液、分泌物等提供阻隔、防护用。产品性能指标采用下列参考标准中的适用部分，如： GB 19082 医用一次性防护服技术要求。 | Ⅱ
686 | 14-14-04 | 手部防护用品 | 通常采用聚氯乙烯、橡胶或不锈钢等材料制造。有足够的强度和阻隔性能。无菌提供，一次性使用。用于戴在医生手上或手指上对患者病情进行检查或触检，或用于防止医生手部被咬伤。 | Ⅱ
687 | 14-15-01 | 婴儿光疗防护眼罩 | 婴儿光疗防护眼罩可由眼罩和眼罩固定两部分组成；眼罩可由优质无纺布/蓝黑物理复合布等制成，固定部分可用弹力带或柔软魔术贴等材料制成；供婴儿蓝光照射治疗时的眼部防护用。 | Ⅱ
688 | 14-15-02 | 眼贴 | 对应《医疗器械分类目录》中14-15-02中的Ⅱ类产品。可由基层（无纺布等）、医用聚乙烯醇水凝胶和离型纸组成；可按形状、尺寸等不同分为若干型号和规格；贴敷于全麻或深度昏迷病人的眼外部，通过提供相对密闭的潮湿环境，预防暴露性角膜炎。豁免情况不包括含有活性成分的产品，如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等。 | Ⅱ
689 | 14-15-03 | 鼻腔止血器 | 通常由充气阀、导气管、气囊组成，无菌提供，一次性使用。 通过充气后的气囊压迫，抑制鼻腔出血。 | Ⅱ | 新增 | 新
690 | 14-15-04 | 海水鼻腔喷雾器 | 通常由喷雾器和喷雾液构成。喷雾液由纯净水和海水组成。 用于急慢性鼻炎、过敏性鼻炎、鼻息肉、鼻窦炎等鼻腔疾病患者的鼻腔清洗，也用于鼻炎手术后及化疗后的鼻腔清洗。 | Ⅱ
691 | 14-15-05 | 无菌垫单 | 通常由非织造布和塑料膜复合或缝制而成。无菌提供，一次性使用。病床或检查床上用的卫生护理用品 | Ⅱ
692 | 14-16 | 皮肤注射位点标记用彩格 | 由上面印有化妆品级墨水的纹身纸和乙酸乙烯酯聚合物薄膜组成。其中墨水由食品级色素、乙基纤维素和蓖麻油组成。产品的作用机理及预期用途为：将墨水基质形成的彩色小点组成的多个彩格转印在完好皮肤上，用来给医生提供注射间隔的参考（1平方厘米间隔的网格）。 | Ⅱ
693 | 14-16-01 | 洁净屏 | 通常由箱体、风机组、过滤器、电气控制器等组成。用于手术过程中患者手术部位局部环境空气的净化，以防止感染。 | Ⅱ
694 | 14-16-03 | 预充式导管冲洗器 | 由0.9%氯化钠注射液、外套、芯杆、橡胶活塞组成。用于不同药物治疗的间隙，封闭、冲洗导管的管路末端。 | Ⅲ | 新增 | 新
695 | 14-16-05 | 鼻贴（不含药） | 鼻贴（不含药）将具有弹性的塑料片通过强力胶和基布固定在鼻梁部位，供扩张鼻孔，缓解鼻塞用。 | Ⅱ
696 | 14-16-05 | 上呼吸道扩张器 | 通过扩张鼻腔、上下颌或对其进行矫正，仅用于宣称堵塞式呼吸暂停的辅助治疗，用于扩张鼻孔，缓解鼻塞。接触人体时间小于30天。 | Ⅱ
697 | 14-16-05 | 鼻腔润滑剂 | 通常由聚山梨醇酯80、甘油、氯化钠、依地酸钠、山梨酸钾和纯净水组成。通过润滑和保持粘膜湿润，降低上呼吸道阻力，以改善呼吸受阻状况。 | Ⅱ
698 | 14-16-06 | 无菌咬口 | 手术或检查时患者开口的辅助器械，一般由聚乙烯等高分子材料制成。无菌提供。用于经口腔手术或检查时维持患者的开口状态，防止非预期咬合，或便于插入和固定气管插管。 | Ⅱ
699 | 14-16-08 | 体表器械固定装置 | 通常是能专门为某一种或某一类器械的使用提供固定的装置。与创口接触。无菌提供，一次性使用。用于固定使用过程中的医疗器械 | Ⅱ
700 | 14-16-09 | 医用无菌液体石蜡无纺布/医用石蜡棉球 | 医用无菌液体石蜡无纺布由非织造布/棉球和液状石蜡制成，可按配方、材质、尺寸等不同分为若干型号及规格，可以无菌形式提供；供患者检查部位表面及器械表面润滑用。豁免情况不包括：（1）适应症宣称可以促进上皮化、引导组织再生、促进伤口愈合、减轻疼痛、抗菌、防感染、抗病毒、止血、溶解坏死组织、减少疤痕、防粘连、作为人工皮/皮肤替代物等作用的产品；（2）宣称可以用于体内伤口、三度烧伤、感染创面、坏死组织较多的创面、发生创面脓毒症的患者等情况的产品；（3）含有活性成分的产品：如药品/药用活性成分、生物制品/生物活性成分、银、消毒剂等；（4）其他新型产品,如新材料、新作用机理、新功能的产品。 | Ⅱ
701 | 14-16-09 | 润滑剂 | 润滑剂通常由甘油、黄原胶等成分组成，不含药物成分。临床上用于器械进入人体自然腔道时的润滑。 | Ⅱ
702 | 14-16-10 | 消毒用棉棒/球/片 | 含有酒精、碘酊或碘伏的棉棒/球/片，用于注射、输液前消毒完整皮肤。 | Ⅱ
703 | 14-16-10 | 无菌棉棒/球/片 | 通常包括吸水性材料。为了方便使用，部分产品有供手持的组件。不含消毒剂。无菌提供，一次性使用。用于对皮肤、创面进行清洁处理。 | Ⅱ
704 | 14-16-13 | 无菌接管机 | 通常由主机、控制器、熔接装置、显示屏组成。用于将两根医用管路无菌地接合在一起。 | Ⅱ
705 | 15-00-00 | 耳鼻喉检查治疗台 | 耳鼻喉科检查治疗台通常由操作检查治疗台、主机和附件（喷枪、吸枪、吹枪、除雾、射灯及支架、托盘、棉球缸、监视器、冷光源、CCD摄像机、紫外线消毒灯、观片灯等），可按结构组成、功能与技术参数等不同分为若干型号及规格；可与耳鼻喉内窥镜配套，供耳鼻喉科作检查、诊断、治疗等用。 | Ⅱ
706 | 15-01-01 | 电动手术台 | 电动手术台主要由台面、台座、电气控制系统、配套件等部分组成，按驱动方式、技术参数、辅助功能、配套件等不同分为若干型号及规格，可分别适用于头、颈、胸、腹腔、会阴、四肢等外科、妇科、泌尿科、五官科、骨科等手术承载患者，提供合适的手术体位用。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1106－2008电动手术台，YY 0570-2005 医用电气设备 第2部分：手术台安全专用要求。 | Ⅱ
707 | 15-01-01 | 电动液压手术台 | 电动液压手术台主要由台面、台座、电动液压控制系统（通常台面升降由液压控制）、配套件等部分组成，按驱动方式/部位、技术参数、辅助功能、配套件等不同分为若干型号及规格，可分别适用于头、颈、胸、腹腔、会阴、四肢等外科、妇科、泌尿科、五官科、骨科等手术承载患者，提供合适的手术体位用。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1106－2008电动手术台，YY 0570-2005 医用电气设备 第2部分：手术台安全专用要求。 | Ⅱ
708 | 15-01-02 | 手动手术台（液压） | 通常由床体（包括支撑部分、传动部分和控制部分）和配件组成。支撑部分包括台面、背板、臀板、板等，可调节。升降形式为液压升降式，体位调整均为人力操纵。 | Ⅱ
709 | 15-01-02 | 液压手术台 | 液压手术台主要由台面、台座、液压控制系统、配套件等部分组成，按驱动部位、技术参数、辅助功能、配套件等不同分为若干型号及规格，可分别适用于头、颈、胸、腹腔、会阴、四肢等外科、妇科、泌尿科、五官科、骨科等手术承载患者，提供合适的手术体位用。 | Ⅱ
710 | 15-02-01 | 电动诊疗台及诊疗椅 | 诊疗（查）台通常由床架、床面、枕头等组成。诊疗（查）椅通常由基座、背板、坐椅和搁脚架组成。有源产品。 | Ⅱ
711 | 15-02-01 | 电动检查椅 | 电动检查椅一般由底座、背板、坐板、脚托和脚控开关组成。用于检查和治疗时支撑患者。 | Ⅱ
712 | 15-03-01 | 电动病床 | 电动病床为网电源或内部电源或组合式电驱动式设备，可由床面部分（如背板、坐板、腿板、脚板）、床架部分（前后床头板、左右护栏等）、下床架部分（升降机构、中控机构、脚轮等）、电控部分（升降电机、背板电机、腿板电机、控制器、手控板）、输液架等附件组成，按功能配置等不同分为若干型号，供临床各类病房承载和护理患者用。产品性能指标采用下列参考标准中的适用部分，如：YY 0571-2013 医用电气设备 第2部分:医院电动床安全专用要求。 | Ⅱ
713 | 15-04-01 | 电动患者手术位置固定辅助器械 | 通常由主机、通用连接附件、床夹等组成。有源产品。用于手术中，膝、肩、髋及小关节的固定及术中调整。 | Ⅱ
714 | 15-05-01 | 传染病员运送负压隔离舱 | 通常由负压舱体（活动舱盖、舱底）、排风过滤装置组成。通常工作时充气展开，储运时放气可折叠。用于传染病员的安全隔离转运。可与担架车、生命体征检测仪器配合使用，对传染病患者进行运送途中监护。 | Ⅱ |  | 新
715 | 15-05-02 | 电动推车、担架等器械 | 通常由床面（推车面板、担架面）、支撑架组成。可附加输液架、护栏、定向轮踏板、脚轮、刹车、手摇机构等。为单车或双车。有源产品。 | Ⅱ
716 | 15-05-02 | 转运床 | 转运床可由床架、床板、脚轮、床垫、侧护栏、制动系统、液压升降及控制系统、推动手柄、及附件（如：约束带、输液架）组成。适用于医院对患者的转运或紧急救护。产品性能指标采用下列参考标准中的适用部分，如： YY 0571-2005。 | Ⅱ
717 | 15-05-02 | 电动担架车 | 该产品通常由床面（推车面板、担架面）、支撑架组成。可附加输液架、护栏、定向轮踏板、脚轮、刹车、手摇机构等。为单车或双车。有源产品。用于医疗机构运送、移动患者用。 | Ⅱ
718 | 15-06-01 | 电动防压疮（褥疮）垫 | 通常由充气床垫、气道（连接管）、充气泵等组成。防压疮（褥疮）气垫由若干个气室组成。气垫由气泵充气后，气室维持一定气压，所形成的软性垫，可增加患者身体与垫接触面积，降低身体局部压力。 | Ⅱ
719 | 15-06-01 | 医用防褥疮气垫 | 医用防褥疮气垫主要由气囊、电动充气泵、控制器、按摩电机、床罩组成；可按材质、设计、技术参数、充气模式等不同分为若干型号；供卧床患者褥疮防治用。 | Ⅱ
720 | 16-00-00 | 无源眼科手术用器械 | 无源短期(≤24小时)使用器械，市场应用多年，非创新功能设计，非新材料制成，无涂层，无降解、吸收功能，不含药物、动物源性成分，一般为一次性使用的灭菌产品，用于辅助完成眼科手术、检查。 | Ⅱ
721 | 16-01-01 | 一次性使用眼科手术用刀 | 一次性使用眼科手术用刀由刀柄（聚碳酸酯）与刀片（不锈钢）组成，可按材质、设计、技术参数、预期用途等不同分为若干型号与规格；以无菌形式提供；供显微眼科手术用。 | Ⅱ
722 | 16-02-01 | 一次性使用眼科穿刺系统 | 一次性使用眼科穿刺系统可包括穿刺刀、穿刺套管、手柄、保护帽、自闭阀及灌注管头端，可分别用不锈钢、聚碳酸酯、聚丙烯、硅树脂等材料制成；可按材质、设计、技术参数等不同分为若干型号及规格；产品以无菌形式提供；供玻璃体切割术等眼科手术中，在巩膜平坦部建立器械通道用。 | Ⅱ
723 | 16-03-01 | 验光设备和器具 | 分类编码16-03-01下所有产品均免于进行临床评价。通常由球镜度片、柱镜度片、棱镜度片、辅助镜片和机械换片结构组成。利用被检者对视标成像清晰程度的主观表述，测出被检者的屈光状态，与视力表配合使用。用于主观测量人眼屈光状态。 | Ⅱ |  | 新
724 | 16-03-01 | 检影镜 | 检影镜通常由折光镜、中心通光反射镜的投射系统、观察系统和电源组成，可分为带状光（矩形截面光束，焦距和方位可调）检影镜和点状光检影镜（近圆截面光束，像焦点可调）两类，每类按电源、技术参数、辅助功能等不同分为若干型号；供眼屈光不正客观评估和视网膜反射交叉越过瞳孔时的位移观察用。产品性能指标采用下列参考标准中的适用部分，如：YY 0718-2009眼科仪器 检影镜。 | Ⅱ
725 | 16-03-01 | 验光镜片 | 验光镜片主要由正、负球镜验光镜片，正、负柱镜验光镜片，棱镜验光镜片，辅助验光镜片（马式杆片、裂隙片、针孔片、黑片、磨砂片、平光片、十字片、红色滤光片、绿色滤光片 、偏振片、交叉柱镜片、翻转镜等组合而成，按验光镜片的数量、种类不同可分为若干型号或种类，可附验光镜架，供检查人眼的屈光状态（远视、近视、老视、散光）和斜视及眼科其他方面的疾病用。产品性能指标采用下列参考标准中的适用部分，如：GB 17342－2009眼科仪器 验光镜片。 | Ⅱ
726 | 16-03-01 | 验光仪 | 验光仪由主体部、电源部和鄂托部组成。验光仪可测眼睛的球镜屈光度、柱镜屈光度、散光轴。 | Ⅱ
727 | 16-03-01 | 验光头 | 验光头是能提供球镜片、柱镜片和棱镜片及其他光学器件的设备，可通过手动或电动方式进行控制，供检查人眼屈光状态。产品性能指标采用下列参考标准中的适用部分：YY 0674-2008《眼科仪器 验光头》 | Ⅱ
728 | 16-03-01 | 液晶视力表 | 液晶视力表主要由液晶显示器、主机和电源线组成。由显示器显示的视力表或卡。用于视力测定。 | Ⅱ
729 | 16-03-02 | 同视机 | 通常由两套视标空间方位可调，光亮可调的独立光学系统及可对两套系统进行空间方位变化测量的机械系统组成，并可结合其他辅助部件（如海丁格刷、偏振片）使用。用于检查人眼的同时视、融像、立体视等双眼视觉功能、以及诊断主客观斜视角、异常视网膜对应、隐斜、后像、弱斜视等眼科疾病，也可供弱视训练、治疗。 | Ⅱ |  | 新
730 | 16-03-02 | 瞳距测量仪、瞳距仪 | 通常由光源、视标、读数系统、机械调节系统、观察或显示系统组成。利用光学成像定位原理，测量人眼瞳距。用于测量人眼两瞳孔之间的距离。 | Ⅱ |  | 新
731 | 16-03-02 | 眼像差仪、全眼波前像差仪 | 通常由光学定位系统、像差测量系统、信号探测器和数据处理分析系统组成。由光线追迹导出的光程差通过拟合获得ZERNIK系数的方法进行波前像差分析。用于测量人眼像差。 | Ⅱ |  | 新
732 | 16-03-02 | 视野计 | 视野计主要由背景屏、刺激点显示装置、颌托、额托、瞳孔监测装置、反应装置、非独立升降台、控制电路/计算机主机等组成，可按设计、技术参数、附加辅助功能等不同分为若干型号；用于对视路疾病的判断和诊断，如青光眼、眼底病等。视野随诊还能用以判断某些疾病的发展状况，用于指导治疗。产品性能指标采用下列参考标准中的适用部分，如： YY 0676-2008眼科仪器 视野计。 | Ⅱ
733 | 16-03-02 | 手动角膜曲率计 | 手动角膜曲率计为手动估读式，由曲率计部件、运动底座、仪器台、电箱、头架五部分组成，可按设计、技术参数、附加辅助功能等不同分为若干型号；供检查角膜的曲率及散光轴位测量用。 | Ⅱ
734 | 16-03-02 | 角膜曲率计 | 角膜曲率计由主机(光学系统、观察系统和控制系统)、可移动工作台和头托构成；可按设计、技术参数、附加辅助功能等不同分为若干型号；供测量患者角膜和角膜接触镜的中心曲率和子午方向用。产品性能指标采用下列参考标准中的适用部分。如： YY 0579-2005角膜曲率计。 | Ⅱ
735 | 16-03-02 | 眼电生理设备 | 产品由刺激器、生物信号放大器、显示装置等组成，用于测量眼部电生理信号，为医生诊断提供参考信息。可进行多种测量程序，如ERG、VEP、MERG、MVEP、EOG等。 | Ⅱ
736 | 16-04-03 | 眼科专用超声脉冲回波设备 | 包括眼科A超、B超和A/B超由主机和探头组成，其探头标称频率一般在10MHz以上，利用超声脉冲回波原理，完成眼科诊断信息采集、显示、测量。专用于眼科的超声诊断设备，可实现眼球及眼眶的超声成像、角膜厚度测量、眼轴长度测量等功能。不包括眼科高频超声诊断仪（通常称为超声生物显微镜UBM） | Ⅱ | 修订管理类别描述。
737 | 16-04-05 | 眼用照相机 | 通常由照明系统、观察系统、成像系统等组成。可与单独的查看软件配合使用，并实现附加功能。用于拍摄眼部图像，观察和诊断视网膜病变。 | Ⅱ
738 | 16-04-07 | 裂隙灯显微镜 | 裂隙灯显微镜是一种眼科双目（也可为三目）显微镜，仪器把一个高亮的窄缝投射到眼睛里，用可移动的显微镜从侧面观察其反射光，从而对反射表面进行观察和/或测量。典型结构由显微镜和能产生裂隙（窄缝）的旋转照明系统组成，可以是手持式的，也可是台式，显微镜可以是连续变倍或非连续变倍；若配备照相机和计算机，也可用软件管理图像，可因设计与技术参数的不同分为若干型号；供检查眼前节及其他眼内部病变用。产品性能指标采用下列参考标准中的适用部分，如：YY 0065-2007眼科仪器裂隙灯显微镜。 | Ⅱ
739 | 16-04-08 | 直接检眼镜 | 直接检眼镜由照明系统和观察系统（由透视镜和光阑组组成），产生正向或不翻转的图像，使检查者可以清晰观察视网膜细节和其他结构/介质（角膜、房水、晶状体及玻璃体），多为手持式，电池供电（也可由网电源供电），典型结构由一个安装在内部的光源、一面反光镜和一个拨盘组成，光源方向可由检查者进行改变，通过瞳孔照亮眼球内部，反光镜中间有一个圆孔，检查者可通过其进行观察，拨盘上有补偿透镜和辅助镜片，检查时可进行选择，用于检查眼底病变、屈光介质的异常及视网膜的定位等。产品性能指标采用下列参考标准中的适用部分，如：YY 1080-2009 眼科仪器 直接检眼镜。 | Ⅱ
740 | 16-04-09 | 间接检眼镜 | 间接检眼镜由一个聚光镜（手持式或整体式）来产生一个可被目视观察的中间实像，通常由一个用于照亮眼球内部的主光源、一个镜片及滤片系统，以及一个用于减少反光的照明光束组成，可分为手持式或头带式，可以是内部电源供电，也可由网电源供电；供眼内介质及眼底检查用。产品性能指标采用下列参考标准中的适用部分，如：YY0633-2008 眼科仪器 间接检眼镜。 | Ⅱ
741 | 16-04-12 | 角膜地形图仪 | 角膜地形图仪由数据采集单元，数据校正/处理单元及附件组成，可包括计算机及专用软件，可按工作原理、设计、技术参数、附加辅助功能等不同分为若干型号；通过非接触方式测量人眼角膜表面形状，获得角膜曲率及屈光力等参数用。 产品性能指标采用下列参考标准中的适用部分，如：YY 0787-2010眼科仪器 角膜地形图仪。 | Ⅱ
742 | 16-04-13 | 角膜测厚仪 | 通常由光学发射系统和光学测量系统组成。通过两反射光路的偏离方法或光程差方法测量角膜厚度。用于测量角膜厚度。 | Ⅱ |  | 新
743 | 16-04-17 | 眼压计 | 分类编码16-04-17下所有产品均免于进行临床评价。通常由角膜形状变化发生器、角膜变形测量系统或接触角膜装置和压变传感器组成。通过角膜形状变化（压平式、压陷式、非接触式等）或直接测量角膜血流脉动压力变化，换算获得眼内压。用于测量眼内压力。 | Ⅱ |  | 新
744 | 16-04-17 | 压陷式眼压计 | 压陷式眼压计通常由固定砝码、压针、锤弓和指针组成；可按型式（手持或台式）等不同分为不同型号；通过仪器手动向眼睛施加了一个外力，通过压针测量出眼角膜内陷的程度，并通过换算间接测量眼内压力。产品性能指标采用下列参考标准中的适用部分，如：YY 1036-2004压陷式眼压计。 | Ⅱ
745 | 16-04-17 | 压平眼压计 | 压平眼压计为光学压平式，由压重控制箱、双像棱镜、测量臂、测量旋钮、安装组件和校准组件组成，可按与裂隙灯显微镜的连接方式等不同分为若干型号；供人眼内压测量用。 | Ⅱ
746 | 16-04-17 | 回弹式眼压计 | 回弹式眼压计由探针、测量系统、显示系统组成，可按设计、技术参数、附加辅助功能等不同分为若干型号；供人眼内压测量用。 | Ⅱ
747 | 16-04-17 | 非接触式眼压计 | 非接触式眼压计由显示系统、喷气、测量光路系统及三维运动系统组成，可按设计、技术参数、附加辅助功能等不同分为若干型号；供人眼内压测量用。 | Ⅱ
748 | 16-04-19 | 干眼检测仪 | 通常由光源、CCD相机和控制系统组成，与计算机和配套软件组合使用。通过患者眼表泪膜干涉成像，用于诊断干眼程度。 | Ⅱ
749 | 16-04-21 | 眼科诊断辅助器具 | 16-04-21分类编码下所有Ⅱ类产品均免于进行临床评价。 | Ⅱ |  | 新
750 | 16-04-21 | 泪液分泌检测滤纸 | 泪液分泌检测滤纸由带有荧光素钠标示线的滤纸裁切而成，供临床用于各种泪液分泌障碍的检测 | Ⅱ
751 | 16-05-05 | 眼科手术显微镜 | 该产品一般由光学系统、控制系统、照明系统、支架系统构成,不包括具有特殊功能的选配件或部件,临床用于为眼科手术提供放大和照明。 | Ⅱ
752 | 16-05-06 | 眼内照明器光纤探头 | 通常由光纤和插入头组成。无菌提供。与眼科照明光源连接使用，可直接插入人眼，也可和套管针系统一起使用。用于眼内手术时传输照明光源发出的光，进行眼内照明。 | Ⅲ
753 | 16-05-06 | 玻切头 | 通常由玻切刀头、接头、导管和柄部等组成，刀头一般采用不锈钢材料制成。与气动设备配合使用，用于眼科手术时切除玻璃体。与已获准境内注册产品具有等同性，产品组成材料成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ |  | 新
754 | 16-05-06 | 一次性使用眼内激光光纤 | 由激光器连接头、光纤及手柄组成，与激光器连接后，用于向眼内传输激光。光纤出光端为直射平切端面，不含抽吸或照明等其他功能。 | Ⅲ |  | 新
755 | 16-05-06 | 接触式眼科观察镜片 | 通过放大或倒像等光学原理辅助眼科治疗、手术的光学器具（与角膜接触）。用于辅助眼科治疗和手术。 | Ⅱ
756 | 16-05-06 | 角膜板层刀片 | 通常为片状设计，带刃口。无菌提供，一次性使用。与角膜板层刀配合使用，用于剖层、切割眼角膜。 | Ⅱ
757 | 16-05-06 | 眼内照明器 | 通常由眼内照明器、眼内照明光纤和插入头（无菌提供，一次性使用）、可拆下和可消毒的旋钮组成。用于眼科手术期间对眼内的照明。 | Ⅲ
758 | 16-06-01 | 软性接触镜 | 日戴、单焦设计，矫正近视或远视。与已获准境内注册产品具有等同性，产品配方成熟。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
759 | 16-06-01 | 硬性接触镜 | 日戴、单焦设计，矫正近视或远视。与已获准境内注册产品具有等同性，产品配方成熟。新型结构设计、新型作用机理、药械组合产品、角膜塑形用硬性透气接触镜除外。 | Ⅲ
760 | 16-06-02 | 接触镜盐溶液 | 含量0.9％氯化钠的水溶液，用于软性或/和硬性接触镜的冲洗、储存，及片剂类护理产品的溶解稀释等。与已获准境内注册产品具有等同性。 | Ⅲ
761 | 16-06-02 | 接触镜护理产品 | 用于软性或硬性接触镜的清洁、消毒、冲洗、储存等。接触镜润滑液除外。不含有我国同类已上市产品中未含有的组成成分。产品有效成分的浓度不超过已上市产品有效成分的浓度。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ | 新增
762 | 16-06-03 | 医用光辐射防护眼镜 | 通常由镜架和镜片组成。镜片采用能反射或吸收辐射线，但能透过一定可见光的材料制成。用于在诊断或手术过程中防护紫外、蓝光、红光和红外危害。 | Ⅱ |  | 新
763 | 16-06-03 | 激光防护眼镜 | 一般采用玻璃或塑料镜片加入吸收剂制成。用于在诊断或手术过程中防止激光辐射对人眼的伤害。 | Ⅱ
764 |  | 义眼台 | 植入体内的球状或椭球状物。由羟基磷灰石材料制成。药械组合产品除外。适用于眼球、眶壁缺失、摘除或萎缩后的填充，眼眶内支撑。产品与已获准境内注册产品具有等同性。新型结构设计、新型作用机理、药械组合产品除外。 | Ⅲ
16-07-06
765 | 16-07-15 | 一次性使用折叠式人工晶状体推注器 | 一次性使用折叠式人工晶状体推注器可采用聚丙烯等适用材料制成，可由推进器、植入器两部分组成；可按材质、设计、技术参数等不同分为若干型号；以无菌形式提供；供人工晶状体植入术中推注折叠式人工晶体用。 | Ⅱ
766 | 17-00-00 | 固位纤维 | 固位纤维主要成分为高分子材料如聚对苯二甲酸乙二醇酯或玻璃纤维，可按材质、设计、物理性能、尺寸等不同分为若干种；供松牙固定、缺失牙齿的临时修复、牙列缺隙保持、固定桥的临时或永久加固和修复等用。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
767 | 17-00-00 | 义齿密合剂 | 义齿稳固剂可由山羧甲基纤维素钠和聚甲基乙烯醚钠盐、钙盐和缩苹果酸混合物及其他辅料等制成的膏状物，可按组方、技术参数等不同分为若干种；供义齿修复中，涂覆在义齿组织面与黏膜聒隙间。 | Ⅱ
768 | 17-00-00 | 一次性口腔麻醉导管 | 产品由软管、软管接头、手柄、套筒组成，可带或不带一次性使用无菌注射针。与口腔麻醉助推仪配合使用，适用于口腔局部麻醉时输送麻醉剂用。 | Ⅲ
769 | 17-01-01 | 牙周袋探测设备 | 通常由探针、手柄、脚踏开关和电源组成。通过使用压力敏感电子探针对牙周进行触诊，测量牙周袋深度。用于探测牙周袋深度。 | Ⅱ
770 | 17-01-06 | 口腔检查灯 | 通常由照明装置和检测观察装置组成。照明装置通常包括照明手柄、电源、患者护目镜；检测观察装置通常包括观察镜。用于口腔照明及检测观察，并且辅助增强口腔检查中粘膜异常和口腔病变的可视化程度。 | Ⅱ |  | 新
771 | 17-02-02 | 一次性使用无菌口镜 | 通常由柄、带有连接杆或不带有连接杆的镜子组成。镜面一般采用不锈钢或玻璃制成。无菌提供，一次性使用。 | Ⅱ
772 | 17-03-01 | 牙科综合治疗机 | 电驱动设备，供口腔诊断、治疗、手术用。可由牙科治疗机、牙科病人椅组成，牙科治疗机可由三用喷枪、观片灯、器械盘、吸唾组件、强吸组件、医用脚踏开关、地箱、净水系统等组成。产品性能指标采用下列参考标准中的适用部分，如：YY/T 1043.1 牙科治疗机、YY/T 0058 牙科病人椅等。 | Ⅱ
773 | 17-03-02 | 牙科病人椅 | 用于支撑患者保持坐姿，以便进行牙科检查、治疗和/或监测外科手术过程。高度通常可调，以使医护人员可站立操作。通常包括头枕和扶手，可躺式椅背可从垂直位置倾斜到水平位置，可具有旋转功能；牙科检查和/或治疗（如灯、冲洗装置）器械可作为手术椅的组件固定，也可以独立式壁式或悬吊式安装装置单独安放，上述装置可为电驱动设备。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0058 牙科病人椅。 | Ⅱ
774 | 17-03-02 | 牙科用椅 | 通常由底座支撑系统、控制系统、脚踏开关、椅面和头托组成。有源产品。用于牙科临床诊疗时承载患者（牙科椅）、医护人员（医师椅）。 | Ⅱ
775 | 17-03-03 | 超声洁牙机 | 超声洁牙机将超声能量通过手柄尖端作用于牙齿，用于牙齿表面、根管、牙周等部位的清洁；通常由超声波发生器、手柄、与手柄相配接的各种作用头（尖端）组成，可有冲洗部分；可按设计、技术参数、是否带工作头、附加辅助功能等不同分为若干型号；供临床去除牙垢用。产品性能指标采用下列参考标准中的适用部分，如： YY 0460-2009超声洁牙设备。 | Ⅱ
776 | 17-03-03 | 气动洁牙机 | 气动洁牙机由气动主机/手机、连接器、洁牙头、附件等组成，可有辅助照明、喷水等功能，可按设计、技术参数、辅助功能、预期用途等不同分为若干型号；可用于牙齿表面、根管、牙周等部位的清洁。 | Ⅱ
777 | 17-03-03 | 气动牙科喷砂装置 | 气动牙科喷砂装置由手柄 、连接装置等组成，可按材质、设计、技术参数等不同分为若干种；借助人工控制的高压气流及其内含的微小氧化铝细粒，去除牙体表面和牙周、根管等部位所粘着的牙垢、着色斑或抛光牙表面用。 | Ⅱ
778 | 17-03-04 | 牙科手机 | 牙科手机是一种用于向牙科工具或器具传递工作所需能量的手持件。产品涵盖符合YY1045.1 《牙科手机 第1部分：高速气涡轮手机》的要求的高速气涡轮手机，以及符合YY 1045.2《牙科手机 第2部分：直手机和弯手机》的要求的牙科弯手机或牙科直手机，以及抛光手机。 | Ⅱ
779 | 17-03-04 | 牙科低压电动马达 | 牙科低压电动马达是由主机控制或作为独立的移动装置。它通过马达电缆与主机完成电气相连并由后者供电及控制。牙科低压电动马达用于提供牙科手机进行牙科手术的驱动力。产品性能指标参照YY0836-2011制定。 | Ⅱ | 与原759合并。
780 | 17-03-04 | 牙科气动马达 | 牙科气动马达由牙科治疗机控制。牙科气动马达用于提供牙科手机进行牙科手术的驱动力。产品性能指标参照YY0837-2011制定。 | Ⅱ
781 | 17-03-04 | 牙科手机及附件 | 根据动力来源不同分为气动和电动。其中气动马达通过压缩空气推动叶片旋转产生动力，电动马达通过电磁原理产生动力。与牙科手机配套使用。用于驱动牙科手机。 | Ⅱ
782 | 17-03-06 | 牙科光固化机 | 通过产品发射出的蓝光照射修复材料使之固化，分为石英钨卤素灯为光源的光固化机和发光二极管（LED）灯为光源的LED光固化机两种。产品性能指标采用下列参考标准中的适用部分，如：YY 0055光固化机系列标准。 | Ⅱ
783 | 17-03-07 | 牙科种植机 | 牙科种植机通常由主机、马达和脚踏开关组成。用于牙科种植手术。 | Ⅱ | 新增
784 | 17-03-09 | 根管预备机 | 根管预备机为电力驱动式的根管手术设备。根据其技术结构不同，通常有两种组成结构情况。第一种由主机、马达电缆、马达、手机、脚踏开关等组成；第二种由主机、电池、充电器、手机等组成。组成中均不包含根管锉。根管预备机临床上用于根管治疗中，根管预备阶段成形和清理。产品不含牙根尖定位功能（或叫做根管长度测量功能）。 | Ⅱ
785 | 17-03-09 | 牙胶充填仪 | 牙胶充填仪为电力驱动式的根管手术设备。产品通常由主机、手柄（或手持件）、工作尖（推送牙胶的通道）、电源部分等组成。牙胶充填仪临床上使用于根管手术中的根管充填阶段，预期用途是加压软化牙胶，并将牙胶充填至制备完成的根管内。产品不含牙根尖定位功能（或叫做根管长度测量功能）。 | Ⅱ
786 | 17-03-09 | 牙根尖定位仪 | 通常由主机、唇钩、测量导线、探针和电源组成。用于牙科临床根管治疗时辅助确定工作长度。 | Ⅱ | 新增 | 新
787 | 17-03-10 | 口腔麻醉推注设备 | 通常由助推管和主机组成。通过设定程序控制注射的速度和流量，实现自动注射。用于口腔麻醉剂的注射。 | Ⅱ
788 | 17-03-12 | 牙齿研磨机 | 通常由研磨机和一次性无菌研磨容器组成。通常为电动。用于口腔科手术过程中将牙齿碎骨研磨成骨粉以供手术时使用。 | Ⅱ
789 | 17-04-01 | 软组织环切刀 | 与牙科手机配合使用，由不锈钢、钛合金等适用材料制成，或为无菌提供。用于牙科种植手术中牙龈的成形。产品使用原材料须符合相关国家和行业标准要求，产品性能指标需满足下列标准中的适用部分，如： YY91010《牙科旋转器械配合尺寸》等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。无菌提供。 | Ⅱ
790 | 17-04-06 | 机用根管器械 | 牙科根管治疗器械中的一种，与有源医疗器械配合使用。由柄和头部工作端组成，本体材质通常为不锈钢、镍钛合金等适用材质。用于牙科根管治疗中进行根管预备、辅助充填。可按柄和头部的材质、型式、尺寸等不同分为若干型号及规格。产品本体材料须符合相关国家和行业标准要求，性能指标需满足下列标准中的适用部分，如：YY 0803.1《牙科学 根管器械 第1部分：通用要求和试验方法》，YY/T 0803.2《牙科学 根管器械 第2部分：扩大器》等,且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
791 | 17-04-06 | 牙科锉 | 一般采用不锈钢、镍钛合金等金属材料制成的机用器械。工作端有刻纹或螺旋刃口等结构设计，起切削、平整的作用。该产品连接手机使用，由主机提供动力。配合有源器械使用，用于牙科治疗中对牙骨、根管进行切削、平整、清洁、塑形。 | Ⅱ
792 | 17-04-07 | 牙科车针 | 牙科车针是牙科旋转器械的一种，由柄（钢或适合的材料）和头部工作端（钢或硬质合金、金钢石等制成）组成，可按柄及头部的材质、型式、尺寸等不同分为若干型号及规格，用来切削牙体组织，以去除病变组织，治疗钻孔或制备牙体。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY 0761.1 牙科学 金刚石旋转器械 第1部分：尺寸、要求、标记和包装，YY 0302.1 牙科旋转器械 车针 第1部分：钢质和硬质合金车针。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
793 | 17-04-07 | 根管桩预备用牙钻 | 牙科旋转器械的一种，由柄和头部工作端组成，本体材质通常为不锈钢、钛合金等适用材质。用于口腔根管桩修复过程中，去除部分根管充填材料，制备与根管桩相匹配的根管形态。可按柄和头部的材质、型式、尺寸等不同分为若干型号及规格。产品本体材料须符合YY/T0294.1 《外科器械 金属材料 第1部分：不锈钢》等相关国家和行业标准要求，性能指标需满足下列标准中的适用部分，如：YY 91064 牙科旋转器械 钢和硬质合金牙钻技术条件，YY 91010 牙科旋转器械 配合尺寸等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
794 | 17-04-07 | 种植手术用牙钻 | 牙科旋转器械的一种，由柄和头部工作端组成，本体材质通常为不锈钢、钛合金等适用材质。用于口腔种植修复手术过程中，制备与使用型号种植体相匹配的种植窝洞。可按柄和头部的材质、型式、尺寸等不同分为若干型号及规格。产品本体材料须符合YY/T0294.1 《外科器械 金属材料 第1部分：不锈钢》等相关国家和行业标准要求，性能指标需满足下列标准中的适用部分，如：YY 91064《牙科旋转器械 钢和硬质合金牙钻技术条件》，YY91010《牙科旋转器械 配合尺寸》等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
795 | 17-04-07 | 种植体附件取出用牙钻 | 牙科旋转器械的一种，由柄和头部工作端组成，本体材质通常为不锈钢、钛合金等适用材质。用于种植手术过程中取出基台、种植体内破损螺丝及清理螺纹等。可按柄和头部的材质、型式、尺寸等不同分为若干型号及规格。产品本体材料须符合YY/T0294.1 《外科器械 金属材料 第1部分：不锈钢》等相关国家和行业标准要求，性能指标需满足下列标准中的适用部分，如：YY 91064《牙科旋转器械 钢和硬质合金牙钻技术条件》，YY91010《牙科旋转器械 配合尺寸》等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
796 | 17-04-07 | 牙科车针（非金属） | 牙科旋转器械中的一种，由树脂、聚醚醚酮（peek）等材料制成。由柄和头部工作端组成，可按柄及头部的材质、型式、尺寸等不同分为若干型号及规格，用来切削牙体组织，以去除病变组织，治疗钻孔或制备牙体。产品性能指标采用下列国家标准现行版本中的适用部分，如：YY 0761.1 牙科学 金刚石旋转器械 第1部分：尺寸、要求、标记和包装，YY 0302.1 牙科旋转器械 车针 第1部分：钢质和硬质合金车针。 | Ⅱ
797 | 17-04-07 | 正畸支抗配套用钻 | 牙科旋转器械的一种，由柄和头部工作端组成，本体材质通常为不锈钢、钛合金等适用材质。用于植入正畸支抗之前,在颌骨上钻孔以确定正畸支抗的植入位置。可按柄和头部的材质、型式、尺寸等不同分为若干型号及规格。产品本体材料须符合YY/T0294.1 《外科器械 金属材料 第1部分：不锈钢》等相关国家和行业标准要求，性能指标需满足下列标准中的适用部分，如：YY 91064 牙科旋转器械 钢和硬质合金牙钻技术条件，YY 91010 牙科旋转器械 配合尺寸等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
798 | 17-04-10 | 口内研磨抛光器械 | 牙科旋转器械的一种，与牙科手机配合使用。由金属柄和工作端组成，工作端为盘状、杯状等。用于对修复体的抛光、打磨以及多余部分的去除、种植体的清扫。也用于对牙齿表面的除垢、抛光。可按金属柄和工作端的材质、型式、尺寸等不同分为若干型号及规格。产品使用原材料须符合相关国家和行业标准要求，产品性能指标需满足下列标准中的适用部分，如： YY91010《牙科旋转器械配合尺寸》等，且基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
799 | 17-04-10 | 打磨抛光清洁器具 | 分为手动和机动两类。手动类产品通常由手持部分和工作端组成；机动类产品通常由金属柄和工作端组成，金属柄与牙科手机连接，其工作端为毛刷或杯状。用于对修复体的抛光、打磨以及多余部分的去除，或种植体的清扫。也用于对牙齿表面的除垢、抛光。不包括仅用于口腔科技工室的器具。 | Ⅱ
800 | 17-04-11 | 无菌上颌窦提升用球囊 | 本产品是由乳胶材料制成的球囊、封帽、导管、接头和一次性注射器组成，在牙科种植手术中使用，用于提升上颌窦。 | Ⅱ |  | 新
801 | 17-04-11 | 种植体安装辅助器械 | 通常为上颌窦提升用球囊等，在牙科种植过程使用。无源产品。无菌提供。用于牙科种植过程。 | Ⅱ
802 | 17-04-14 | 一次性使用牙科平头冲洗针 | 一次性使用牙科平头冲洗针采用高分子材料和高等级不锈钢材料制成，由针座（带标准鲁尔接头）和平头针管、护帽组成，按针管型式不同可分为直型和弯型两类；以无菌形式提供；与注射器配套使用，用于口腔科诊疗术后口腔压力冲洗用。 | Ⅱ
803 | 17-04-14 | 口腔清洗器具 | 对口腔进行冲洗的无源产品。无菌提供，一次性使用。用于去除口腔中的碎屑或杂物。 | Ⅱ
804 | 17-04-15 | 一次性使用无菌水枪头 | 一般采用不锈钢或塑料等材料制成。使用时安装在牙科综合治疗台的喷枪前端，为喷枪气流、液体的出口端。无菌提供，一次性使用。 | Ⅱ
805 | 17-04-15 | 一次性使用无菌吸唾管 | 一般采用不锈钢或塑料等材料制成。通常与牙科治疗机的抽吸装置一起使用。无菌提供，一次性使用。 | Ⅱ
806 | 17-04-15 | 喷枪 | 通常由枪体、手柄和喷杆组成。按使用时手持方式不同分为弯式和直式两种类型。根据需要可把压缩空气、水或气水喷雾以喷射的形式传送到口腔内某一部位。通常连接在牙科治疗机上使用。用于口腔清洗、吹干。 | Ⅱ
807 | 17-05-01 | 光固化正畸粘结用玻璃离子水门汀/双固化正畸用玻璃离子水门汀 | 一般为粉液状或糊状。作为临时性置入人体材料，用于正畸产品的粘结，短期粘结或临时填充。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
808 | 17-05-01 | 氧化锌丁香酚水门汀 | 一般为粉液状或糊状。粉剂为可析出离子的金属氧化物或金属盐，如氧化锌、玻璃粉、氢氧化钙等；液剂为酸溶液或螯合物等，如磷酸、聚羧酸、水杨酸、丁香酚；部分水门汀含有可聚合树脂成分。大部分为通过酸碱反应固化，含树脂成分的可通过化学反应或光固化反应固化。作为临时性置入人体材料，用于正畸产品的粘接、短期粘接或临时充填。 | Ⅲ
809 | 17-05-01 | 不含丁香酚的氧化锌水门汀 | 主要由氧化锌、矿物油等组成。用于牙科暂时修复和垫底。豁免临床的产品须至少符合YY 0272 牙科学 氧化锌/丁香酚水门汀和不含丁香酚的氧化锌水门汀的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
810 | 17-05-01 | 聚羧酸锌水门汀 | 经氧化锌与聚丙烯酸水溶液或聚链烯酸反应或经氧化锌/聚链烯酸粉与水反应而固化的水门汀。用于窝洞的垫底、衬层及粘固。豁免临床的产品须至少符合YY 0271.1 牙科水基水门汀 第1部分：粉/液酸碱水门汀和YY 0272 牙科学 氧化锌/丁香酚水门汀和不含丁香酚的氧化锌水门汀的适用部分。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
811 | 17-05-01 | 磷酸锌水门汀 | 经氧化物粉末（主要组成为氧化锌）与磷酸水溶液反应而固化的水门汀。用于窝洞的垫底、衬层及粘固。豁免临床的产品须至少符合YY 0271.1 牙科水基水门汀 第1部分：粉/液酸碱水门汀的相关要求。 | Ⅲ
812 | 17-05-01 | 玻璃离子水门汀 | 经硅铝酸盐玻璃粉与链烯酸水溶液反应，或经硅铝酸盐玻璃粉/聚酸粉与水或酒石酸水溶液反应而固化的水门汀。用于垫底、洞衬、粘固、乳牙和恒牙窝洞(Ⅲ类洞、V类洞、契状缺损)的修复充填。豁免临床的产品须至少符合YY 0271.1 牙科水基水门汀 第1部分：粉/液酸碱水门汀的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
813 | 17-05-01 | 氢氧化钙水门汀 | 主要成分为氢氧化钙和螯合剂。用于低强度垫底材料、衬层、间接衬层或盖髓剂。豁免临床的产品须至少符合YY 0271.1 牙科水基水门汀 第1部分：粉/液酸碱水门汀和YY 0272 牙科学 氧化锌/丁香酚水门汀和不含丁香酚的氧化锌水门汀的适用部分。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
814 | 17-05-01 | 桩核用树脂水门汀 | 一般由双酚A-二甲基丙烯酸缩水甘油酯（Bis-GMA）等组成。与根管桩配合使用，用于桩核修复时牙体组织的充填修复。性能、结构、组成等属于YY 1042《牙科学聚合物基修复材料》可完全涵盖的树脂水门汀，且基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括与已上市产品相比使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
815 | 17-05-01 | 粘固用树脂水门汀 | 一般由双酚A-二甲基丙烯酸缩水甘油酯（Bis-GMA）等组成。用于冠、桥等修复体的粘固或固位。性能、结构、组成等属于YY 1042《牙科学聚合物基修复材料》可完全涵盖的聚合物基粘固材料，且基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
816 | 17-05-02 | 正畸粘接剂 | 主要由树脂、填料、溶剂等组成。用于口腔正畸治疗的粘接处理。 | Ⅱ
817 | 17-05-02 | 牙科粘接剂 | 为粘稠液体性状，主要由甲基丙烯酸酯等制成。用于修复体或修复材料与牙齿粘接、修复材料之间粘接使用。性能、结构、组成、用途等属于YY/T 0518《牙科修复体用聚合物基粘接剂》可完全涵盖的牙科粘接剂，且基本原理、适用范围、性能和主要组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
818 | 17-05-03 | 牙胶尖 | 主要由古塔胶、氧化锌、硫酸钡等组成。属于牙科固体根管充填材料。用于根管治疗过程中与根管糊剂配合使用进行根管充填。豁免临床的产品须至少符合YY 0495 牙根管充填尖的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
819 | 17-05-06 | 银汞合金 | 粉液双组份或胶囊型。液剂为汞，粉剂为银合金。通过汞齐化反应生成银汞合金。用于牙体缺损的直接充填修复。豁免临床的产品符合YY 1026 牙科学 汞及银合金粉的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
820 | 17-05-07 | 临时充填材料 | 固体、单组份或双组份糊剂。固体通常为古塔胶；糊剂通常由树脂基质、经过表面处理的无机填料、引发体系等组成，或由硫酸钙/锌、氧化锌、树脂基质等组成。用于牙体缺损的临时修复。 | Ⅱ | 新增 | 新
821 | 17-05-08 | 光固化氢氧化钙间接盖髓剂 | 光固化氢氧化钙间接盖髓剂是一种光固化阻射性氢氧化钙间接盖髓剂，主要组成成分通常为氢氧化钙、氨基甲酸酯双甲基丙烯酸酯（UDMA）、三乙二醇二甲基丙烯酸酯（TEDMA）等。产品不用于直接盖髓，作为间接盖髓剂，为其他充填材料垫底。豁免情况不包括与已上市产品相比采用新材料、新技术或具有新作用机理、新功能的产品。 | Ⅲ |  | 新
822 | 17-06-00 | 牙科增材制造用金属材料 | 牙科增材制造用纯钛、钛合金、钴铬合金金属材料，用于通过激光选区熔化工艺加工制作牙冠、牙桥等非植入修复体，产品本体材料及性能指标须至少满足YY∕T 1702-2020 《牙科学 增材制造 口腔固定和活动修复用激光选区熔化金属材料》、GB 17168 《牙科学 固定和活动修复用金属材料》等相关国家和行业标准要求。豁免情况不包括用于制作颌面赝复体、种植体、基台及附件的产品及与已上市产品相比采用新材料、新技术或具有新作用机理、新功能的产品。 | Ⅲ |  | 新
823 | 17-06-01 | 铸造合金 | 定制式义齿制作用材料，应符合GB 17168 牙科学 固定和活动修复用金属材料标准 和YY 0621.1 牙科学 匹配性实验 第1部分：金陶-陶瓷体系。贵金属应符合YY 0620. 牙科学 铸造金合金相关标准要求；用于制作嵌体、支架、牙冠、桥、基托、卡环、金属烤瓷修复体等的镍铬基合金、钴铬合金、纯钛与钛合金等，不可用于基台的定制。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
824 | 17-06-01 | 牙科切削用钛及钛合金 | 牙科切削用钛和钛合金金属材料，用于通过CAD/CAM加工制作牙冠、牙桥等修复体，产品本体材料及性能指标须满足相关国家和行业标准要求，如：GB17168 《牙科学 固定和活动修复用金属材料》、GB/T 13810《 外科植入物用钛及钛合金加工材》的适用部分。豁免情况不包括用于制作种植体及附件的产品。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
825 | 17-06-01 | 牙科附着体 | 通常由阴型和阳型两部分部件组成，按结构组成分为栓道式，杆卡式，按扣式和球帽式等类型。用于可摘局部义齿、覆盖义齿等修复体的辅助固位。产品使用原材料须符合相关国家和行业标准要求。产品基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括使用了新材料、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
826 | 17-06-01 | 牙科切削用合金 | 牙科切削用钴铬合金、镍铬合金。用于通过CAD/CAM加工制作牙冠、牙桥等修复体，产品本体材料及性能指标须满足相关国家和行业标准要求，如：GB17168 《牙科学 固定和活动修复用金属材料》的适用部分。豁免情况不包括用于制作种植体及附件的产品，不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
827 | 17-06-02 | 定制式义齿用烤瓷粉 | 定制式义齿制作用材料，I型粉状陶瓷，可按组分、预期用途等不同分为若干类，可分别供定制式义齿的牙核、牙本质瓷/牙体、牙釉质、牙颈部、染色、透明、添加、上釉用。产品性能指标采用下列参考标准中的适用部分，如：YY 0716-2009 牙科陶瓷。 | Ⅱ
828 | 17-06-02 | 全瓷义齿用氧化锆瓷块 | 全瓷义齿用氧化锆瓷块主要成分为：氧化锆、氧化钇、氧化铝及其他氧化物，其中，氧化锆（含氧化铪、氧化钇）（ZrO2+HfO2+Y2O3）材料成分（重量百分比）应大于99%。可按成分、成型工艺、形状等不同分为若干型号、规格；供采用CAD/CAM(计算机辅助设计/计算机辅助制造)方法切削制作牙科固定义齿，如冠、桥、嵌体、贴面。产品性能指标采用下列国家标准现行版本中的适用部分，如：GB30367 牙科学 陶瓷材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
829 | 17-06-02 | 牙科陶瓷修复用全瓷材料 | 定制式义齿制作用材料，I型粉状陶瓷，可按组分、预期用途等不同分为若干类，可分别供定制式义齿的牙核、牙本质瓷/牙体、牙釉质、牙颈部、染色、透明、添加、上釉用。用于全瓷单冠、长桥和嵌体的制作及全瓷修复。产品性能指标采用下列国家标准现行版本中的适用部分，如：GB30367 牙科学 陶瓷材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
830 | 17-06-02 | 全瓷义齿用玻璃陶瓷 | 定制式义齿制作用材料，通常为块状，主要成分为二氧化硅、五氧化二磷、三氧化二铝、氧化锂、氧化钾，氧化锆及其他氧化物等。用于全瓷贴面、嵌体、冠、桥的制作。一般采用粉浆涂塑、计算机辅助设计/制造（CAD/CAM）、热压铸等工艺制成。豁免临床的产品须至少符合GB 30367 牙科学 陶瓷材料的相关要求。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
831 | 17-06-02 | 全瓷义齿用染色液 | 通常为液剂，主要成分为聚乙二醇、氯化氢、氯化铁、氯化铒等。用于义齿制作时陶瓷材料的染色。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
832 | 17-06-03 | 义齿稳固剂 | 主要由聚甲基乙烯醚钠盐、羧甲基纤维素钠等组成。用于涂覆于义齿组织面，使义齿与牙龈紧密贴合，维持义齿稳固。产品基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
833 | 17-06-03 | 临时冠桥材料 | 产品为双组份糊剂或粉液剂，或单组份糊剂，或预成制品。主要成分为聚丙烯酸酯类、聚乙烯基酯类、聚碳酸酯类等。用于制作牙科临时修复体。产品性能指标采用下列参考标准中的适用部分，如：YY/T 0826牙科临时聚合物基冠桥材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
834 | 17-06-03 | 合成树脂牙 | 主要由丙烯酸酯类聚合物制作而成，适用于局部义齿和全口义齿的制作。产品本体材料及性能指标须满足相关国家和行业标准要求，如：YY 0300 牙科学 修复用人工牙的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
835 | 17-06-03 | 义齿基托聚合物 | 一般由双组份糊剂或粉液剂、粉剂，或单组份糊剂组成，主要由聚丙烯酸酯类、聚酰胺类等材质组成。用于义齿基托的制作。产品性能指标须至少符合YY 0270.1《牙科学基托聚合物第1部分：义齿基托聚合物》的相关要求，且基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ | 修订产品描述。
836 | 17-06-03 | 聚合物基冠桥材料 | 双组份糊剂或粉液剂、粉剂，或单组份糊剂，工作原理为引发剂和催化剂的混合而固化（自凝）。主要成分为丙烯酸酯类。用于制作与金属表面接触或不接触的前牙冠，不用于牙体缺损的修复，且不用于承力区的后牙修复。基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免临床的产品须至少符合YY/T 0710《牙科学 聚合物基修复材料》的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
837 | 17-06-03 | 造牙粉及造牙水 | 主要由甲基丙烯酸甲酯（MMA）的均聚物或丙烯酸酯类的共聚物以及甲基丙烯酸甲酯单体制成，用于义齿的制作。豁免临床的产品须至少符合YY 0270.1 牙科学 基托聚合物 第1部分：义齿基托聚合物的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
838 | 17-06-04 | 定制式义齿 | 一般采用钴铬合金、镍铬合金、纯钛、钛合金、贵金属合金、瓷块、瓷粉、基托树脂、合成树脂牙等材料制成，根据需要而定。可以是固定或活动的，如冠、桥、嵌体、贴面、桩核、可摘局部或全口义齿等。制作过程中所使用的医疗器械材料全部为具有注册证的材料。 | Ⅱ
839 | 17-06-04 | 定制式固定义齿 | 定制式固定义齿主要由金属、陶瓷、高分子材料制成，用于牙体牙列缺损缺失的固定修复，包括冠、桥、贴面、嵌体等。制作过程中所使用的义齿材料均为具有注册证的材料。豁免情况不包括定制式种植体、基台及附件，及与已上市产品相比采用新材料、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
840 | 17-06-04 | 定制式活动义齿 | 定制式活动义齿包括局部义齿和全口义齿，主要由金属、陶瓷、高分子材料制成，用于牙体牙列缺损缺失的活动修复。制作过程中所使用的义齿材料均为具有注册证的材料。豁免情况不包括定制式种植体、基台及附件，及与已上市产品相比采用新材料、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
841 | 17-06-05 | 金属桩 | 采用金属材料制成。固定于根管或牙冠内以保证修复体的固位，防止冠或充填修复体因无足够的固位而折断或脱落，用于根管治疗后牙体大面积缺损时对重建的冠核进行固位，参见导则。常与树脂或桩核材料共同使用。豁免临床的产品须至少符合YY/T 0517 牙科预成根管桩的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
842 | 17-06-05 | 固位钉 | 一般采用金属材料制成。常与树脂等材料配合使用，用于辅助修复牙体缺损，为修复牙齿提供固位力。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅲ
843 | 17-06-05 | 纤维桩 | 采用环氧树脂等高分子制成，为具有标准尺寸的非定制产品。用于根管治疗后牙体大面积缺损时对重建的冠核进行可靠固位参见导则。豁免临床的产品须至少符合YY/T 0517 牙科预成根管桩的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
844 | 17-07-01 | 正畸托槽 | 正畸托槽可采用不锈钢、陶瓷适用材料制成，可按材质、牙位、技术参数、使用部位等不同分为若干型号与规格；与正畸丝等配套，供牙齿正畸治疗用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY/T 0915 正畸用托槽及颊面管。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
845 | 17-07-02 | 正畸丝 | 正畸丝采用不锈钢、镍钛合金、钛合金、钛钼合金、铜镍钛合金制成，可按材质、弹性、截面形状（圆形和矩形）和尺寸等不同分为若干型号及规格，与正畸托槽等配套，供口腔正畸用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY/T 0625 牙科学 正畸产品：正畸丝。 | Ⅱ
846 | 17-07-02 | 正畸结扎丝 | 正畸结扎丝可由不锈钢等适用材料制成，可按材质、物理性能、丝径/长度等不同分为若干种，主要用于口腔科牙列正畸时将正畸丝与托槽或其他附件相结扎，以达到固定及牵引牙齿移动的目的。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
847 | 17-07-03 | 正畸颊面管 | 一般采用金属材料制成。颊面管为管状，由主体和拉钩组成，可按材质、设计、技术参数等不同分为若干型号及规格；与正畸托槽、正畸丝等配套，正畸治疗中用于固定正畸丝，也可传递矫治力量。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY/T 0915 正畸用托槽及颊面管。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
848 | 17-07-03 | 正畸带环 | 由不锈钢、钛合金等适用材料制成，可按材质、物理性能、尺寸等不同分为若干种。用于正畸治疗过程为正畸丝,扩弓器,弹性体或其他附件提供支撑点,从而对牙齿进行矫正。产品使用原材料须符合相关国家和行业标准要求。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
849 | 17-07-04 | 正畸基托聚合物 | 一般由粉剂和液剂组成。主要由甲基丙烯酸甲酯（MMA）的均聚物或丙烯酸酯类的共聚物以及甲基丙烯酸甲酯单体制成，适用于正畸基托的制作。豁免临床的产品须至少符合YY 0270.2-2011 牙科学 基托聚合物 第2部分：正畸基托聚合物的相关要求。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
850 | 17-07-06 | 正畸弹性体附件 | 正畸弹性体附件包括正畸弹性牵引圈、正畸弹力线/弹力管、正畸橡皮链、正畸结扎圏、正畸分牙圏、旋转橡皮垫，用天然或合成橡胶制成，可按材质、形状、尺寸等不同分为若干型号及规格；在口内与口外与固定或活动矫治器起协同矫治作用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY/T 0624 牙科学 正畸产品：正畸弹性体附件。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
851 | 17-07-07 | 游离牵引钩 | 游离牵引钩可由不锈钢等适用材料制成，可按材质、物理性能、尺寸等不同分为若干种，与其他正畸产品配套，供口腔正畸用。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
852 | 17-07-07 | 螺旋扩弓器 | 螺旋扩弓器用不锈钢等适用材料制成，可由丝杆、螺母、导向杆、调节杆组成，可按材质、设计、技术参数、适用部位等不同分为若干种，供牙弓狭窄者扩大牙弓间隙等用。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
853 | 17-07-07 | 颌前方牵引口外装置 | 颌前方牵引口外装置由前额托架、可调式支架、调节器和下颌兜等组成，前额托架和下颌兜采用聚丙烯树脂制成，可调式支架采用不锈钢制成；可按材质、设计、技术参数等不同分为若干种；配合口内正畸装置，供人体上颌矫正时辅助牵引用。不包括口内正畸装置。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
854 | 17-07-07 | 磨牙推进器 | 磨牙推进器用不锈钢等适用金属材料制成，由各类支杆或支杆和弹簧组成，可按材质、设计、技术参数等不同分为若干种；与托槽、带环、种植钉等配套，用于推动磨牙，获得错牙排齐所需间隙。不包括种植钉、托槽、带环。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
855 | 17-07-07 | 舌侧扣 | 通常由不锈钢、钛合金等材料制成。用于口腔正畸治疗，与正畸托槽、正畸颊面管、正畸丝等配套使用。产品使用原材料须符合相关国家和行业标准要求，产品基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品相比使用了新材料、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
856 | 17-07-07 | 矫治保持器 | 产品由医用级聚氯乙烯材料、聚对苯二甲酸乙二醇酯（PET）、热塑性聚氨酯或医用硅胶制成，为具有标准尺寸的非定制产品。用于巩固牙颌畸形矫治完成后的疗效。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ | 修订产品描述。
857 | 17-08 | 定制式牙科种植 | 通常由基板和导环组成。用于牙科种植手术中种植体的定位、导向等，从而辅助安装种植体。基板通常采用环氧树脂或丙烯酸酯材料，根据CAD模型，通过3D打印技术制成。导环采用纯钛钛合金等金属材料制成。产品本体材料及性能指标须满足相关国家和行业标准要求，如：GB/T 3620.1《钛及钛合金牌号和化学成分》、GB/T 13810 《外科植入物用钛及钛合金加工材》、YY/T0911《牙科学 聚合物基代型材料》的适用部分。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ | 修订产品描述。
用导板
858 | 17-08-02 | 保护帽 | 种植基台用保护帽由聚甲基丙烯酸甲酯和塑料聚甲醛制成，帽状，用于口腔科种植手术后的临时修复过程，保护种植体的上部结构或基台。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
859 | 17-08-02 | 直形基台、愈合基台、临时基台及附件 | 直形基台、愈合基台由钛和钛合金材料制成，也包括临时基台。附件包括中央螺钉、基台螺丝、封闭螺丝、覆盖螺钉、愈合帽、套筒及螺丝、基台帽。为牙缺失后颌骨内植入的牙种植体的配套用部件，安装在锚固于骨内的种植体平台上，用于连接、支持和固位修复体或种植体上部结构。对于钛和钛合金材料制成的免于临床评价产品须至少符合 YY/T 0520《钛及钛合金材质牙种植体附件》的相关要求，钛和钛合金材料本体材料须符合GB/T 3620.1《钛及钛合金牌号和化学成分》、GB/T 13810《外科植入物用钛及钛合金加工材》的相关要求。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅲ | 修订产品名称及产品描述。
860 | 17-08-04 | 种植体密封胶 | 单组份糊剂或其他形式提供。通常由聚硅氧烷和无水硅酸组成，也可由其他成分组成。用于封闭种植体与中央螺钉之间的空隙。防止致病菌渗入种植体内部引发其周围组织二次感染。基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
861 | 17-08-05 | 转移帽/印模帽 | 指无菌提供的转移帽或印模帽，为牙科种植修复用辅助器械的一种，采用钛合金等适用材质制成。用于种植修复时将种植体和基台在口腔的位置转移到工作模型上。产品使用原材料须符合相关国家和行业标准要求。产品基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括使用了新材料、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
862 | 17-09 | 牙齿干燥剂 | 牙齿干燥剂由乙醇、丙酮、蒸馏水等组成；敷于待诊疗牙齿表面，使其快速清洁、干燥，以利进一步的诊疗操作。 | Ⅱ
863 | 17-09-01 | 氢氧化钙根管消毒剂 | 氢氧化钙根管消毒剂为单组分或双组分糊剂，所含成分不具有药理学作用；单组分由氢氧化钙与赋型剂组成，双组分由氢氧化钙和丙二醇水溶液组成；可按组方、技术参数及附件等不同分为若干种；仅供口腔内短期使用，用于根管消毒。 豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
864 | 17-09-01 | 根管扩大液/根管清洗液/根管润滑剂 | 通常由乙二胺四乙酸/依地酸(EDTA)，羧甲基纤维素等组成。用于根管预备和扩大时润滑、清洁、冲洗根管。产品基本原理、适用范围、性能和主要组成成分与已经上市产品相同。豁免情况不包括与已上市产品相比使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
865 | 17-09-02 | 一次性使用无菌牙科吸潮纸尖 | 通常为纸质或纯棉纤维质的锥形尖。具有良好的吸水性、硬且有韧性、容易放进牙根管内。无菌提供，一次性使用。 | Ⅱ
866 | 17-09-03 | 牙科用酸蚀剂 | 通常由单组份或双组份液体或凝胶。一般为磷酸、乳酸、柠檬酸、氢氟酸、草酸、聚丙烯酸等。利用酸的腐蚀性发挥作用。用于口内修复或正畸治疗时，利用酸蚀剂的腐蚀性对牙体、金属、陶瓷等修复体表面进行处理，以去除污染层、粗糙表面、提高其表面性能。作用机理及主要组成成分须与已上市产品相同。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY 0769 牙科用磷酸酸蚀剂。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
867 | 17-09-04 | 硅烷偶联剂 | 单组份液体。通常为亲水基团的丙烯酸酯功能单体、硅烷偶联剂或其他成分。用于牙齿、树脂、陶瓷、金属修复体等的表面处理。利用其化学改性作用改变牙齿、修复体表面性状。 产品基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
868 | 17-09-04 | 预处理剂 | 预处理剂通常为亲水基团的丙烯酸酯功能单体、硅烷偶联剂或其他成分。用于牙齿、树脂、陶瓷、金属修复体等的表面处理，利用其化学改性作用改变牙齿、修复体表面性状，以提高修复体之间或者正畸治疗中的粘结力。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。预处理剂、正畸预处理剂、通常为亲水基团的丙烯酸酯功能单体、硅烷偶联剂或其他成分。用于牙齿、树脂、陶瓷、金属修复体等的表面处理，利用其化学改性作用改变牙齿、修复体表面性状，以提高修复体之间粘结力。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
869 | 17-09-05 | 排龈线 | 不含血管收缩或止血功能的棉线，可按材质、线径、长度等不同分为若干种，用于在牙体预备、取印模或粘固牙冠时排开牙龈。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
870 | 17-09-05 | 排龈膏 | 不含血管收缩或止血功能的膏状材料。一般由氯化铝、高岭土和辅料等组成，用于在牙体预备、取印模或粘固牙冠时排开牙龈。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
871 | 17-09-05 | 排龈线（带硫酸铝/氯化铝的排龈线） | 通常为辫子结构的棉线等适用材料制成，含有硫酸铝或氯化铝等成分。用于牙体预备、制取印模或粘冠等操作前减少牙龈损伤排开牙龈使用。基本原理、适用范围、性能和组成结构与已经上市产品相同。豁免情况不包括与已上市产品相比使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅲ
872 | 17-09-06 | 口内研磨抛光材料 | 口内研磨抛光材料主要由二氧化硅或碳酸氢钠及辅料等制成，一般为粉剂、糊剂，也可为其他形式，可按组方、磨料颗粒大小、装量、附件等不同分为若干种；用于在口腔内研磨抛光牙体组织或修复体，使其表面平滑均匀。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
873 | 17-09-06 | 洁牙粉/喷砂粉 | 与牙科专用喷砂机配套使用, 通常由碳酸氢钠、二氧化硅、氢氧化铝、碳酸钙、氟化钙、硫酸氢钠等组成。用于祛除牙齿上的菌斑、色素及牙石。产品基本原理、适用范围、性能和组成成分与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
874 | 17-09-07 | 琼脂基水胶体印模材料 | 琼脂基水胶体印模材料为可逆性凝胶材料，可按稠度等技术参数不同分为若干型号和规格，供牙齿和口腔组织印模用。若用于技工室复制模型，可重复使用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY 0494牙科琼脂基水胶体印模材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
875 | 17-09-07 | 弹性体印模材料 | 弹性体印模材料是经塑性后，印模具有弹性的材料，如硅橡胶、聚醚橡胶，通常由基质、交联剂、催化剂和填料组成，可按材料、配方、印模技术参数（如弹性、化学稳定性、凝固时间等）等不同分为若干型号和规格，经聚合反应后，对牙齿和口腔组织取模用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY 0493 牙科学 弹性体印模材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
876 | 17-09-07 | 齿科藻酸盐印模材料 | 齿科藻酸盐印模材料为粉末状，外观均匀无杂质，调和后的材料应是均匀，不结团，不成粒，且表面光滑，能形成光滑的可塑体，一次性使用产品，无口腔粘膜刺激性，无致敏毒性，主要组成成分为海藻酸盐、硫酸钙、硅藻土、滑石粉等，可按组方、调和时间、总工作时间、固化时间等不同分为若干规格，供制取牙齿和口腔组织印模用。产品性能指标采用下列行业标准现行版本中的适用部分，如：YY 1027 齿科藻酸盐印模材料。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
877 | 17-09-07 | 牙科光学喷粉 | 单纯的粉剂通常由乙醇、氟化烃、薄荷香料组成，喷涂在病人牙齿上，起辅助成像作用。用于制作记录口腔各组织形态及关系的阴模，或者辅助获取清晰的牙齿3D图像。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
878 | 17-09-07 | 牙科光学显影膏 | 产品通常由聚二甲基硅氧烷，硫酸钡，二氧化硅等组成。通过涂布的义齿或非阻射性组件，在口腔扫描仪（CT或DVT）下以便获取牙齿的3D图像。用于义齿导向种植规划和牙科诊断。 | Ⅱ
879 | 17-09-12 | 硅橡胶咬合记录材料 | 硅橡胶咬合记录材料主要成分为二氧化硅、聚二甲基硅氧烷，用于口腔修复治疗中，记录上下牙列咬合关系。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ
880 | 17-09-12 | 义齿贴合点指示剂 | 通常由基质、催化剂和阻聚剂组成，主要成份为乙烯基硅橡胶、二氧化硅等。用于涂布于牙科修复体组织面，经咬合后通过固化、变色、变形等方式对修复体不良接触部位等进行检测。产品性能指标等须符合相关国家和行业标准要求，豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
881 | 17-09-12 | 咬合蜡 | 双组份糊剂或粉液剂或片。通常由硅橡胶或软质塑料等材料组成。所含成分不具有药理学作用，所含成分不可被人体吸收。咬合检查材料经咬合后通过固化、变色、变形或对施力部分进行染色起到指示作用。用于口腔修复治疗中，记录上下牙列咬合关系，并根据记录结果，制作口腔各组织形态及关系的阴模，或者辅助获取清晰的牙齿3D图像。 | Ⅱ
882 | 17-10-03 | 脱敏剂 | 产品通常为单组份或双组份液剂、糊剂、凝胶状。主要成分为氟化物等。用于缓解因牙本质暴露而引起的牙齿过敏症状。主要组成成分、基本原理、性能需与已经上市产品相同。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
883 | 17-10-04 | 氟防龋材料 | 一般采用含氟材料制成。用于预防龋齿。产品性能指标采用下列国家标准现行版本中的适用部分，如：YY/T 0823《牙科氟化物防龋材料》。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
884 | 17-10-04 | 窝沟封闭剂 | 一般由树脂基材料制成，用于封闭牙齿窝沟点隙，预防牙齿窝沟点隙处龋齿的形成。基本原理、性能和主要组成成分须与已经上市产品相同，产品性能指标采用下列国家标准现行版本中的适用部分，如：YY 0622-2008《牙科树脂基窝沟封闭剂》。豁免情况不包括使用了新材料、活性成分、新技术、新设计或具有新作用机理、新功能的产品。 | Ⅱ
885 | 17-10-05 | 牙科膜片 | 牙科膜片可采用聚乙烯、醋酸乙烯酯、可食用色素等适用材料制成，单层片状结构或多层片状结构，可按材质、形状、尺寸等不同分为若干种，用于制作磨牙垫、运动护齿（牙合）垫、保持器等。豁免情况不包括与已上市产品材料、技术、设计或作用机理、功能等不同的产品。 | Ⅱ | 修订产品描述。
886 | 18-01-02 | 一次性使用脐带剪夹器 | 一次性使用脐带剪夹器通常由高分子材料（如聚乙烯等）和金属材料（如用奥氏体不锈钢制作刀片）制成，可由脐带夹剪器、留置脐带夹及附属部件等构成，以无菌形式提供，供切断并封闭新生儿脐带残端用。 | Ⅱ
887 | 18-01-03 | 一次性使用子宫颈钳 | 高分子塑料制品，由雌雄片钳子及鳃轴组成，用于牵拉子宫颈。无菌，一次性使用。 | Ⅱ
888 | 18-01-04 | 一次性使用无菌脐带夹 | 一次性使用无菌脐带夹采用高分子材料制成，可按材质、设计、技术参数等不同分为若干型号，以无菌形式提供；供新生儿闭合脐带残端用。不包括脐带剪。 | Ⅱ
889 | 18-01-05 | 一次性使用无菌阴道扩张器 | 一次性使用无菌阴道扩张器可采用聚丙烯等适用高分子材料制成，主要由上叶、下叶和手柄组成，可按材质、型式、尺寸不同分为若干型号和规格；产品以无菌形式提供；供妇产科作阴道诊察用。产品性能指标采用下列参考标准中的适用部分，如：YY 0336 一次性使用无菌阴道扩张器。 | Ⅱ
890 | 18-01-05 | 一次性使用宫颈扩张棒 | 一次性使用宫颈扩张棒一般由高分子材料制成，可根据外径不同分为若干规格，产品以无菌形式提供。供临床妇产科宫腔诊查前的子宫颈扩张用。 | Ⅱ
891 | 18-01-05 | 一次性使用子宫颈扩张球囊导管 | 产品由导管、球囊、充盈头组成，导管、球囊、充盈头材质为高分子材料。该产品为无菌产品，一次性使用。该产品主要用于机械扩张子宫颈。 | Ⅱ
892 | 18-01-06 | 一次性使用胎头吸引器 | 一次性使用胎头吸引器采用聚丙烯等适用高分子材料制成，由手柄和吸盘组成；可按材质、设计、尺寸等不同分为若干型号和规格；可以无菌形式提供；当产妇因种种原因不能正常分娩时，通过吸盘对胎头头皮产生的负压吸引以辅助经阴道或剖腹产分娩。不包括气动负压泵式胎头吸引器。 | Ⅱ
893 | 18-01-07 | 一次性使用妇科冲洗治疗头 | 一次性使用妇科冲洗治疗头采用塑料经注塑而成，由管身和接头组成；可按材质、技术参数（如长度、管径等）等不同分为若干型号和规格；产品以无菌形式提供；连接于医用臭氧冲洗治疗仪的治疗或冲洗手柄上，供建立阴道内输送臭氧气体或液体的通道用。 | Ⅱ
894 | 18-01-07 | 一次性使用阴道给药器 | 一次性使用阴道给药器采用聚丙烯等适用材料制成，可由药筒（可有刻度）和推杆组成；产品通常以无菌形式提供；可按材质、设计、容量等不同分为若干型号及规格；供女性阴道给药用。 | Ⅱ
895 | 18-01-08 | 一次性使用子宫刮匙 | 通常由刮匙头、刮匙颈与刮匙柄组成，一般由高分子材料制成。无菌提供。 用于刮、擦方式提取或除去子宫内物质。 | Ⅱ
896 | 18-01-09 | 宫腔镜举宫器 | 产品为宫腔镜附件，可由摆头、摆头座、闭塞器、连杆、挂钩、手柄及注射管等部件组成。用于宫腔镜或宫/腹腔镜联合下手术时拨动子宫位置用。无菌提供。 | Ⅱ
897 | 18-01-09 | 子宫操纵器 | 该产品通常由不锈钢和高分子材料制成的棒状器械。无菌提供。在腹腔手术中作变动子宫体位用，也可用于控制、提升子宫。 | Ⅱ
898 | 18-01-09 | 子宫定位手柄 | 产品由把手、扳机、轴杆和卡夹式接头组成。用于腹腔镜手术中对子宫的瞄准以及对子宫、输卵管和卵巢的定位，并在需要染色输卵管灌注法的手术中用于给付染料。 | Ⅱ
899 | 18-01-10 | 输卵管通液诊断仪 | 输卵管通液诊断仪主要由控制系统、电源系统、助推系统、打印机等组成，可按设计、技术参数、附加辅助功能等不同分为若干型号；供输卵管造影注液诊断和输卵管疏通用。 | Ⅱ
900 | 18-01-10 | 一次性使用宫腔通液造影管 | 一次性使用宫腔通液造影管可采用硅橡胶等高分子材料制成，主要由气囊、导管、充气接头、通液接头组成；可按材质、管腔外径等不同分为若干型号及规格；产品以无菌形式提供，供临床人体宫腔通液造影用。 | Ⅱ
901 | 18-01-10 | 一次性使用输卵管造影导管 | 一次性使用输卵管造影导管可采用塑料、不锈钢丝等组成，由弯头引导管、牵引杆、内导管、导丝组成；也可用硅胶和塑料等制成，由导管（单腔或双腔）、导管座（可含阀门和三通）、球囊组成；导管可带不透X射线标记；可按材质、设计、技术参数（如管径、球囊压力）等不同分为若干型号和规格；产品以无菌形式提供；通过宫腔镜操作通道或其他子宫介入器械通道插入输卵管，注入色素液或造影剂来评估输卵管的通畅性，并进行注液疏通操作。 | Ⅱ
902 | 18-02-01 | 超声多普勒胎儿监护仪 | 超声多普勒胎儿监护仪利用超声多普勒原理，在围产期对胎儿进行连续监护，并在出现异常时及时提供报警信息的超声仪器。通常由超声探头（一般采用梅花式探头）、宫缩压力传感器及与之相连的主机组成；可按机型、探头配置、技术参数、附加辅助功能、软件处理、预期用途等不同分为若干型号；主要用于围产期胎儿心率和孕妇宫缩压力等的连续监护。 | Ⅱ
903 | 18-02-02 | 超声多普勒胎儿心率仪 | 超声多普勒胎儿心率仪根据多普勒原理从孕妇腹部获取胎心运动信息的超声仪器。通常由探头（一般采用单元探头）、超声波发射/接收电路、信号输出部分组成。可按机型、探头配置、技术参数、附加辅助功能、软件处理、预期用途等不同分为若干型号；供胎儿心率测量用。产品性能指标采用下列参考标准中的适用部分，如：YY 0448-2009超声多普勒胎儿心率仪。 | Ⅱ
904 | 18-02-03 | 一次性使用子宫探针 | 一次性使用子宫探针采用聚丙烯或适用材料制成，由柄部和细棒状有刻度的探针针身组成，可按材质、有无不透X射线标记、针径、长度等不同分为若干型号和规格；以无菌形式提供；供宫腔诊疗时，探测子宫宫腔的方向和深度用。 | Ⅱ
905 | 18-03-02 | 光学阴道镜 | 通常由观察系统、照明系统组成，观察系统是具有目镜、物镜的短工作距的体视光学显微系统，可外接图像采集显示系统。利用显微放大原理，观察物体细节。用于外阴、阴道、子宫颈的观察成像。 | Ⅱ |  | 新
906 | 18-03-02 | 电子阴道显微镜 | 由观察系统、照明系统组成，可外接图像采集显示系统。结构型式可为便携式、分体式、集成式，用于对宫颈、阴道外阴进行放大观察。 | Ⅱ
907 | 18-03-03 | 宫腔电切内窥镜 | 一般由硬性光学内窥镜、鞘套、操作器、手术电极组成，在内窥镜观察下，利用高频电流对宫腔组织进行切除。 | Ⅲ | 修订产品管理类别描述。
908 | 18-03-03 | 宫腔内窥镜 | 硬性光学内窥镜，一般由光学成像系统和照明系统组成。光学成像系统由物镜、转像系统、目镜三部分组成。照明系统一般为光学纤维。被观察物经物镜所成的倒像，通过转像系统将倒像转为正像，并传输到目镜，再由目镜或摄像系统放大用于观察。用于宫腔观察成像。 | Ⅱ
909 | 18-03-04 | 一次性使用妇科刮板 | 一次性使用妇科刮板由竹、木、高分子材料等制成，由手柄及头端组成，头端（病人端）为月牙形光滑平板；可按尺寸不同分为若干规格；产品以无菌形式提供，供妇科检查宫颈取样用。 | Ⅱ
910 | 18-03-04 | 子宫颈样本采集器 | 用于采集子宫颈上皮细胞、细胞团和细胞块。无菌，一次性使用。 | Ⅱ
911 | 18-04-01 | 一次性使用球囊子宫支架 | 通常由由导管、球囊、导管座组成。用于子宫内操作及手术之后的宫内放置，以减少子宫出血。置留人体时间小于24小时。 | Ⅱ
912 | 18-04-02 | 尿失禁悬吊带/尿道悬吊带 | 包括植入体内的吊带及植入工具。吊带由聚丙烯材料制成，可带有部分可吸收材料，植入工具由手柄及金属杆组成。作用原理为植入尿道附近形成对尿道的支撑。用于因女性的尿道过度移动和/或括约肌功能障碍而造成的压力性尿失禁。与已获准境内注册产品具有等同性。新型结构设计、药械组合产品除外。 | Ⅲ
913 | 18-04-02 | 子宫托 | 一般由金属或非金属材料制成。放置于妇女阴道（宫颈口），起到支持、支撑作用。在人体的滞留时间大于24小时小于30天。 | Ⅱ
914 | 18-04-02 | 子宫托 | 一般由金属或非金属材料制成。放置于妇女阴道、宫颈、子宫或输卵管中，起到支持、支撑作用。在体内滞留时间超过30天。 | Ⅲ
915 | 18-05-01 | 产床 | 产床一般由床面、床架、电气控制系统、配套件组成。用于孕妇分娩前、过程中、过程后使用。该产品性能指标采用下列参考标准中的适用部分：YY0571-2013医用电气设备 第2部分：医院电动床安全专用要求。 | Ⅱ
916 | 18-05-02 | 电动妇产科手术床 | 通常由台面、升降柱、底座、控制系统（控制台面调节）、升降立柱和电机组成。按传动原理可分为液压、机械和气动三种传动结构形式。有源产品。 | Ⅱ
917 | 18-06-01 | 宫内节育器及取放器械 | 通常是钩状、钳状或环状的器械。一般由高分子材料制成。无菌提供。用于宫内节育器/阴道夹持、放置和/或取出。 | Ⅱ
918 | 18-06-03 | 天然胶乳橡胶避孕套 | 天然胶乳橡胶避孕套采用天然胶乳制造，可按模型差异分为平面型、浮点型、螺纹型、异型；按添加成分不同分为若干种；可按标称宽度不同分为若干规格；彩色型添加食品级颜料，香味型添加食用香精，润滑型增加适量湿型润滑剂硅油用量；标称宽度、模型差异与添加成分可进行一定的组合。供男性用于避孕和预防性传播疾病。产品性能指标采用下列参考标准中的适用部分，如：GB 7544 天然胶乳橡胶避孕套 技术要求与试验方法。 | Ⅱ
919 | 18-06-03 | 聚异戊二烯合成橡胶避孕套 | 用聚异戊二烯橡胶制造，可按模型差异分为平面型、浮点型、螺纹型、异型；按添加成分不同分为若干种；可按标称宽度不同分为若干规格；彩色型添加食品级颜料，香味型添加食用香精，润滑型增加适量湿型润滑剂硅油用量；标称宽度、模型差异与添加成分可进行一定的组合。供男性用于避孕和预防性传播疾病。产品性能指标采用下列参考标准中的适用部分，如： GB 7544 天然胶乳橡胶避孕套 技术要求与试验方法。 | Ⅱ
920 | 18-06-03 | 女用聚氨酯避孕套 | 女用聚氨酯避孕套用由带内、外环和聚氨酯薄膜囊组成，内环封闭，内外表面涂有二甲基硅油，可按设计、技术参数等不同分为若干种；供女性用于避孕和预防性传播疾病。 | Ⅱ
921 | 18-06-06 | 超声引导宫腔手术监视仪 | 超声引导宫腔手术监视仪由主机、专用凸阵探头和监视器组成，超声成像部分可以是模拟或数字化成像，主机可以包括人流用负压吸引系统和相关专用附件，设备可按设计、技术参数、辅助功能、附件等不同分为若干型号；供人工流产、取放节育环等妇产科手术的实时监测及腹部超声检查等用。 | Ⅱ
922 | 18-06-06 | 一次性使用宫腔组织吸引管 | 一次性使用宫腔组织吸引管一般由高分子材料制成，可由手柄、管体组成，可按管径不同分为若干规格，以无菌形式提供，供临床与吸引器配合，吸取宫腔内组织用。 | Ⅱ
923 | 18-06-06 | 一次性使用流产吸引管 | 一次性使用流产吸引管一般由高分子材料制成，由管身和接头组成，吸引管上可有刻度，可为套管，可按材质、设计、技术参数等不同分为若干型号和规格；产品以无菌形式提供；与吸引器配套，供临床吸引法终止妊娠用。 | Ⅱ
924 | 18-06-06 | 流产吸引器 | 通常由吸引泵、开关、安全阀、止回阀、储液瓶、控制电路组成、与吸引管道、流产吸引管配套使用。可在其他影像设备监视下（如超声诊断仪）使用。用于对早期妊娠的孕妇施行人工流产手术。 | Ⅱ
925 | 18-07 | 胚胎转移导管辅助器 | 该产品用于补充和辅助经认可的胚胎移植器械插入子宫，以便置放体外受精(IVF)胚胎进入子宫腔。该产品由硬芯，导管和手柄等组成。 | Ⅱ
926 | 18-07-01 | 一次性使用胚胎移植导管 | 一次性使用胚胎移植导管由导引导管、带鲁尔接头的移植导管和护套等组成；一般由高分子材料、不锈钢材料制成；可按材质、设计、技术参数等不同分为若干型号及规格；以无菌形式提供；与注射器配合，经阴道向子宫内转移经过体外受精(IVF)的胚胎或受精卵用。 | Ⅱ
927 | 18-07-01 | 一次性使用带鞘胚胎移植管 | 一次性使用带鞘胚胎移植管由保护套管、导引导管和移植导管组成。一般由高分子材料、不锈钢材料制成；可按材质、设计、技术参数等不同分为若干型号及规格；以无菌形式提供；与注射器配合，经阴道向子宫内转移经过体外受精(IVF)的胚胎或受精卵用。 | Ⅱ
928 | 18-07-01 | 辅助生殖导管 | 通常由导引导管、外套管、移植导管、连接件、保护套管组成。一般由高分子材料制成。无菌提供。授精导管用于经阴道插入子宫腔内，注入精液，进行人工授精。胚胎移植导管用于经阴道向子宫内移植经过体外授精（IVF）后的胚胎。 | Ⅱ
929 | 18-07-02 | 辅助生殖穿刺取卵/取精针 | 通常由针管、针套、针柄、软管以及胶塞接口、负压连接头、吸引管、真空管、冲洗管组成。无菌提供。用于经由阴道对卵巢穿刺及从卵巢卵泡中对卵母细胞进行抽吸和冲洗，或用于经皮睾丸/附睾穿刺取精。 | Ⅱ
930 | 18-07-03 | 辅助生殖微型工具 | 辅助生殖用微型工具，包括IVF显微操作用注射、持卵、剥离、活检、辅助孵化和取精用的微细管状和针状工具。还包括与生殖细胞和胚胎接触的器皿、盘板。无菌提供。用于体外环境下操作或储存人类生殖细胞以及胚胎，包括精子显微注射，胚胎活检以及辅助孵化。 | Ⅱ
931 | 18-07-04 | 体外辅助生殖培养用油 | 由石蜡油（矿物油）制备而成，用于体外辅助生殖技术中覆盖培养液或培养液滴，建立配子、合子或胚胎与外界环境间的保护屏障。产品与已获准境内注册产品具有等同性。新型作用机理、药械组合产品除外。 | Ⅲ
932 | 18-07-04 | 精子密度梯度分离液 | 一般由硅烷包裹的硅胶颗粒、无机盐成分、能量底物、酸碱缓冲成分、氨基酸、人血清白蛋白、抗生素、水等成分组成。不含有我国同类已上市产品中未含有的组成成分。用于作为密度梯度分离介质分离精子。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ
933 | 18-07-04 | 精子制动液 | 一般由聚乙烯吡咯烷酮、无机盐成分、能量底物、酸碱缓冲成分、氨基酸、人血清白蛋白、抗生素、水等成分组成。不含有我国同类已上市产品中未含有的组成成分。用于在对精子进行显微操作时降低精子活动力。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ
934 | 18-07-04 | 颗粒细胞去除液 | 一般由透明质酸酶、无机盐成分、酸碱缓冲成分、抗生素、人血清白蛋、水等成分组成的溶液。用于辅助生殖技术中去除卵子颗粒细胞。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ
935 | 18-07-04 | 体外辅助生殖用操作液 | 用于取卵、取精、配子/胚胎洗涤、操作的短时接触溶液，如取卵液、卵泡冲洗液、配子缓冲液、操作液、精子洗涤液、器皿冲洗液等。不含有我国同类已上市产品中未含有的组成成分。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ
936 | 18-07-04 | 体外辅助生殖用冷冻/解冻液 | 一般由冷冻保护剂、无机盐成分、能量底物、酸碱缓冲成分、氨基酸、人血清白蛋白、抗生素、水等成分组成。不含有我国同类已上市产品中未含有的组成成分。产品与已获准境内注册产品具有等同性。预期用于人配子、胚胎、卵母细胞的冷冻/解冻，包括程序化及玻璃化冷冻解冻。产品与已获准境内注册产品具有等同性。新型作用机理除外。 | Ⅲ
937 | 18-07-05 | 胚胎培养箱 | 通常由培养箱主机、温度控制系统、供气控制系统组成，能够提供一个可温控（接近人体温度）的环境。用于辅助生殖技术（ART）中对配子、胚胎及其他细胞进行培养。 | Ⅱ
938 | 18-07-05 | 时差培养箱 | 通常由培养箱主机、温度控制系统、供气控制系统组成。能够提供一个可温控（接近人体温度）的环境。时差培养箱可具有时差照相系统，图像捕获分析应用软件，服务器，客户端等。用于辅助生殖技术（ART）中对细胞、配子、胚胎进行培养及动态监测和分析，评估胚胎发育能力。 | Ⅱ
939 | 18-07-05 | 辅助生殖用恒温台、程序冷冻仪、程控降温仪、体外受精（IVF）超净工作台 | 通常由操作台、热台或控制器、冷冻装置等组成。 | Ⅱ | 修订分类编码。
940 | 19-01-07 | 盒式助听器 | 盒式助听器主要由输入换能器、信号调理单元、输出换能器和电源组成，多为模拟式助听器，可按设计、技术参数、预期用途、附加辅助功能等不同分为若干型号；经验配，供气导性听力损失患者补偿听力用。产品性能指标采用下列参考标准中的适用部分，如： GB/T 14199-2010 电声学 助听器通用规范。 | Ⅱ
941 | 19-01-07 | 耳背式助听器 | 耳背式助听器主要由输入换能器、信号调理单元、输出换能器和耳钩外壳等组成，可按设计、技术参数、预期用途、附加辅助功能等不同分为若干型号（模拟/数字/可编程等）；经验配，供气导性听力损失患者补偿听力用。产品性能指标采用下列参考标准中的适用部分，如： GB/T 14199-2010 电声学 助听器通用规范。 | Ⅱ
942 | 19-01-07 | 耳内式助听器 | 耳内式助听器包托耳甲腔式和耳道式助听器，主要由输入换能器、信号调理单元、输出换能器和电源等组成，可按设计、技术参数、预期用途、附加辅助功能等不同分为若干型号（现多为数字化可编程助听器）；经验配，供气导性听力损失患者补偿听力用。产品性能指标采用下列参考标准中的适用部分，如： GB/T 14199-2010 电声学 助听器通用规范。 | Ⅱ
943 | 19-02-01 | 步态训练设备 | 通常由减重装置、主机、跑台、控制装置、固定装置等组成。通过训练患者步态促进康复，可附带步态评估功能。用于对下肢步行障碍患者进行步态康复训练。不包括四肢联动设备，不包括外骨骼康复训练设备。 | Ⅱ
944 | 19-02-02/05 | 肢体康复训练设备 | 通常由床架、绑带、机械支撑部件、电动控制装置、固定保护装置等组成。通过改变体位、起立角度对患者进行训练促进康复，用于对脑中风、脑外伤、脑卒中等导致肢体运动功能障碍的患者进行肢体运动康复训练。不包括生物电信号采集、反馈和电刺激功能。不包括四肢联动设备，不包括外骨骼康复训练设备。 | Ⅱ
945 | 19-02-03 | 平衡训练设备 | 通常由测量平台、辅助支架、平衡训练软件等组成。通常对站立或坐在测试平台上的患者进行平衡能力训练，可附带平衡能力评估功能。用于对平衡能力障碍患者进行康复训练。不包括四肢联动设备，不包括外骨骼康复训练设备。 | Ⅱ
946 | 19-02-04 | 振动训练设备 | 通常由训练平台、控制装置、固定架等组成。通过周期机械振动方式，达到肌肉或关节康复的目的。用于改善运动功能障碍患者的肌肉功能、平衡性和协调性。不包括四肢联动设备，不包括外骨骼康复训练设备。 | Ⅱ
947 | 19-02-05 | 关节持续被动活动仪（CPM仪） | 关节持续被动活动仪（CPM仪）为有源设备，通常由主机、控制部分和关节固定部分组成，可按固定关节、技术参数、附加功能等不同分类若干型号；通过电动机械结构使患者相应关节反复衩动屈伸，持续被动运动可以刺激关节，使其生理性关节液产生良性循环，减轻关节肿胀，可防止或治疗因创伤或运动不足等原因导致或可能导致的关节运动幅度减小，关节僵硬和静脉栓塞等。 | Ⅱ
948 | 19-03-01 | 手动轮椅车 | 手动轮椅车由人力驱动，可由车架、前轮、后轮、坐、背垫、扶手、挂脚、刹车、脚踏板等组成，可按驱动方式、轮椅材质、配置、功能、尺寸、适用范围等不同分为若干型号，供行动困难的残疾人、病人及年老体弱者作代步工具用。产品性能指标采用下列参考标准中的适用部分，如：GB/T 13800-2009 手动轮椅车。 | Ⅱ
949 | 19-03-01 | 电动轮椅车 | 电动轮椅车为有配套椅式支持系统的两轮个人移动设备，有一个或数个电机提供驱动力。行进速度由电机控制，行进方向由陪同人员或乘用者手动或电动控制；可按材质、功能、附件、技术参数等不同分为若干型号；供残疾或无完全行走能力者（不含肥胖症）代步用。产品性能指标采用下列参考标准中的适用部分，如：GB 12996-1991 电动轮椅车和GB/T 18029 轮椅车 电动轮椅系列。 | Ⅱ
950 | 19-04 | 鼻中隔固定夹 | 鼻腔手术后插入鼻腔，起到固定鼻中隔用。 | Ⅱ
951 | 19-04-01 | 矫形器 | 通常由高分子材料、织物、金属等材料制成。穿戴于头部、躯干或四肢体表，用于矫正或预防畸形。 | Ⅱ
952 | 20-02-04 | 拔罐设备 | 通常由电动负压源、导管、罐体等组成。通过负压源使罐体内产生负压，从而吸附在肌肉上。 | Ⅱ
953 | 20-02-05 | 熏蒸仪 | 熏蒸仪由药液箱、进/出液管路、加热装置、温度/液位等控制单元、显示单元、治疗头等组成（不包括药物），可按设计、技术参数、适用部位、预期用途、辅助功能等不同分为若干型号；可配合药液，用于人体局部熏蒸治疗。 | Ⅱ
954 | 20-02-05 | 熏蒸床 | 熏蒸床可由床体、控制单元、药槽、温度控制器、液位控制器、加热装置、液晶显示屏等组成（不包括药物）；可按设计、技术参数、适用部位、预期用途、辅助功能等不同分为若干型号；可配合药液，用于人体躯干等部位熏蒸治疗。 | Ⅱ
955 | 20-03-01 | 针灸针 | 针灸针由针体和针柄构成，针体材料可以用奥氏体不锈钢等适用材料制成，针柄可以用不锈钢丝（管）或铜丝（管）或铝丝（管）或聚丙烯等制成，按是否带进针管，针柄型式（如环柄针、平柄针、花柄针、塑柄等）、针体直径及长度等不同分为若干型式和规格；可以灭菌或非灭菌状态提供；供专业人员用于人体体表穴位（含耳穴）不同深度的侵入式刺激用。产品性能指标采用下列参考标准中的适用部分，如： GB 2024-1994 针灸针。 | Ⅱ
956 | 20-03-03 | 小针刀 | 小针刀疗法是一种介于手术方法和非手术疗法之间的闭合性松解术，其操作的特点是在治疗部位刺入深部到病变处进行切割，剥离有害的组织，以达到止痛祛病的目的。小针刀是形状上似针又似刀的一种针灸用具，分手持柄、针身、针刀三部分，通常刀与针同宽，刃口锋利，可按形状和长短等不同分为若干型号和规格；可以以灭菌/非灭菌形式提供；供软组织损伤性病变和骨关节病变非直视下松解术用。 | Ⅱ
957 | 20-03-04 | 皮肤针 | 通常由针盘、针体、针尖和针柄组成。外形似小锤状，一端附有莲蓬状的针盘，在针盘下规则嵌有不锈钢短针。根据针的数目多少不同，分别称为梅花针（五支针）、七星针（七支针）、罗汉针（十支针）。用于叩刺穴位及其他部位的皮肤。无菌提供。 | Ⅱ | 新增
958 | 20-03-05 | 滚针 | 通常由支架、滚轮、不锈钢针、手柄等组成。用于体表特定部位的局部刺激，实施滚针疗法。无菌提供。 | Ⅱ | 新增
959 | 20-03-06 | 揿针 | 揿针通常由针体和针盘构成，也可粘贴于适形的敷贴上以方便留置，针体、针座多采用奥氏体不锈钢丝等适用材料经磨抛绕制而成，按预期用途、结构型式、针径、长度等不同分为若干型号和规格；多以非灭菌状态提供；供专业人员用于人体体表穴位（含耳穴）的表浅侵入式剌激。产品性能指标采用下列参考标准中的适用部分，如：YY 0105-1993揿针。 | Ⅱ
960 | 20-03-06 | 皮内针 | 皮内针由针尖和针盘两部分组成，也可粘贴于适形的敷贴上以方便体表留置，多采用奥氏体不锈钢丝等适用材料经磨抛绕制而成，按预期用途、结构型式、针径、长度等不同分为若干型号和规格，可以灭菌和非灭菌状态提供，供专业人员用于人体体表穴位（含耳穴）的表浅侵入式留置剌激。产品性能指标采用下列参考标准中的适用部分，如：YY 0105 揿针。 | Ⅱ
961 | 20-03-07 | 一次性使用埋线针 | 通常由衬心座、针座、针管、衬芯和保护套组成.用于穴位的穿刺埋线。 | Ⅱ
962 | 20-03-10 | 一次性使用浮针 | 通常由针芯、针座、软管和保护套组成。用于浮针疗法 | Ⅱ
963 | 21-00-00 | 迹法分析软件 | 印迹法分析软件用于将经平板扫描仪扫描获得的膜条信息保存为电子图片，识别膜条位置，鉴别条带特异性。不包括自动诊断功能。 | Ⅱ
964 | 21-00-00 | 临床管理软件 | 用于心电、产科诊疗活动管理。 | Ⅱ | 修订产品描述。
965 | 21-00-00 | 临床管理软件 | 用于肿瘤、重症与麻醉诊疗活动管理。 | Ⅲ | 修订产品描述。
966 | 21-01-02 | 放射治疗记录与验证系统软件 | 用于放射治疗之前的治疗参数验证和治疗过程中的治疗参数记录。 | Ⅲ
967 | 21-01-02 | 放射治疗轮廓勾画软件 | 用于放射治疗计划制定之前的肿瘤及相关器官、组织的轮廓勾画。 | Ⅲ
968 | 21-02-01 | 医学影像存储与传输系统软件 | 用于医学影像的传输、显示、输出和存储。 | Ⅱ
969 | 21-02-02 | 医学图像处理软件 | 用于对来源于单模式或多模式的医学影像进行处理。 | Ⅱ
970 | 21-03 | 数据处理软件 | 用于监护数据的传输、集中实时显示、报警或者用于生理数据的传输、显示和处理。 | Ⅱ
971 | 22-01-02 | 血细胞分析仪器 | 通常由血细胞检测模块、血红蛋白测定模块、机械模块、电子模块、计算机系统等组成。原理一般为电阻抗法、比色法、流式激光散射技术等。用于对血液/体液中有形成分进行定量定性分析，并提供相关信息。 | Ⅱ
972 | 22-01-04 | 凝血分析仪器 | 通常由预温模块、加样模块、计时模块、样品传送及处理模块、检测模块和计算机系统等组成。原理一般为凝固法、产色底物法和免疫比浊法等。用于对血液进行凝血和抗凝、纤溶和抗纤溶等功能的分析。 | Ⅱ
973 | 22-01-05 | 血小板分析仪器 | 通常由液路模块、样品处理模块、检测模块、计算机系统等组成。原理一般为比浊法等。用于分析血液样本中血小板数量、体积、聚集率等相关功能参数。 | Ⅱ
974 | 22-01-06 | 血液流变仪/黏度计 | 一般分为毛细管粘度计和旋转式粘度计。毛细管粘度计通常由毛细管、样品池、控温装置、驱动装置、计时器等组成；旋转式粘度计通常由加样模块、样本传感器、转速控制与调节模块、力矩测量模块、恒温模块等组成。原理一般为泊肃叶定律或粘滞定律等。用于临床对全血、血浆或血细胞流变特性进行分析。 | Ⅱ
975 | 22-01-08 | 流式细胞分析仪器 | 通常由流动室和液流系统、激光源和光学系统、光电管和检测系统、计算机和分析系统组成。原理一般为通过液流系统使单个粒子通过流动室并分析单个粒子的荧光标记信号，实现对其生物特性的分析。 | Ⅱ
976 | 22-02-01 | 生化分析仪器 | 通常由样品器、取样装置、反应池或反应管道、保温器、检测器、微处理器等中的一种或几种组成。原理一般为分光光度法、浊度比色法、离子选择性电极法、荧光法、反射光度法、差示电位法等。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
977 | 22-02-02 | 血糖及血糖相关参数分析仪器 | 通常由主机模块、电源模块、软件模块等组成。原理一般为电化学法、光反射技术、比色法等。不包含采血器具及适配试剂。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
978 | 22-03-01 | 电解质分析仪 | 通常由电极模块、测量模块、管路模块、电路模块和数据输出模块组成。原理一般为离子选择电极法等。用于分析血液及体液中的电解质含量。 | Ⅱ
979 | 22-03-02 | 血气分析仪 | 通常由电极模块、测量模块、管路模块、电路模块、数据输出模块等组成。原理一般为离子选择电极法和微电子和生物芯片技术等。用于测定血液及体液的pH、二氧化碳分压、氧分压等血气参数。 | Ⅱ
980 | 22-03-03 | 电解质血气分析仪器 | 通常由电极模块、测量模块、管路系电路模块和数据输出模块组成。原理一般为离子选择电极法等。 | Ⅱ
981 | 22-03-04 | 电解质、血气、生化分析用检测电极 | 通常由单项或多项电极块组成。与电解质分析仪、血气分析仪和含电解质模块的生化分析仪配套使用，用于电解质、血气或代谢物的分析。 | Ⅱ
982 | 22-04 | 全自动免疫分析仪 | 通常由取样中心、处理中心、废液和供应中心、系统控制中心等组成，通过以抗原抗体相互结合的免疫学反应为基础，使用酶标记或化学发光剂标记抗原抗体，通过一系列级联放大反应，将光信号或电信号与分析物浓度等相联系，分析人体样本中的待测的抗原或者抗体。用于对人类体液中的各分析物，如肿瘤标志物，病原体抗原抗体等进行定量、半定量或定性检测。 | Ⅱ
983 | 22-04 | 过敏原半定量分析软件 | 过敏原半定量分析软件内含标准曲线，配合普通扫描仪使用。用于对免疫印迹法和免疫层析法的过敏原特异性抗体和自身免疫抗体的反应试条进行半定量检测。不包括自动诊断功能。 | Ⅱ
984 | 22-04 | 全自动免疫发光仪 | 通常由加样模块、反应模块、光学检测模块（光电倍增管）、数据处理模块、温育温控模块和清洗分离模块等中的一种或几种组成。原理一般为将化学发光反应发出的光信号转变为数字信号，由数据处理系统经过计算得出浓度值。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
985 | 22-04-01 | 酶联免疫分析仪器 | 通常由传输模块、试剂加注模块、孵育模块、光学模块、清洗模块和数据处理模块中的一种或几种组成。原理一般为单色光经标本吸收后通过光电检测器将光信号转换成电信号，由数据处理系统经过计算得出浓度值。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
986 | 22-04-02 | 化学发光免疫分析仪 | 通常由加样模块、反应模块、光学检测模块（光电倍增管）、数据处理模块、温育温控模块和清洗分离模块等中的一种或几种组成。原理一般为将化学发光反应发出的光信号转变为数字信号，由数据处理系统经过计算得出浓度值。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
987 | 22-04-03 | 荧光免疫分析仪器 | 通常由加样模块、反应模块、光学检测模块（荧光）、数据处理模块、温育温控模块、清洗分离模块等中的一种或几种组成。原理一般为将荧光信号转变为数字信号，由数据处理系统经过计算得出浓度值。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
988 | 22-04-04 | 胶体金试纸分析仪 | 通常由光电检测模块、机械扫描控制模块、控制主板模块、信息采集模块等组成。原理一般为通过传感器将检测试剂卡的反射率信号转为光电信号，通过校准信息将光电信号转为相应的浓度值或阈值，对待测物进行分析。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
989 | 22-04-05 | 免疫印迹仪 | 通常由蠕动泵模块、加样模块、孵育模块、温控模块等组成。原理一般为用电转移等方法将蛋白转移到固相膜上，最后进行免疫学检测。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
990 | 22-04-06 | 特定蛋白免疫分析仪 | 通常由光学模块、检测模块、计算机系统等组成。原理一般为免疫比浊法。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
991 | 22-05-03 | 核酸扩增分析仪器 | 通常由控制部件、热盖部件、热循环部件、光电部件、传动部件、嵌入式软件和分析软件、电源部件等组成。原理一般为利用温度控制，为核酸的体外扩增提供适宜环境，采集和分析扩增过程中产生的光、电信号。与适配试剂配合使用，用于样本基因的核酸体外扩增与分析。 | Ⅲ
992 | 22-05-04 | PCR扩增仪 | 通常由控制部件、热盖部件、热循环部件、光电部件、传动部件、嵌入式软件和电源部件组成。原理一般为利用温度控制，为核酸的体外扩增提供适宜环境。与适配试剂配合使用，用于样本基因的核酸体外扩增。 | Ⅱ
993 | 22-05-05 | 核酸分子杂交仪 | 通常由控温模块和控制面板模块等组成。原理一般为碱基互补原则。用于核酸分子的杂交。 | Ⅱ
994 | 22-06-02 | 微生物培养监测仪器 | 通常由孵育模块、检测模块、控制/报警模块、显示模块、随机软件模块等组成。还可包括空气过滤、条码扫描等，原理一般为通过测量光散射，光密度，电阻抗，压力感应，产色（二氧化碳）或是通过细菌直接计数的变化来确定细菌悬浮在液体培养基的浓度。用于临床培养、检测血液和体液等标本中需氧菌、厌氧菌、真菌和分枝杆菌等。 | Ⅱ
995 | 22-06-06 | 微生物鉴定药敏分析仪器 | 通常由自动接种器模块、孵育模块、鉴定模块、药敏分析模块、计算机模块、条码阅读器模块、随机软件模块等一种或几种组成。原理一般为通过形态学、生长、生理学及临床化学的手段鉴定从生物样本（如：血液、尿液、脑脊液、唾液或粪便）中分离出的传染性和/或病原性微生物。药敏部分：通过与含不同浓度抗菌剂的试剂配合使用，来确定从临床样本分离出的细菌病原体的药物敏感性。用于对临床分离出的微生物鉴定和/或药敏分析。 | Ⅱ
996 | 22-06-07 | 细菌内毒素/真菌葡聚糖检测仪器 | 通常由光道模块、自动恒温器模块和控温模块组成。原理一般为凝胶法、光度法等。用于细菌内毒素和/或（1，3）-β-D葡聚糖的检测。 | Ⅱ
997 | 22-06-08 | 13C、14C呼气分析仪 | 通常由进样模块、空气净化干燥模块、光学模块、测量模块、电气控制模块和气体采集模块组成。原理一般为呼气试验检测法。用于临床诊断由于幽门螺旋杆菌感染引起的疾病。 | Ⅱ
998 | 22-07-02 | 图像扫描仪器 | 通常由光学成像系统、图像采集系统、计算机、软件等组成。原理一般为以照相扫描的方式将载片或者切片上的细胞呈现为扫描图像。用于对临床样本的显微图像进行扫描、观察等。 | Ⅱ
999 | 22-07-03 | 图像分析仪器 | 通常由光学成像系统、图像采集系统、计算机、图像分析软件等组成。原理一般为以照相扫描的方式将载片上的细胞呈现为图像，并能对图片上的细胞进行分类标记及分析。用于对临床样本的显微图像进行观察、筛选、标记及分析等。 | Ⅱ
1000 | 22-08-02 | 液体闪烁计数器 | 通常由探测模块、测量模块和计算机系统组成，原理一般为采用液体闪烁体（闪烁液）接受射线并转换成荧光光子。用于测定氚、碳14等放射性核素发射出的β放射性射线。临床上与放免试剂盒配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
1001 | 22-09-01 | 干化学尿液分析仪 | 通常由机械模块、光学模块、电路模块等组成。原理一般为反射光度法等。与适配试剂配合使用，用来测量尿液中蛋白、葡萄糖、尿pH值、酮体、尿胆原、胆红素、亚硝酸盐等参数。 | Ⅱ
1002 | 22-09-02 | 尿液有形成分分析仪 | 一般分为流式细胞式和影像式。流式细胞式通常由光学检测模块、液流模块、电阻抗检测模块和电路模块组成。原理一般为流式细胞分析术。影像式通常由标本处理模块、光学计数池模块、显微摄像模块、数据处理模块等部分组成。原理一般为数字成像自动识别原理。用于尿液中有形成分的识别和分析。 | Ⅱ
1003 | 22-09-04 | 粪便分析工作站 | 通常由样本处理模块、显微镜模块、结果处理软件模块等组成。原理一般为化学法、免疫法及显微镜检法等。用于粪便标本的有形成分、潜血和病原微生物等的检测。 | Ⅱ
1004 | 22-09-05 | 精子质量分析仪 | 通常由显微图像扫描模块、温控系统模块、计数池模块、计算机系统、软件等组成。用于精子质量的分析。 | Ⅱ
1005 | 22-10-02 | 微量元素分析仪 | 通常由主机模块、微量分析工作台模块、软件、计算机系统等组成。原理一般为电化学分析方法、电位溶出法、原子吸收法及质谱法等。用于检测血液，尿液，毛发等样品中的各种微量元素。 | Ⅱ
1006 | 22-10-04 | 糖化血红蛋白分析仪 | 通常由主机模块、电源模块、软件模块等组成。原理一般为电化学法、光反射技术、比色法、免疫比浊法等。不包含采血器具及适配试剂。与适配试剂配合使用，用于人体样本中待测物的定性和/或定量分析。 | Ⅱ
1007 | 22-10-05 | 糖化血红蛋白层析柱 | 通常由柱体和固定部件等组成。与分析设备配套使用，用于对人体样本中的被测物进行分离。 | Ⅱ
1008 | 22-10-06 | 冰点渗透压测定仪 | 一般分为冰点渗透压测定和胶体渗透压测定。冰点渗透压测定仪通常由制冷模块、搅动模块、测温传感器模块和计算机系统组成；原理一般为振动原理结晶。胶体渗透压测定仪通常由半透膜及其固定装置参比液室电压传感器和计算机系统组成。原理一般为渗透原理。用于测量尿液、血液等样本的晶体渗透压和胶体渗透压。 | Ⅱ
1009 | 22-10-08 | 生物芯片分析仪器 | 通常由主机模块、光电信号采集器模块、计算机系统等组成。原理一般为采集生物芯片上的光、电信号，通过软件进行分析。用于临床实验室对多种医学检测项目进行定性、半定量或定量检测。 | Ⅱ
1010 | 22-11-01 | 动静脉采血针及连接件 | 通常由动静脉采血针、采血器、保护套和其他部件组成。用于采集动静脉血样。如产品具有特殊性能或结构等不能豁免临床。豁免情况不包括使用了新材料、新作用机理、新功能的产品。 | Ⅲ
1011 | 22-11-02 | 末梢采血针 | 通常由针、针柄、保护套等组成。可包括激发装置（如弹簧等）。无菌提供。一次性使用。用于临床医学上皮肤穿刺，以采集人体末梢血样。 | Ⅱ
1012 | 22-11-04 | 一次性使用真空采血容器 | 一次性使用真空采血容器通常由通常由管和头盖组成，管内壁附着或不附着添加剂或附加物，与一次性使用采血针配合使用，用于人体静脉血的收集、运输、存储。产品性能指标采用下列参考标准中的适用部分，如：YY 0314《一次性使用静脉血样采集容器》。 | Ⅱ
1013 | 22-11-05 | 末梢采血管 | 通常由毛细管、吸管、接头等组成。无菌提供时，管内壁有或无添加剂；非无菌提供时，管内壁有添加剂。用于人体末梢血的采集、存储。 | Ⅱ
1014 | 22-11-06 | 末梢血采集容器 | 通常由容器（管或瓶或管瓶或试管）、盖子和添加剂组成。用于人体末梢血样的采集、运输和存储等。 | Ⅱ
1015 | 22-11-08 | 一次性使用隐血采样胶囊 | 一次性使用隐血采样胶囊由医用空心胶囊、医用脱脂棉、棉线组成，供提取胃液作隐血检查用。 | Ⅱ
1016 | 22-11-09 | 无菌样本采样拭子、一次性使用无菌微生物拭子、一次性使用无菌采样拭子 | 通常由柄部和取样头组成，拭子可按材质、设计、尺寸等不同分为不同型号、规格，可带取样容器；通常以无菌形式提供；供人体自然腔道，如鼻腔、口腔、阴道、尿道口等部位沾取生物样本检验用。 | Ⅱ
1017 | 22-11-11 | 足跟采血器 | 通常由弹簧、刀片、弹出结构和外壳组成，刀片一般由不锈钢制成。无菌提供。用于早产儿或新生儿足跟采血。 | Ⅱ
1018 | 22-14-01 | 医用培养/恒温箱 | 通常由温湿度、气体浓度控制系统、电子显示系统、箱体等组成。用于人体来源样本的培养。 | Ⅱ
1019 | 22-14-02 | 厌氧培养系统 | 通常由取样室、操作室、厌氧罐、培养室等组成。用于厌氧、兼性厌氧微生物的培养。 | Ⅱ
1020 | 22-15-03 | 自动加样系统 | 通常主要由精密加样系统组成，可以包含传输系统、清洗系统、温育系统、混匀系统、软件系统等其他功能连接件。用于临床检验分析仪器分析前试剂或样本的精密加样。 | Ⅱ
1021 | 22-15-04 | 低温储存设备 | 通常由制冷装置、绝热箱体、电控机构等部件组成。用于组织、细胞、血液和血液制品等的低温储存或转运。 | Ⅱ
1022 | 22-15-04 | 脏器冷藏装置 | 无源器具，通常由冰块等冷媒以及箱体组成，用于人体脏器的冷藏转运。 | Ⅱ
1023 | 22-16-01 | 生物安全柜 | 产品一般由柜体、前窗操作口、支撑脚及脚轮、风机、集液槽、过滤器、控制面板及紫外灯和照明光源组成。要求前窗操作口流入最低平均流速的气流，下降气流经过高效过滤器过滤后送至工作区，污染气流经过高效过滤器过滤后可以排到下一环节，安全柜内所有生物污染部位均处于负压状态或者被负压通道和负压通风系统环绕。生物安全柜是具有前窗操作口的安全柜，操作者可以通过前窗操作口在安全柜内进行操作，对操作过程中的人员、产品及环境进行保护。 | Ⅲ
1024 | 22-16-02 | 洁净工作台 | 由柜体、风机、高效过滤器、脚轮、照明灯、紫外线杀菌灯及控制系统组成。本产品工作区内洁净度可达ISO 5级，用于临床实验室化验及实验，使局部操作环境达到一定洁净等级。 | Ⅱ
1025 | 通用02-14-01；眼用06-01-07；口腔17-04-14 | 一次性手术用冲洗针 | 一次性手术用冲洗针采用适用的高分子材料和金属材料制成，由针座和平头针管组成，可按材质、容量及针径等不同分为若干型号和规格；以无菌形式提供；与注射器配套使用，供显微外科、眼科、口腔科、整形外科等手术术中冲洗创口用。 | Ⅱ

## ENG
Serial number | Classification Code | Product Name | Product Description | Category | Remark | Serial Number | New
1 | 0 | Medical device products upgraded to Class II after sterilization in the catalog of Class I medical devices | Products that are upgraded to Class II after sterilization in the catalog of Class I medical devices, if no new risk occurs after sterilization or the degree of risk is reduced after sterilization, the clinical evaluation may be exempted . | II |  | 1
2 | 02 | Passive Surgical Instruments | Class II passive products under the 02 subcategory are exempt from clinical evaluation. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New mechanism of action and drug-device combination products are excluded; if the new structural design does not generate new risks or the degree of risk is reduced, clinical evaluation may be exempted. | II |  | 2
3 | 03 | Neuro and Cardiovascular Surgical Instruments | Class II passive products under the 03 subcategory are exempt from clinical evaluation. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New mechanism of action and drug-device combination products are excluded; if the new structural design does not generate new risks or the degree of risk is reduced, clinical evaluation may be exempted. | II |  | 3
4 | 11 | Medical equipment disinfection and sterilization equipment | Class II products under 11 subcategories are exempt from clinical evaluation. | II |  | 4
5 | 16 | Ophthalmic Instruments | Class II passive products under 16 subcategories are exempt from clinical evaluation. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New mechanism of action and drug-device combination products are excluded; if the new structural design does not generate new risks or the degree of risk is reduced, clinical evaluation may be exempted. | II |  | 5
6 | 18 | Obstetrics, assisted reproduction and contraceptive devices | Class II passive products under 18 subcategories are exempt from clinical evaluation. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New mechanism of action and drug-device combination products are excluded; if the new structural design does not generate new risks or the degree of risk is reduced, clinical evaluation may be exempted. | II |  | 6
7 | 01-02-02 | Medical Laser Fiber | It is composed of a laser connector, an optical fiber transmission body and an application end. After being connected to a laser, it is used to transmit laser light for laser surgery. The light output end of the optical fiber is a straight cut end face. | II |  | 7
8 | 01-02-02 | Disposable intranasal irradiation fiber optic head | It is composed of a laser connector, an optical fiber, and a patient end, and the patient end is U-shaped. Connected to a suitable laser, the laser light is delivered to the nasal cavity to irradiate the capillaries in the nasal cavity. | II |  | 8
9 | 01-03-01 | High frequency electrocautery | The high-frequency electrocautery device is composed of a main unit, a variety of therapeutic devices, and accessories. Usually, the rated frequency and output power are much lower than that of high-frequency surgical equipment. It can be divided into several models according to the design, type, technical parameters, and intended use; Coagulation, denaturation and/or necrosis of corresponding tissues during superficial operations in dermatology, ENT, gynecology and anorectal. | II |  | 9
10 | 01-03-01 | High Frequency Surgical Equipment | It usually includes a high-frequency generator host and other related accessories (foot switch, surgical electrode, neutral electrode, etc.); the use of high-frequency current (frequency range 200k-5MHz) directly flows through the human body to generate heat to perform routine cutting and repair of human tissue. Coagulation; used for routine cutting and coagulation of human tissue during conventional electrosurgery. The frequency range of high-frequency generators should be limited to the range 200k-5MHz specified in GB 9706.4; the scope of application is limited to routine tissue cutting and coagulation operations (general surgery, ENT, neurosurgery, gynecology, laparoscopic surgery, etc.); Including special clinical applications or usage methods, such as plasma cutting coagulation, large blood vessel closure, etc. | III |  | 10
11 | 01-03-03 | Sterile pipeline high frequency connection instrument | The high-frequency connection instrument for sterile pipelines adopts high-frequency radiation heating method to realize aseptic connection of two sterile high-molecule pipelines; the instrument is usually composed of a high-frequency generator, output and controller. With the help of molds, different tubes can be heat-bonded Diameter sterile tubing; used for aseptic tubing such as heat-sealed infusion tubing on the clinical site. | II |  | 11
12 | 01-03-03 | Argon Controller | Including accessories such as controller and argon spray pen; through the gas pressure reducing valve and argon control module, the high-pressure argon in the high-purity argon bottle is converted into the adjustable low-flow argon required for work, and it is used together with the high-frequency knife to become Argon Knife; used in conjunction with high-frequency electrosurgical knife to achieve blood coagulation in argon environment during electrosurgery, and can also be used to isolate air in cutting mode (without argon ionization effect). | II |  | 12
13 | 01-03-04 | Plasma Surgical Electrodes | Bipolar attachment for plasma surgery equipment. It usually consists of electrode head, electrode rod, handle, cable, connector and liquid pipeline. Used in conjunction with the main unit of plasma surgery equipment and perfusion pumps, it is suitable for cutting and coagulating soft tissues in the environment of physiological saline or other electrolyte solutions under open surgery or endoscopic surgery. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.4 Medical Electrical Equipment Part 2-2: Special Requirements for Safety of High Frequency Surgical Equipment, YY/T 1409-2016 Plasma Surgical Equipment. | Ⅱ/Ⅲ |  | 13 | New
14 | 01-03-04 | Cathode plates for disposable high-frequency surgical equipment | The cathode plate for disposable high-frequency surgical equipment is usually composed of release paper, conductive adhesive, aluminum foil, base lining (such as polyethylene foam or non-woven fabric coated with acrylic pressure-sensitive adhesive) and electrode wires, according to the conductive medium (metal electrodes conduct electricity) / ion conductive glue conduction), user population (adult type/child type), number of electrodes (unipolar/bipolar) and shape (square/elliptical), whether with wires, etc. are divided into different models and specifications; connecting the human body And high-frequency surgical equipment, providing a high-frequency current circuit with low current density to prevent human body burns; the other pole of the bipolar is used for the alarm circuit of high-frequency surgical equipment. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.4 Medical Electrical Equipment Part 2-2: Special Requirements for Safety of High Frequency Surgical Equipment. | II |  | 14
15 | 01-03-04 | High Frequency Surgical Accessories | The accessories of high-frequency surgical equipment are usually composed of surgical handles, cables of surgical accessories, surgical connectors and surgical electrodes, which may include manual switches and cooling water supply systems. Accessories can be divided into several categories, models and specifications according to working principles, additional functions such as cooling, technical parameters, and expected functions; they can be provided in sterile/non-sterile forms; including monopolar and bipolar forms. Matched with high-frequency surgical equipment, it is used for tissue cutting, separation, electrocoagulation to stop bleeding, and tissue coagulation in various clinical open surgical operations. Excludes plasma surgical electrodes and closure instruments for vessel closure. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.4 Medical Electrical Equipment Part 2-2: Special Requirements for Safety of High Frequency Surgical Equipment. | II | Modify product description. | 15
16 | 01-03-04 | High Frequency Endoscopic Surgical Instruments | The high-frequency electrode used to complete the surgical operation under the endoscope belongs to the application part of the high-frequency surgical equipment; under the endoscopic operation, it enters the human body through the endoscopic instrument channel or other instrument channels; in the endoscopic operation For routine cutting and coagulation of human tissue. | III |  | 16
17 | 01-03-04 | Needle electrodes for radiofrequency ablation | Needle electrodes for radiofrequency ablation are usually used in sets with puncture needles, cannula, etc., and belong to the application part of radiofrequency ablation equipment. The ablation needle contains a temperature sensor; it transmits the high-frequency current generated by the radio frequency generator to the designated human tissue, so as to achieve the purpose of ablation/coagulation/necrosis; it is used in conjunction with the radio frequency generator (radio frequency controller) for Ablation of human tissue. | III |  | 17
18 | 01-03-04 | capacitive neutral electrode | The product usually consists of external wires, insulating film, polymer gel and conductive mesh. Used in conjunction with high-frequency surgical equipment, it is designed with the principle of capacitor structure to form an alternating electric field between the patient and the electrode, so that the high-frequency current flowing through the human body passes through the capacitor in the form of a field, providing a return channel for the high-frequency current to prevent Excessive temperature rise or undesired burns in patient tissue. | II |  | 18
19 | 01-03-05 | Priming pump | It is generally composed of a host, a controller and cables, and is used in conjunction with radiofrequency ablation equipment to reduce the temperature of the part of the ablation area in contact with the patient. | II |  | 19
20 | 01-07-03 | Operating room equipment overall control system | It usually consists of a host, a control module, a communication device connected to third-party equipment, an expansion device, and a connecting cable. It is used for centralized control and display of related equipment (including medical devices and non-medical devices) in the operating room. The system does not change the original structure and function of the controlled equipment, and the realized control function and display parameters do not exceed the original function and parameter adjustment range of the controlled equipment itself. | III |  | 20
21 | 01-08-00 | Disposable spinal surgery lighting fiber optic | The product is composed of light source connector, optical fiber and head end. Supplied sterile. After connecting with the lighting source, it transmits light energy to the surgical site to provide lighting for minimally invasive spinal surgery. | II |  | 21
22 | 01-08-01 | Surgical shadowless lamp | The surgical shadowless lamp is composed of a fixed seat, a beam, a balance body, a balance elbow, a lamp head bend, a lamp head, a dimming handle, and a handle. It can be divided into several models according to the light source, structure type, technical parameters, configuration, and additional functions. For surgical lighting. The product performance index adopts the applicable parts of the following reference standards, such as: YY 0627-2008 Medical Electrical Equipment Part 2: Specific Safety Requirements for Surgical Shadowless Lamps and Diagnostic Lighting Lamps. | II |  | 22
23 | 01-09-01 | Endoscopic surgery power system | The product is generally composed of a host, a handle, and a knife head. The motor drives the handle and the knife head to perform mechanical movements, and is used for endoscopic surgery to achieve surgical functions such as mincing or resecting tissue (except nasal cavity surgery). Apply or refer to the industry standard YY/T 0955 "Planer for Endoscopic Surgical Equipment". | II |  | 23
24 | 01-09-01 | Power system for sinus surgery | The product is generally composed of a host, a handle, and a cutter head. The motor drives the handle and the cutter head to perform mechanical movements, and is used for sinus surgery to achieve surgical functions such as resection of nasal tissue. Apply or refer to the industry standard YY/T 0955 "Planer for Endoscopic Surgical Equipment". | II |  | 24
25 | 01-10-00 | preoperative skin preparation | It generally consists of an electric handle host and a battery charger. Used in conjunction with single-use blades to remove hair from the patient's body and head prior to surgery. | II |  | 25
26 | 01-10-04 | Surgical Power System | It usually consists of a host, a control device, an electric motor, a handle and various cutters. Used for cutting/cutting, grinding, drilling, percussion, etc. during open surgery. | II |  | 26
27 | 01-10-05 | Skin grafting net machine | It is composed of mesh base, hob assembly, locking assembly, sterile slide, etc., and is driven by external active instruments, and is used to expand skin grafts for burn treatment. The slides of different specifications can realize the expansion of different proportions of the skin slices. | II |  | 27
| #N/A |  | Working principle:
| #N/A |  | Put the skin to be expanded on the sterile slide, place the slide on the mesh base, and then drive the external active device (motor) connected to this product, so that the power of the motor is transmitted to this product, and finally Make the hob part of this product to realize the crushing of the carrier in the way of rotating motion.
28 | 01-10-05 | Skin grafting machine | Usually consists of a host and a handpiece. For skin removal, skin grafting or skin rubbing. | II |  | 28
29 | 02 | Reusable Endoscopic Surgical Instruments | Reusable endoscopic surgical instruments are made of stainless steel and are composed of handles, jaws, scissor heads, etc.; they can be divided into several models and specifications according to different materials, designs, technical parameters, working parts, and intended uses; they are provided non-sterilized, Reusable; cooperate with hysteroscope, thoracoperitoneal endoscope, nephroscope and cystoscope, etc., to grasp, bite, separate tissues or foreign bodies through natural or surgical channels. | II |  | 29
30 | 02 | Surgical Instruments for Thoracic and Laparoscopic Endoscopy | Surgical instruments for thoracoabdominal endoscopy are passive components under the endoscope, made of polysulfone resin, stainless steel and other materials, consisting of pipelines, puncture needles, puncture needle sleeves, forceps, knotted push rods, fibroid screws , probe rod, tissue retractor, cotton swab applicator, dilator, surgical hook, irrigation device, cap, connector, adapter, etc.; it can be divided into several types according to different materials, designs, technical parameters, working parts, and intended uses. Model and specifications; can be provided in sterile/non-sterile form; used in conjunction with thoracic and abdominal endoscopes for the diagnosis and treatment of patients with thoracic and abdominal diseases. | II |  | 30
31 | 02 | Disposable intracranial suction tube | The disposable intracranial suction tube can be made of suitable materials such as polyvinyl chloride, and can be designed as a single tube or a cannula. The head end and side tube of the tube have smooth round holes, and can be divided into several types according to different types and diameters. Model and specifications, provided in sterile form; used for brain surgery to absorb intracranial body fluid, blood, flushing fluid, etc., and to clean up the surgical field of view. | II |  | 31
32 | 02 | Suction/perfusion system (tube) | This product can be composed of handle, suction perfusion tube, connector, sealing cap, spare valve tube, sealing ring and other components. This product is sterilized and disposable. Cooperating with the suction and perfusion control host, it is used for suction and irrigation during thoracic and laparoscopic operations. | II |  | 32
33 | 02 | Non-orthopedic surgical instruments under passive endoscope | Passive short-term (≤24 hours) use devices, not used in conjunction with active devices, marketed for many years, non-innovative functional design, not made of new materials, no coating, no degradation, absorption function, no drugs, animal sources Sexual components, used to assist in the completion of endoscopic examination, guidance, extraction, ligation (except staplers), imaging, dilation, clipping, separation, cutting tissue or stone removal/crushing. | II |  | 33
34 | 02 | Renal dilator/dilator sheath | Products are generally made of polytetrafluoroethylene, polyethylene, polyvinyl chloride, polypropylene and other materials, and can be divided into several models and specifications according to material, design, technical parameters, etc., with probes or without probes. Supplied sterile for single use. For urological surgery, the establishment of percutaneous access. | II |  | 34
35 | 02 | endoscope bridge | The material of the product is stainless steel, which is an accessory of urology endoscope. It is used to connect endoscopes and tube mirrors to provide channels for surgical instruments for surgical treatment or observation and diagnosis. | II |  | 35
36 | 02 | endoscopic obturator | The product is made of stainless steel, inserted into the mirror sheath, and used with the mirror sheath. In urological and gynecological endoscopic diagnosis and surgery, when the endoscopic sheath is inserted into the human body, the front end of the endoscopic sheath is closed to avoid damage to human tissue. | II |  | 36
37 | 02 | Gastrointestinal endoscope spiral outer tube (without lubricant) | This product consists of a hose with a helix at the end and a handle at the top. The top end of the hose also comes with a swivel collar and connection assembly. Aids in inserting and advancing the endoscope to the target site. | II |  | 37
38 | 02 | Sleeve Skin Extender | It consists of a sleeve, an inner core and a casing. According to the size and location of the wound, the product can cooperate with the traction wire to stretch and extend the skin around the wound, so that the distance between the wound and the margin can be gradually reduced or moved closer. | II |  | 38
39 | 02 | flushing line | This product consists of irrigation tubing for irrigation and suction during surgery. | II |  | 39
40 | 18-01-10 | intrauterine balloon catheter | This product consists of balloon, catheter, catheter seat, check valve, connector, support rod, etc. It is used to inject contrast medium into the uterus and fallopian tubes for hysterosalpingography, or to inject methylene blue for tubal formation in abdominal surgery surgery and/or tubal anastomosis. | II | Revised classification codes. | 40
41 | 02 | Disposable laparoscopic soft instrument sheath | Disposable laparoscopic soft instrument sheath (abbreviated as soft sheath) is a single-incision laparoscopic surgical instrument, which can be composed of incision protective sleeve and multi-channel sealing body. The incision protection sleeve retracts and protects the incision. The clasp is the base, the multi-channel sealing body is sealed and fitted with the base in the form of a cap, and the valve channel of the sealing body is used as a laparoscope, forceps, scissors and other instruments for single-hole laparoscopic surgery. Sealed operating channel in and out of the incision. The product is sterilized and disposable. Product performance indicators can use the applicable parts of the following reference standards, such as: YY 0672.1-2008 Endoscopic Instruments Part 1 Laparoscopic Trocar. | II |  | 41
42 | 02 | Disposable endoscope sheath | The disposable endoscope sheath can be composed of multiple tubes, plugs, joints, clips, flanges, sprinklers, nozzles (lenses) with fixed distances, silicone tube seals and other components. The product is sterilized and disposable. Provides a sterile disposable protective layer for endoscopes used during flexible endoscopy of the upper respiratory tract, vocal cords or nasopharynx. | II |  | 42
43 | 02 | Disposable incision protector | Disposable incision protectors are generally made of polymer materials. It can be composed of outer ring, insert ring and channel. The products can be divided into fixed-height type and variable-height type, and are divided into several specifications according to different sizes. It is suitable for endoscopic surgery and small incision surgery, to expand the incision field, protect the incision from damage, and reduce incision infection. | II |  | 43
44 | 02 | medical rinse head | This product can be composed of flushing head seat, flushing head and other components, and does not contain flushing liquid. Made of polymer material. This product is provided sterile and for single use. It is used for nasal irrigation of patients with rhinitis, and external auditory canal irrigation of patients with cerumen and otitis media. | II |  | 44
45 | 02 | Bracket Retriever | Used to capture and retrieve ureteral stents under endoscopic direct vision or X-ray. | II |  | 45
46 | 02 | Surgical Tissue Bags and Introducers | Consists of a surgical tissue bag and an introducer. The surgical tissue bag is composed of a tissue bag and a pull wire, and the introducer is composed of introducer pins and a handle. The product is sterilized and disposable. Surgical tissue bags are used to separate tissue during or prior to surgical removal and/or comminution of tissue; introducers are used to aid in the introduction of surgical tissue bags during endoscopic surgery. | II |  | 46
47 | 02 | Three-way cock | This product consists of a cock seat, a knob and a screw interface. Sterile and pyrogen-free. Single use. If the product has special properties or structure, it cannot be exempted from clinical practice. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 47
48 | 02-01-01 | Disposable surgical blade | Disposable surgical blades are made of materials suitable for making blades such as carbon tool steel or alloy tool steel, and may include handles; blades are divided into several types according to different types, and each type can be divided into different types according to the matching handle and blade size. Several sizes; product supplied sterile; for cutting soft tissue. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0174-2005 surgical blade, GB 8662-2006 matching dimensions of surgical blade and scalpel handle. | II |  | 48
49 | 02-01-03 | Disposable sterile skin preparation knife | It usually consists of a blade, a knife holder and a protective cover. Supplied sterile. Single use. For preoperative preparation, hair removal. | II |  | 49
50 | 02-03-03 | Disposable endoscope ligature scissors | Disposable endoscope ligature scissors are made of polyvinyl chloride, polyphenylsulfone, stainless steel, etc., and usually consist of a handle, tube body, cutter head/scissors storage part; it can be selected according to material, design, technical parameters, and applicable parts It is divided into several models and specifications; it can be provided in sterilized or non-sterilized form; it is matched with the endoscope and used for cutting ligature ends, etc. | II |  | 50
51 | 02-04-09 | Endoscopic Graspers | The endoscope grasping forceps can be composed of three-claw forceps, sheath tube and control handle and other components. This product is provided sterile, single use. Used for grasping and holding tissue during endoscopic surgery. | II |  | 51
52 | 02-04-09 | Endoscopic forceps | It usually consists of a head, a rod or a soft catheter and a handle, and the work of the head is transmitted and controlled through the handle operation. Generally the head is made of stainless steel. It is operated under the endoscope during the operation, and is used for clamping, clamping, and separating tissues, clamping foreign bodies, and clamping instruments. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 52
#N/A
53 | 02-04-10 | Disposable hose type biopsy forceps for endoscope | Disposable hose-type biopsy forceps for endoscopes are mainly composed of forceps head, spring outer tube and operating handle, and may have positioning pins, which can be divided into several models and specifications according to material, design, size, etc.; It is provided in the form of bacteria; it is used for biopsy sampling or forceps and removal of foreign bodies under endoscopes such as the digestive tract and respiratory tract. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T1076-2004 Endoscopic hose-type biopsy forceps. | II |  | 53
54 | 02-04-12 | Foreign body forceps for endoscope | Foreign body pliers for endoscopes are made of stainless steel and plastic, and are composed of pliers head, outer sleeve, connecting rod, and pliers handle. , cup-shaped head, pen holder type, reverse two-claw type, peanut type, three-claw type, etc., can be divided into several models and specifications according to different materials, designs, technical parameters, etc.; can be provided in sterilized or non-sterilized state; can be used with Bronchoscope, esophagoscope and laryngoscope, etc., are used for clamping foreign bodies in the trachea or upper gastrointestinal tract. | II |  | 54
55 | 02-04-13 | Endoscopic Looper Needle Holder | The endoscope looper needle holder is made of stainless steel, silica gel and resin, etc. It is composed of a needle holder groove, a push rod, and a control handle. It can be divided into several types according to different materials, designs, and technical parameters; it can be sterile Provided; for suture operation in laparoscopic surgery. | II |  | 55
56 | 02-04-13 | Endoscopic clip applier and clip remover | This product can be composed of pliers head, pliers rod, turntable, pliers handle and other components, usually made of stainless steel. reusable. This product is used as a delivery device for delivering or removing ligature staples/clips through a cannula of a specific size. | II |  | 56
57 | 02-06 | Minimally Invasive Fascial Closer | The product includes two parts: closure firmware and suture passer. The closed firmware is composed of a device body, a driver, a proximal wing, a bracelet, and a window, and the device body is composed of a seat plate, a guide hole, an outlet, a finger grip point, and a control panel. A suture passer consists of a needle and a handle. The product is sterilized and single-use. Used to gather tissue and suture percutaneously to close surgical incisions during laparoscopic surgery. | II |  | 57
58 | 02-06-01 | closing clip | Made of metal, high molecular polymer or other materials, it is used to temporarily clamp blood vessels or tissues during surgical operations and remove them immediately after surgery. Supplied sterile, single use. | II | Revised product name. | 58
59 | 02-06-01 | closing clip | For marking, hemostasis, or closure of tubular structures. Available with conveyor. Not removed after surgery. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. Absorbable materials, new structural designs, new mechanisms of action, and drug-device combination products are excluded. | III | Revised product description. | 59
60 | 02-07-01 | medical suture needle | Usually the shape is straight or curved, the needle point has round, triangular, shovel shape, and the needle tail has a hole. Generally made of stainless steel. Used for suturing tissue and skin. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0043 medical suture needles. | II |  | 60
61 | 02-07-01 | Disposable purse bag suture needle | Single-use purse-string suture needles are made of metal and suture and consist of two needles and one non-absorbable suture. The product is provided in sterile form and is suitable for purse-string suturing in digestive anastomosis procedures. The product performance index refers to the applicable part of the standard of YY 0877 purse-string suture needle. | II |  | 61
62 | 02-07-02 | Disposable sterile puncture needle | Disposable sterile pricking needles are made of applicable stainless steel and consist of two parts: a flat handle and a triangular needle tip; they are used to puncture the skin during skin testing for immediate hypersensitivity reactions, allowing the antigen solution on the skin to penetrate into the subcutaneous area . Puncture needles containing antigen solution are not included. | II |  | 62
63 | 02-07-02 | Disposable endoscopic pneumoperitoneum trocar | The disposable endoscopic pneumoperitoneum puncture trocar is made of applicable polymer materials such as ABS, and is composed of an outer sleeve and a puncture inner core that can be locked; it can be divided into several models according to different materials, designs, and sizes. It is used as a supporting surgical instrument for the endoscope, and is used to form the pneumoperitoneum gas injection channel and the working channel of the laparoscope or other surgical instruments by puncturing and fixing the abdominal wall. | II |  | 63
64 | 02-07-02 | skin pricking box | Made of polymethyl methacrylate polymer etc. Consists of a test head and a disc. Allergen testing kit for people with allergies. | II |  | 64
65 | 02-07-02 | Veress needle | It consists of needle tube, handle, plunger, locking joint, stylet, etc., and does not include Class III medical device components. Used to establish pneumoperitoneum during abdominal endoscopic surgery. The basic principle, scope of application, performance and composition are the same as those of the products already on the market. The exemption does not include: products using new materials, products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application. | II |  | 65
66 | 02-07-03 | breast tissue markers | Usually consists of a pushing tool (eg, a trocar) and a marker. The markers are generally made of stainless steel, titanium, or titanium alloy materials, and may contain polymer materials to assist positioning. Supplied sterile, single use. The markers were exposed to human tissue for more than 30 days. For breast tissue localization markers. Products with new materials, new designs, and new functions are excluded. | III |  | 66 | New
67 | 02-07-03 | Disposable breast positioning wire and guide needle | The guide needle tube in the product should be made of the materials specified in GB 18457, and the positioning wire should be made of stainless steel or other alloy materials. This product is used clinically in order to facilitate surgical resection of small local breast lesions. It is used under X-ray or ultrasound. The breast positioning wire and its guide are used to locate the lesion, and provide clear guidance to the clinical surgeon to narrow the incision and reduce the surgical injury. The product performance index refers to the applicable part in YY/T 0880 Disposable Breast Positioning Wire and Its Guide Needle. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 67
68 | 02-07-04 | Disposable endoscopic suction biopsy needle | It usually consists of a head, a rod or a soft catheter and a handle. The head is needle-shaped, and the work of the head is transmitted and controlled through the handle. Generally the head is made of stainless steel. Supplied sterile. The exemption does not include products using new materials, new mechanisms of action, and new functions, as well as products used for injection, positioning, and treatment. | II |  | 68
69 | 02-08-02 | Tissue Retractors for Endoscopy | Tissue hooks for endoscopes are made of stainless steel and aluminum, and are composed of working end (tie rod), instrument rod, handle, etc., and can be divided into several models according to different materials, designs, technical parameters, etc.; provided in sterile form ; For temporary stretching of tissues during various endoscopic procedures. | II |  | 69
70 | 02-09 | Disposable skin curette | It consists of a curette head, a handle, and may have a protective cap. The curette head is made of stainless steel or other non-toxic metal materials, and the handle and protective cap are made of plastic materials; it can be divided into multiple specifications and models according to the inner diameter of the curette head and the shape of the handle; the product is sterilized and disposable; applicable Scrape off necrotic tissue and dandruff during surgery. | II |  | 70
71 | 02-09-02 | endoscopic curette | This product can be composed of a tip part, an insertion part, a handle and other parts. The product is supplied sterile for single use. Inserted into the body cavity through the endoscope, through the forward and backward operation of the product and the bending operation of the tip of the curette, it is used to collect tissues or cells in the body cavity. | II |  | 71
72 | 02-11-00 | Endoscopic extra-abdominal dilator | The product can be composed of tube body, expansion bag, tube seat, sealing ring, three-way one-way gas injection valve (seat), introduction column, stop clip and protective cap and other components. Before laparoscopic minimally invasive surgery, an operating space required for surgery is established between the extraperitoneal cavity and the surgical site. | II |  | 72
73 | 02-11-03 | Disposable urethral dilators | Usually a long, slender design consisting of a rigid guide wire and a soft catheter. Supplied sterile, single use. Used to dilate the urethra. | II |  | 73
74 | 02-11-03 | Esophageal cardia stricture dilator | The dilator for esophageal cardia stenosis is composed of expansion strips (with X-ray imaging signs), guide steel wires, safety soft springs, etc. The expansion strips are made of soft polyvinyl chloride plastic, and the guide steel wires are made of nickel-titanium alloy. It is divided into several models and specifications according to different design, diameter/length, etc.; it can be provided in sterilized/non-sterilized form; it is used for endoscopic treatment of lower esophagus and cardia stenosis and dilatation. | II |  | 74
75 | 02-12-01 | Disposable laparoscopic trocar | Disposable laparoscopic trocars can be made of suitable metal materials and polymer materials. Typical structures can include puncture needles, puncture sleeves, gas injection valves, gas choke valves, and sealing caps; they can be made according to material, design and technology. The parameters are divided into several models and specifications; the product is provided in a sterile form; it is used to puncture the abdominal wall tissue of the human body during laparoscopy and surgery, and to establish a working channel for abdominal surgery. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0672.1-2008 Endoscopic Instruments Part 1 Laparoscopic Trocar. | II |  | 75
76 | 02-12-01 | laparoscopic trocar | This product generally consists of a puncture sleeve and a puncture core. The product is sterilized and disposable. It is used in conjunction with an endoscope to puncture human tissue and establish an abdominal cavity during endoscopic surgery. The exemption does not include: the use of new materials; the use of non-DEHP plasticizers; products containing special coatings such as functional polymers, drugs, and biological products, containing active ingredients, new technologies, new designs or new mechanisms of action, and products with Products with special structure, function, scope of application, etc. | II |  | 76
77 | 02-12-02 | Disposable skin tissue punch | Consists of a trephine bit and a handle, and may have a protective cap. Sterile, single use; suitable for skin tissue drilling. | II |  | 77
78 | 02-12-03 | guide wire for urinary tract | The guide wire for urinary tract is composed of a guide wire and a push tube. The materials and shapes of the guide wire are different: such as metal guide wire, metal elbow guide wire, PE guide wire, PTFE guide wire, PTFE elbow guide wire, PTFE coated Guide wire, PTFE-coated elbow guide wire, etc., the guide wire can be coated with hydrophilic/hydrophobic coating; it can be divided into several models and specifications according to product material, design, technical parameters, intended use, etc.; provided in sterile form ; Under the endoscope, it is used together with the J-shaped catheter and the minimally invasive dilation and drainage kit to play a supporting and guiding role. | II |  | 78
79 | 02-12-03 | Disposable ureteral introducer sheath | The disposable ureteral introducer sheath is made of nylon, polytetrafluoroethylene, stainless steel, polyethylene, polypropylene, etc., and usually consists of a sheath tube coated with a hydrophilic coating and a dilator, and may also contain a guide wire. The sheath tube can be single-lumen or double-lumen, and can be divided into several models and specifications according to different materials, designs, technical parameters, etc.; it is provided in a sterile form; it is used to establish channels for endoscopes and other instruments to enter the urinary tract during urological operations use. | II |  | 79
80 | 02-12-03 | Disposable cystostomy tube | : Disposable cystostomy tube is usually composed of a fistula tube, a guide needle and a connecting tube; it is provided in a sterile form; it is used for clinical temporary urinary diversion, suprapubic cystostomy, and urine drainage. | II |  | 80
81 | 02-12-03 | Non-vascular stent pusher | It is used to push ureteral stent or other non-vascular stent under endoscope. | II |  | 81
82 | 02-12-05 | Introducer for endoscope | Introducers for endoscopes usually consist of a head, a rod or a flexible catheter and a handle. The head is a guide, which transmits and controls the work of the head through the operation of the handle. Operated under the endoscope during surgery, it is used to guide instruments into the cavity or tissue. | II |  | 82
83 | 02-12-05 | Endoscope Introducer Sheath | This product can be composed of a sheath tube, a main body and a sealing cap (if any). This product is inserted into the trocar or body cavity, and serves as a sheath for the insertion channel of the endoscope, preventing the endoscope from touching the tip of the trocar and the inside of the trocar, and protecting the tip and bending part of the endoscope. | II |  | 83
84 | 02-12-05 | Disposable guide wire for digestive endoscopy | The product can be composed of guide wire body, protector and other components. Enter the cavity through the endoscope of the digestive tract, and use it to guide various catheters to the designated positions. | II |  | 84
85 | 02-13-00 | Spacer/Non-absorbable suture with spacer | Spacers used with non-absorbable sutures or non-absorbable sutures with spacers are used in soft tissue or organ sutures to support the sutures and prevent the sutures from tearing human tissue. The material of the gasket is polytetrafluoroethylene, expanded polytetrafluoroethylene, polyester, etc. New structural designs, new materials, etc. are excluded. | III | Add | 85 | New
86 | 02-13-01 | Disposable Automatic Purse Stapler | Disposable automatic purse-string sutures are made of metal, polymer materials, and non-absorbable surgical sutures. They are composed of clip holders, nail holders, nails, and purse strings. They are divided into several models and specifications according to materials, designs, and technical parameters. ; Provided in sterile form; for purse-string suture in open digestive tract surgery. | II |  | 86
87 | 02-13-01 | Disposable circumcision stapler | Disposable circumcision staplers are made of stainless steel and ABS, and staples can be made of stainless steel or suitable metals, usually including circular cutters, positioning springs, metal gaskets, fixed pins and movable connecting pieces. 316L stainless steel is Suture staples made of ABS material, concave seat made of ABS, annular knife anvil, nail bin protective cover, nail bin, top body, main body shell, mover, main body fixing ring, positioning adjustment nut, main body upper and lower covers, movable hand Put, insurance pin. Products should be sterile. It is suitable for clinical circumcision and suture surgery. | II |  | 87
88 | 02-13-01 | Disposable skin stapler | Disposable skin staplers are generally composed of a push piece, a shell, a handle, and staples, and a staple remover is optional; it can be divided into several models and specifications according to the design, the number of staples, and the shape and size of the staples after forming. ; The product is provided in a sterile form; it is used for superficial skin suture of wounds and surgical incisions. | II |  | 88
89 | 02-13-03 | Endoscopic Stapler | Composed of a needle insertion device, a suture needle, a transmission mechanism, and a control handle, the stapler can be used with sutures or with a thread box. It is suitable for minimally invasive sutures with natural orifices, and intermittent sutures or continuous sutures in soft tissues. Disposable or reusable. Products with absorbable sutures or for vascular sutures are excluded. . | II |  | 89
90 | 02-13-07 | Nonabsorbable Surgical Sutures | Non-absorbable surgical sutures can be made of silk, polyamide 6/66, polyethylene terephthalate, polypropylene, stainless steel wire, titanium-nickel memory alloy wire, etc., with or without suture needles. Divided into several models and specifications according to material, number of strands/braiding, coating, needle, bundle/cluster, wire diameter, use site, etc.; it can be provided in a sterile or non-sterile state; it is used for suturing and ligation of human tissues . Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0167 Non-absorbable surgical sutures. | II |  | 90
91 | 02-14-01 | Aspirator | It usually consists of flushing pipe, pipeline and connection port. For flushing tissue or aspiration. | II |  | 91
92 | 02-15 | Disposable finger-operated surgical instruments | Disposable finger-operated surgical instruments consist of polycarbonate sleeves and stainless steel end effectors, which can be divided into several models and specifications according to design, technical parameters and intended use; provided in sterile form; used for colorectal, Urology, gynecology and other laparoscopic surgery (HALS) and open surgery, for grasping, moving, separating, cutting and suturing tissues. | II |  | 92
93 | 02-15 | Disposable funnel for breast augmentation surgery | Disposable breast augmentation funnel made of elastic nylon material with inner surface coated with a hydrophilic coating of polyurethane and polyvinylpyrrolidone blend; supplied sterile; during surgical breast augmentation, place the implant in The inner surface of the device contains a hydrophilic lubricating coating, and through extrusion, the implant can easily enter the predetermined site through the small opening of the funnel, thereby reducing the patient's incision. | II |  | 93
94 | 02-15-01 | Disposable foreskin ligation device | It usually consists of a collar, a clamping ring, and a connecting device. Supplied sterile. Single use. For circumcision, seam-free healing. | II |  | 94
95 | 02-15-01 | Disposable anorectal ligation device | The disposable anorectal ligation device can be made of applicable metal, plastic and silicone, and can be composed of a gun body (built-in spring), a gun barrel, a silicone ring, a negative pressure suction connector, a winding wheel, a ligation device switch and a cap, It can be equipped with single or multiple silicone rings, and can have an auxiliary lighting device. It can be divided into several models and specifications according to different materials, designs, technical parameters, and intended uses; it is provided in a sterile form; it is suitable for various internal hemorrhoids and mixed hemorrhoids or Ligation of benign polyps of the rectum. | II |  | 95
96 | 02-15-02 | Suture Knot Pusher | The suture knot pusher is made of metal and other materials. It consists of a head end and a handle. There is a suture penetration hole at the head end. The extracorporeal suture knot is pushed onto the body surface or into the chest cavity. Supplied sterile. | II |  | 96
97 | 02-15-03 | support rod | Consists of a nylon rod and swivel head. Used in loop ostomy surgery, passing through the mesentery to support the intestinal tube in a fixed position, preventing the intestinal loop from returning to the abdominal cavity. | II |  | 97
98 | 02-15-04 | Disposable Surgical Electrode Cleaning Tablets | Disposable surgical electrode cleaning sheets can be composed of (sponge) base layer, (SiO2, etc.) fine abrasive layer, pressure-sensitive adhesive layer, (barium line) development line, (silicon oil) release paper, and are divided into several specifications according to shape and size. ; Supplied in sterile form; used for cleaning burnt tissue and other attachments on high-frequency surgical electrodes during surgery. | II |  | 98
99 | 02-15-04 | Disposable sterile fistula brush | Disposable sterile fistula brush can consist of a handle, bristles and silk thread, provided in a sterile form; for identification and cleaning of rectal fistula. | II |  | 99
100 | 02-15-04 | blade cleaning sheet | It usually consists of aluminum oxide layer, polyester layer, double-sided adhesive tape and paper liner, etc. This product is a sterile product and is for one-time use. Tip Cleaning Tablets are used to clean the electrosurgical scalpel pen tips. | II |  | 100
101 | 02-15-06 | Disposable glans protector | The disposable glans protective cover is made of ABS plastic and provided in a sterile form; it is placed on the glans during circumcision to prevent the glans from being damaged when the electric knife cuts the foreskin. | II |  | 101
102 | 02-15-06 | incontinence butt plug | The anal plug for incontinence is made of polyurethane foam, which is composed of expansion plug, water-soluble film and gauze rope. | II |  | 102
103 | 02-15-08 | Medical skin marker pen (with ink) | It consists of roll core, ink, polyester fiber nib, pen cap, pen barrel, tail cap and ruler, where the ink is composed of water, alcohol, ethylene glycol, gentian violet and sodium benzoate. Marking and positioning on the patient's skin (without mucous membranes and skin breaks) during surgery, radiation therapy and skin therapy. | II |  | 103
104 | 02-15-13 | Breast Implant Trial | Usually the shell is silicone rubber, filled with silicone gel. Supplied sterile. Used to determine the shape and size of breast implants. | II | Revised product description. | 104
105 | 02-15-14 | Clip device | Used with an endoscope to place clips in the digestive tract. Clips are used for endoscopic marking, hemostasis of digestive tract tissue. Supplied sterile. Available with conveyor. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. Absorbable materials, new structural designs, new mechanisms of action, and drug-device combination products are excluded. | II | Revised product management category description. | 105
106 | 02-15-16 | Laparoscopic Knot Pusher | The laparoscopic thread knot pusher is made of stainless steel and other materials. It is composed of a thread knot fixing groove, a push rod and a handle. There is a suture penetration hole at the head end. It can be divided into several models and specifications according to different materials, designs, and technical parameters. ; Used to place extracorporeal knotted sutures into the laparoscopic field. | II |  | 106
107 | 02-15-17 | Disposable endoscope sampling brush | Sampling brushes for disposable endoscopes are usually composed of brush handles, sheaths, pulling wires, and brushes, and can be divided into different models and specifications according to materials, designs, technical parameters, and applicable parts; they are provided in sterile form; It is used to extract tissue samples from the human body for endoscopy or when used with a navigation system. | II |  | 107
108 | 02-15-17 | Disposable Stain Resistant Sample Brush | Disposable antifouling sample brush is composed of pipeline assembly, brush head assembly, polyethylene glycol and protective plug, which can be divided into several models and specifications according to material, design, technical parameters, applicable parts, and intended use; It is provided in the form; it is used for brushing the tissue or secretion of the diseased part of the patient under the endoscope. | II |  | 108
109 | 02-15-17 | Disposable sample brush | It is generally composed of a pipeline assembly, a brush head assembly and a protective plug, and is used to brush the tissue or secretions of the patient's lesion under the endoscope. | II |  | 109
110 | 02-15-18 | Disposable Endoscope Stone Retrieval Basket | Disposable endoscopic stone retrieval baskets usually consist of a handle with a Luer connector, a guide wire rod (or cannula) with a soft sheath, and a 360° wire self-expanding basket at the distal end, which can be customized according to material, design, and technical parameters. It is divided into several models and specifications according to different parts, applicable parts, and intended uses; it is provided in a sterile form; it is used for grasping, manipulating, and removing stones and other foreign bodies under the endoscope in the diagnosis and treatment of digestion and urology. | II |  | 110
111 | 02-15-18 | Endoscopic Mechanical Lithotripsy/Lithology Auxiliary | Endoscopic mechanical lithotripsy/lithotomy accessories are made of metal and polymer materials, and are composed of outer sheath tube, lithotripsy handle, flushing extension tube, and joint; it can be selected according to material, design, technical parameters, applicable parts, and intended use It is divided into several models and specifications; it can be provided in a sterile/non-sterile state, and can be used repeatedly; it is used for mechanical lithotripsy/stone extraction when other endoscopic stone extraction methods fail. | II |  | 111
112 | 02-15-20 | Disposable stone removal balloon | The product can be composed of balloons, joints, catheters and other components, wherein the catheter is a plastic tube with a porous lumen structure. This product is sterilized and disposable. Used in conjunction with an endoscope, it is used to enter the biliary tract through a duodenal endoscope to remove stones from the biliary tract. | II |  | 112
113 | 02-15-22 | Disposable endoscopic drug delivery tube | Disposable endoscopic drug delivery tubes are made of stainless steel, polyethylene, ABS, etc., and are usually composed of spray heads, catheters, connectors, etc., and may include guide wires, which can be selected according to material, design, technical parameters, applicable parts, and intended use It is divided into several models and specifications; it is provided in a sterile form; it is used in conjunction with an endoscope for lavage, spraying liquid medicine, imaging, and moving gallstones/biliary sand, etc. | II |  | 113
114 | 02-15-23 | Disposable endoscope ligature | Disposable endoscope ligation devices are mainly composed of sleeves, trigger pull wires, control handles, flushing joints, loading catheters, and ferrules, and can be divided into several models and specifications according to materials, designs, and technical parameters; Supplied in sterile or non-sterile form; performed endoscopically during surgery and used to bandage polyps or place ligature loops to prevent or control bleeding. Also used for ligation of varicose veins. | II | Revised product description. | 114
115 | 03-00-00 | Heart Valve Installation Kit | The heart valve installation tool kit consists of a knot pusher, a valve probe, and a valve placer. The knot pusher is used to push the extracorporeal suture knot to the body surface or into the chest cavity; the valve probe is used to test the flexibility of the mechanical valve leaflet; the placer is used to place the mechanical valve. | II |  | 115
116 | 03-00-00 | Fine-tuning device for neurosurgery endoscope | The neuroendoscope fine-tuning device can be composed of a fine-tuning device, a fixing fitting, a fixing arm and the like. Generally, it is made of stainless steel in accordance with YY/T0294.1, and it is a non-sterile product. For supporting and fine-tuning neuroendoscopes or trocars in neurosurgery. | II |  | 116
117 | 03-00-00 | transjugular intrahepatic puncture device | Usually consists of a stylet, catheter, stiffened cannula, dilator, introducer sheath, dilator, and guidewire. It is used for intrahepatic portal vein puncture through the jugular vein and intrahepatic shunt surgery of the portal vein to reduce the portal vein pressure. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 117
118 | 03-01-01 | chisel for neurosurgery | It usually consists of a cutter head and a stem. The cutter head is generally made of stainless steel, diamond, etc. For tissue dissection in neurosurgery. | II |  | 118
119 | 03-02-01 | brain scissors | Brain scissors are made of metal, consisting of a handle and a head, with a single joint. The head of the knife can be designed to be straight, curved up and down, etc. The handle and the head can form a certain two-dimensional or three-dimensional angle to adapt to different anatomical positions The end of the knife head can be pointed or blunt, and the surface is smooth without damaging the surrounding tissue; it can be divided into several models and specifications according to the material, design, technical parameters, and intended use; it is usually a repeatable instrument; It is used to cut out diseased bone or necrotic nerve/brain tissue during internal surgery. | II |  | 119
120 | 03-02-01 | Scissors for meninges | It usually consists of a pair of blades connected in the middle, with a cutting edge on the head and a ring on the handle. Generally made of stainless steel. For cutting meningeal tissue. | II |  | 120
121 | 03-02-01 | coronary scissors | The product consists of a pair of blades connected in the middle, with a cutting edge on the head and a ring on the handle. Generally made of stainless steel. Supplied non-sterile. Used to cut coronary arteries or widen coronary incisions. | II |  | 121
122 | 03-03-01 | intracerebral forceps | The intracerebral forceps are made of metal, consisting of a handle and a head, with a single joint. The head of the forceps can be designed to be straight, left and right, etc. The handle and the head of the forceps can form a certain two-dimensional or three-dimensional angle to adapt to different anatomical positions. Operation, the end of the forceps head can be round or oval or other shapes, the surface is smooth, and will not damage the surrounding tissue; it can be divided into several models and specifications according to the material, design, technical parameters, intended use, etc.; it is usually a repeatable instrument ; Used for clipping intracranial tumors and capsules during intracerebral surgery. | II |  | 122
123 | 03-03-01 | Valve Surgery Forceps | Valve surgery forceps are made of stainless steel and are used for grasping and removing hard substances on the heart structure, such as calcified material and fibrous tissue, during valve surgery. | II |  | 123
124 | 03-03-01 | Thoracic cardiovascular surgery forceps | Consists of a pair of centrally connected leaves with a pincer beak at the head. Generally made of stainless steel. Used for clamping organs, blood vessels, free vena cava and aorta during thoracic surgery. | II |  | 124
125 | 03-03-02/05 | neurosurgery intracerebral forceps | It usually consists of a beak, a stem and a handle. The beak has a cutting edge and the handle has a finger ring. Generally made of stainless steel. It is used for clamping and biting off brain tissue, foreign bodies, growths or removing tumors. | II |  | 125
126 | 03-03-04 | Atraumatic arteriovenous blocking forceps | Non-invasive arteriovenous blocking forceps are made of metal, single-joint, composed of a handle and a head, the head of the forceps can be designed to be straight, curved up and down, etc. For operations in different anatomical positions, there may be semi-rigid inserts inside the forceps head to increase the stability and non-injury of the forceps. The forceps can be locked and can be divided into several models and Specifications; usually reusable instruments; intended for arteriovenous occlusion in procedures involving major vessels. | II |  | 126
127 | 03-04-01 | intracerebral tweezers | The tweezers for the brain are made of metal, consisting of a handle and a head. The tip of the head can be designed to be tapered or round. The handle and the head can form a certain two-dimensional angle. It does not damage the surrounding tissue; it can be divided into several models and specifications according to the material, design, technical parameters, and intended use; it is usually a repeatable instrument; it is used for grasping dressings and soft tissues in the brain during intracerebral surgery. | II |  | 127
128 | 03-04-01 | Forceps for minimally invasive cardiac surgery | Minimally invasive cardiac surgery forceps are made of stainless steel and are used to pinch and manipulate cardiac structures during minimally invasive cardiac surgery. | II |  | 128
129 | 03-04-01 | Forceps for thoracic and cardiovascular surgery | It consists of a pair of superimposed leaves with a stalk and a straight or curved head. Generally made of stainless steel. Supplied non-sterile. Used to clamp myocardial tissue and blood vessels. | II |  | 129
130 | 03-04-01/02 | Neurosurgery forceps | Usually consists of a pair of gun-shaped blades superimposed on the tail. Generally made of stainless steel. For clamping soft tissue or removing tumors during brain surgery. | II |  | 130
131 | 03-05-01 | Disposable sterile scalp clip | The disposable sterile scalp clip can be made of suitable polymer materials such as polystyrene resin or polyoxymethylene. It has a U-shaped structure, and the blades of the two connected clamping arms bite each other. It can be placed manually or with scalp clamps during use. Non-implantable device; the product is provided in a sterile form and is used to fix the scalp and stop bleeding during clinical operations. | II |  | 131
132 | 03-05-02 | Hemostat | Used to clamp small blood vessels and tubular tissue to stop bleeding during brain surgery. It can also be used in minimally invasive surgery such as abdominal cavity. Not removed after surgery. Available with conveyor. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 132
133 | 03-07-01 | brain hook | The intracerebral hook is made of metal and consists of a handle and a head. The tip of the head can be designed as a semicircular arc or a hook with a smooth transition. It has anti-slip design, smooth surface, and does not damage surrounding tissues; it can be divided into several models and specifications according to different materials, designs, technical parameters, and intended uses; it is usually a repeatable instrument; it is used for stretching nerves or brain tissues during intracerebral surgery Wait for it. | II |  | 133
134 | 03-07-01 | aortic retractor | The product consists of a head and a handle. Head with hook. Generally made of stainless steel. Supplied non-sterile. Used to stretch the aorta. | II |  | 134
135 | 03-08-01 | brain curette | Brain curettes are made of metal, consisting of a handle and a head. The tip of the head can be designed as a spoon. The handle and the head are usually coaxial. The handle can have a non-slip design. It is divided into several models and specifications according to the material, design, technical parameters, and intended use; it is usually a repeatable instrument; it is used for scraping diseased bone and necrotic tissue during intracerebral surgery. | II |  | 135
136 | 03-08-01 | neurosurgery curette | Usually of elongated design, with a handle at the proximal end and a spoon-shaped distal end. Generally made of stainless steel. It is used for peeling and curettage during craniocerebral surgery. | II |  | 136
137 | 03-09-01 | brain dissector | The intracerebral dissector is made of metal, etc. It consists of a handle and a head. The tip of the head can be designed in various shapes with different widths. The handle and the head are usually coaxial. The handle can have a non-slip design, and the surface of the head is smooth. It does not damage surrounding tissues; it can be divided into several models and specifications according to different materials, designs, technical parameters, etc.; it is usually a repeatable instrument; it is used for peeling meninges during intracerebral surgery. | II |  | 137
138 | 03-09-01 | Stripper for thoracic and cardiovascular surgery | It can be composed of parts such as stripping head, stem and handle. Generally made of stainless steel. It is used to peel off soft tissue, or to peel off thrombus in veins, etc., so that the veins can be unblocked. | II |  | 138
139 | 03-10-01 | intercostal soft tissue retractor | The intercostal soft tissue retractor consists of a retractor ring, retractor wall and adhesive tape. The retracting ring can be made of polyester and other elastic materials, and the retracting wall can be made of suitable materials such as polyether polyurethane. It is divided into several models and specifications; it is provided in a sterile form; it is placed in the intercostal space in cardiac surgery, and the tissue is retracted to form the operation port. | II |  | 139
140 | 03-10-01 | Atrial retractor | The atrial retractor is made of stainless steel and consists of a triangular bracket, retractor blades, and forceps. For valve surgery, the anterior wall of the right atrium is retracted to expose the mitral valve for visualization of internal and external cardiac structures. | II |  | 140
141 | 03-10-01 | Disposable intracranial retractor fixation catheter | It usually consists of an introducer and dilator, made of polycarbonate polymer material. Supplied sterile, single use. It is suitable for expanding the soft tissue around the surgical field in clinical intracranial surgery, and provides a visual working channel in brain surgery. | III | Add | 141 | New
142 | 03-10-03 | mitral valve dilator | The mitral valve dilator is made of suitable metal materials, and consists of a handle with a spreader rod, a body and a dilator head. The expansion width can be adjusted. It can be divided into several models and specifications according to different materials, designs and technical parameters. ; Usually can be used repeatedly; for the dilation of the mitral valve during cardiac surgery. | II |  | 142
143 | 03-10-03 | Vascular dilator | It is usually composed of a head, a stem and a handle, and generally can be made of stainless steel or titanium alloy, and is used to expand the caliber of blood vessels in cardiothoracic surgery or vascular surgery. The basic principle, scope of application, performance and structural composition of the product are the same as those of the products already on the market. The exemption does not include: products using new materials, products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application. | II |  | 143
144 | 03-10-03 | percutaneous renal dilator | When used for percutaneous pelvic puncture, expand the subcutaneous tissue, muscle, and renal parenchymal tissue to establish and maintain a percutaneous access. | II |  | 144
145 | 14-00-00 | Percutaneous nephrostomy catheters and accessories | The product consists of percutaneous ostomy catheters, dilators, puncture needles, guide wires, introducer sheaths, balloon dilation catheters, pressure pumps and accessories. For percutaneous nephrostomy. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II | Revised classification codes, product names and product descriptions. | 145
146 | 03-11-01 | hole punch | Usually consists of a punching head and a shank. Supplied sterile. Single use. Exposure to central nervous system or circulatory system. It is used to punch holes in tissues and establish pathways during interventional procedures. | III |  | 146 | New
147 | 03-11-01 | blood vessel punch | Vascular punches are divided into one-time use and repeated use. Disposable blood vessel punches are made of suitable materials such as stainless steel and ABS, and are provided in a sterile manner; reusable blood vessel punches are made of stainless steel that conforms to ASTM F899, and are provided in a non-sterile manner; both types of products can be made according to the material It is divided into several models and specifications according to different types, designs, technical parameters, and applicable blood vessels; it is used to make a circular opening on the wall of arterial/venous blood vessels during surgery for subsequent blood vessel transplantation/anastomosis, etc. | II |  | 147
148 | 03-11-01 | Drills for neurosurgery | Generally made of stainless steel. reusable. For neurosurgery craniotomy. Supplied sterile. | II |  | 148
149 | 03-11-01 | electric bone drill | Usually consists of a drill bit and a drive unit. Drill bits are generally made of stainless steel. Used for tissue drilling and establishing access. | II |  | 149
150 | 03-12-01 | Cerebroventricular fluid suction tube | The cerebroventricular fluid suction tube is made of metal, consisting of a tube body and a Luer connector. The tube body is long and thin, and the head can have end holes or side holes. Specifications; usually reusable instruments; used for aspirating ventricular fluid for examination or drainage of ventricular fluid during intracerebral surgery. | II |  | 150
151 | 03-12-01 | Cardiac head | The intracardiac suction head is made of metal, consisting of a head end, a tube body and a joint. The head end can be designed with various angles and shapes to suit the suction of different anatomical positions. It can be divided into several types according to different materials, designs and technical parameters. Model and specification; reusable; matched with suction device, used for suctioning pericardial/cardiac fluid during cardiac surgery. | II |  | 151
152 | 03-12-01 | Vascular Irrigation/Suction Tube | The vascular flushing/suction tube is made of stainless steel and other metal materials. It is composed of a head end, a tube body and a joint. It can be divided into several models and specifications according to different materials, designs, technical parameters, etc.; it can be used repeatedly; the flushing device is matched for surgical operation It is used for perfusing/suctioning and flushing blood vessels in vivo or in vitro. | II |  | 152
153 | 03-12-02 | Catheter Subcutaneous Tunneling Tool | This product is composed of cleaning rod, handle, obturator and outer casing and other components. The product is sterilized and disposable. Used to create a subcutaneous tunnel to help place a catheter under the skin that connects to an implanted pump. | II |  | 153
154 | 03-13-00 | balloon guide catheter | It can be used to establish intraoperative vascular access, deliver interventional equipment, and temporarily block blood vessels at the same time. It can be composed of catheter, balloon, seat, etc. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 154
155 | 03-13-01 | Contrast Catheter | For injection or infusion of control media and/or fluids. It usually consists of a tube body, catheter tip, radiopaque marker and seat, and can have accessories and can be coated. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III |  | 155
156 | 03-13-02 | Guiding catheter | For the establishment of intraoperative vascular access, or for the injection or infusion of contrast media and/or fluids and/or embolic materials. It usually consists of a tube body, catheter tip, radiopaque marker and seat, and can have accessories and can be coated. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III |  | 156
157 | 03-13-04 | Catheter disinfection connector (infusion joint disinfection cap, alcohol protection cap) | usually consists of joints, 70% isopropanol or ethanol aqueous solution, etc. It is not only used for physical barrier infusion (injection) equipment needle-free joints (including outer cone lock joints and inner cone lock joints); it is also used to sterilize the needle-free joints of infusion (injection) equipment. | II | Revised product management category description. | 157
158 | 03-13-05 | perfusion catheter | Usually consists of a catheter tube, catheter tip, radiopaque marker, catheter hub, and may have a side port at the distal end of the catheter. Available with coating and accessories. For intravascular delivery of various diagnostic and therapeutic solvents. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 158
159 | 03-13-06 | balloon dilatation catheter | For narrowing of arteriovenous vessels or endovascular stents (implantable tubular structures used to support the lumen of vessels), endovascular prostheses, stenosis, or artificial arteriovenous dialysis fistulas. Can consist of balloon (except special designs), catheter, catheter adapter and radiodetectable markers, can have accessories, can be coated. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. Special designs such as spinous process balloons, cutting balloons, etc. are excluded. | III |  | 159
160 | 03-13-06 | Anchored Balloon Dilatation Catheter | The anchoring balloon dilatation catheter is used in percutaneous coronary angioplasty. Through the expansion of the balloon, the guide wire in the guide catheter is fixed to realize the exchange of other interventional devices. It is mainly composed of a balloon, a catheter and a tube seat. The anchoring balloon has a single-lumen Luer connector that is used to inflate and deflate the balloon. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III | Add | 160
161 | 03-13-08 | contrast balloon | the loading material of the control heart occluder into the blood vessel for intravascular angiography or pressure measurement. It usually consists of balloon, radiopaque marker, catheter tube body, seat and other structures, which can have accessories and can be coated. It is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 161
162 | 03-13-08 | Contrast Balloon Catheter | Usually consists of a balloon, radiopaque markers, joints and other structures. It is used to inject contrast fluid into blood vessels for intravascular angiography. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 162
163 | 03-13-09 | Occlusion Balloon Catheter | It usually consists of a catheter, a balloon, and a connector. It is used to temporarily block peripheral blood vessels or nerve vessels, and can also selectively block or control blood flow. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 163
164 | 03-13-10 | thrombus aspiration catheter | It usually consists of a catheter, a joint, and a suction device. The catheter has a lumen for aspiration of intravascular substances. Used to aspirate blood clots in blood vessels. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structure design and mechanism of action; except drug-device combination products; except neurovascular applications. | III |  | 164
165 | 03-13-12 | transseptal puncture system | A septal puncture system usually includes a septal puncture needle and a septal puncture sheath. It is suitable for interatrial septal puncture to establish both pathways. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Drug-device combination products are excluded. | III |  | 165 | New
166 | 03-13-12 | Puncture needle | Consisting of a puncture needle tube and a tube base, it establishes a percutaneous access pathway that facilitates intravascular devices. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | II | Revised product management category description. | 166
167 | 03-13-13 | Guide cannula | Used to insert a catheter or guide wire, used in conjunction with a puncture needle. It usually consists of a pipe body and a joint. Available with accessories and coated. It is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 167
168 | 03-13-14 | Introducer sheath | It is used in conjunction with dilators to insert medical devices such as guide wires and catheters into blood vessels. It usually consists of components such as a sheath tube, a tube seat, and side branches, and can be equipped with accessories such as a hemostatic valve and can be coated. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | II | Revised product management category description. | 168
169 | 03-13-14 | hemostasis valve | Auxiliary instrument used in percutaneous angioplasty. It can be composed of hemostatic valve shell, valve, locking joint, side pipe, three-way cock and other components. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. | II | Revised product management category description. | 169
170 | 03-13-15 | dilator | Consists of dilator tube and socket. The Seldinger technique is used to establish a percutaneous access pathway that is helpful for intravascular devices, and does not have the role of intravascular positioning or establishment of intravascular access. Vascular puncture in Seldinger multi-digit digital subtraction angiography. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | II | Revised product management category description. | 170
171 | 02-11-03 | dilator | Consists of dilator tube and socket. Airway access was established by percutaneous puncture. The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. | II | Revise product management categories and classification code descriptions. | 171
| #N/A |  | The exemption does not include: the use of new materials; products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application.
172 | 03-13-15 | expansion sheath | The dilation sheath consists of an inner sheath and an outer sheath. Suitable for percutaneous dilation of tissue surrounding cardiac lead wires, indwelling catheters, and foreign objects. | II | Revised product management category description. | 172
| #N/A |  | The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market.
| #N/A |  | The exemption does not include: the use of new materials; products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application.
173 | 03-13-16 | guide wire | It is used to guide the insertion of other devices into blood vessels, establish percutaneous access pathways that facilitate intravascular devices, or perform intravascular positioning, or establish intravascular access. It can be composed of core wire, wrapping wire, safety wire, sheath and other components, can have accessories, can be coated, and can include extended guide wire. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | II | Revised product management category description. | 173
174 | 03-13-17 | balloon pressurization device | It usually consists of an inflation device, a pressure gauge, connecting pipes and joints. Used in cardiovascular interventional operations, it is used in connection with the balloon dilation catheter, and the balloon is accurately inflated and deflated by the pressure value displayed by the pressure gauge, so as to achieve the purpose of expanding or contracting the balloon. | II |  | 174
175 | 03-13-18 | connection valve | It is used to connect the pipeline and establish a channel to assist the device to enter the human body while reducing the outflow of blood. There are usually two or more ports to which catheters or other devices entering the blood vessel can be attached, and accessories may be attached. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | II | Revised product management category description. | 175
176 | 03-13-19 | vena cava filter recovery unit | Used to remove a retrievable vena cava filter out of the body or to convert a convertible filter. It usually consists of a recovery snare matched with the vena cava filter and a recovery sheath, which can capture and remove the convertible vena cava filter and the retrievable vena cava filter under the imaging device. Available with accessories and coated. It is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 176
177 | 03-13-20 | Cardiac Occluder Loader | It is used to store the heart occluder, and after being connected with the sheath, the heart occluder can be pushed into the sheath. It usually consists of a catheter and a connector. The catheter is short and its length can cover the matched heart occluder. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Except for new structural design and mechanism of action. | III | Add | 177
178 | 03-13-21 | Cardiac Occluder Delivery Device | The cardiac occluder is delivered via the femoral vein/artery. It can be composed of a loader, a pusher, a guider, a snare, etc., and can have accessories and can be coated. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III | Revised product name and product description. | 178
179 | 03-13-22 | Intravascular Retrieval Device | It is used to capture and remove foreign objects that fall off after the failure of intravascular interventional devices through the catheter, such as stents, broken catheters and guide wires. It usually consists of a grasping device, a rod, a control device, and a sheath. It has the function of grasping foreign objects, and can have accessories and can be coated. It is equivalent to products that have been approved for domestic registration. Products are composed of mature materials, except for new structural design, new mechanism of action, drug-device combination products, and neurovascular applications. | III |  | 179
180 | 03-13-23 | remote protection device | It usually consists of a sheath tube, a guide wire, and a filter mesh structure. The sheath tube is generally made of polymer materials, the guide wire is generally made of stainless steel, the filter screen structure is generally made of nickel-titanium alloy, and the filter screen can be covered with a film. For containment and removal of embolic material during neurovascular interventional procedures. The product composition materials are mature, the mechanism of action is clear, and it is equivalent to the product that has been approved for domestic registration. Excluding new structural design and mechanism of action; excluding drug-device combination products. | III |  | 180
181 | 03-13-24 | Ring handle syringe | It is generally composed of a jacket, a 6% outer conical joint, a mandrel, a mandrel handle, a collar, a piston, and a sealing "O" ring. It is expected to be used for manual injection of contrast solution and drug solution in minimally invasive interventional therapy or diagnostic surgery. . The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 181
182 | 03-13-25 | High pressure contrast injection extension tube | The product consists of a catheter , a swivel Luer connector, an O -ring / fixed Luer connector , and a seat. Single-use product, supplied sterile. This product acts as a conduit for a bolus of contrast media during angiography. The exemption does not include products with new materials, new mechanisms of action, or new functions. | II | Revised product management category description. | 182
183 | 03-13-26 | microcatheter | For injection or infusion of control media and/or fluids and/or embolization materials, and/or appropriate devices (such as stents, coils), etc. It usually consists of components such as a pipe body and a seat, which can have accessories and can be coated. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III |  | 183
184 | 03-14-02 | Disposable cardiometer stabilization plate | Disposable heart meter stabilization plate consists of stabilization plate (Chinese style, left style, right style, plastic), arm (stainless steel), locking bracket, locking bracket knob (ABS), base and arm adjustment knob, which can be customized according to material, design, Different technical parameters are divided into several models and specifications; provided in sterile form; used to stabilize the local heart surface during heart bypass surgery and provide conditions for coronary anastomosis. | II |  | 184
185 | 03-14-02 | Disposable Heart Holder | When used in heart bypass surgery, a fixed surgical field is created at the target vessel in the heart. Divided into two types, Stable-V1 (suction fixation type) and Stable-M1 (mechanical fixation type), the surgical area is fixed by the pressure and negative pressure suction of the fixation head or only by the pressure of the fixation head. Generally, it is composed of adjusting handle, fixed card seat, universal support arm, suction cup, negative pressure suction tube, and can also be composed of presser foot leaf (or suction foot fixing seat, suction foot), negative pressure connecting pipe, three-way valve, directional chuck The assembly is composed of an adjustment arm, a fixed seat, an adjustment handle, a rotary handle, and an inner tension steel wire assembly. Supplied in sterile form; product does not enter the blood circulatory system. Also known as a heart stabilizer. | II |  | 185
186 | 03-14-07 | Cerebral Plane Positioning Ruler | It consists of three parts: coronal, sagittal, and middle plates. Single use sterile product. It is used for auxiliary measurement and positioning plane in brain surgery. | II |  | 186
187 | 04 | Passive Sterile Orthopedic Surgical Instruments | The passive orthopedic surgical instruments whose management category is Class I can be exempted from clinical evaluation if the management category is upgraded to Class II after they are provided in sterile packaging. | II |  | 187
188 | 04 | Meniscus Suture Surgical Instruments | The product consists of a meniscus suture needle, a guide tube, and a meniscus file. The main body should be made of stainless steel that meets the standards of ASTM F899, ASTM A269, and ASTM A564. | II |  | 188
189 | 04 | Passive Endoscopic Orthopedic Surgical Tools | Orthopedic surgical tools for temporary use under passive endoscopy. Generally made of metal materials or polymer materials. It can be divided into several models and specifications according to different materials, designs, technical parameters, working parts, and intended uses. It is used with endoscope for orthopedic surgery and inspection. | II |  | 189
190 | 04-01-01 | osteotomy knife | Usually consists of a blade and a handle, with a hard, sharp, cutting blade in a single-edged configuration at the distal end, and a handle at its proximal end. Generally made of stainless steel. Supplied sterile. Used for resection and amputation of bone. | II |  | 190 | New
191 | 04-01-02 | arthroscopic scalpel | It usually consists of a head, a rod and a handle. The head is a cutting edge piece, which transmits and controls the work of the head through the operation of the handle. The head is generally made of stainless steel. It is operated under the endoscope during the operation, and is used in minimally invasive orthopedic surgery to resect and strip the diseased tissue. | II |  | 191
192 | 04-01-03 | Spine Surgery Reamer | These products are usually made of stainless steel. For reaming, reaming, or removing the cartilage layer of the endplate during posterior spinal surgery. Supplied sterile. | II |  | 192
193 | 04-01-03 | discectomy knife | The discectomy knife is made of stainless steel that complies with YY/T 0294.1 or ASTM F899. Used to cut the outermost annulus of the intervertebral disc. Supplied sterile. | II |  | 193
194 | 04-02-01 | Arthroscopy Scissors | It usually consists of a head, a rod or a flexible catheter and a handle. The head is a pair of blades with blades, which transmit and control the work of the head through the operation of the handle. Generally the head is made of stainless steel. During surgery, an endoscopic operation is used to cut tissue. | II |  | 194
195 | 04-03-01 | Arthroscopy forceps | It usually consists of a head, a rod or a soft catheter and a handle. The head is a pair of blades with a pincer beak, and the work of the head is transmitted and controlled through the handle. The head is generally made of stainless steel. Operated under the endoscope during surgery, it is used to clamp tissue or instruments. | II |  | 195
196 | 04-03-02 | Scoliosis Correction Pliers | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1127 rongeur. It is used for cutting bone tissue and repairing bone during spinal surgery. Supplied sterile. | II |  | 196
197 | 04-03-03 | Vertebral rongeur | Vertebral rongeurs are made of applicable metal materials (such as 40Cr13, 32Cr13Mo or 40Cr13Mo for the clamp head, 20Cr13, 30Cr13 or 32Cr13Mo for the clamp handle, strontium shaft screw, spring and spring pulley, and 12Cr18Ni9 for the spring screw. or 06Cr19Ni10), which can be composed of pliers head, pliers handle, strontium shaft screw, support spring screw, support spring, support spring pulley, etc., and can be divided into several types such as single joint, double joint, and side angle. There are several types of different forms, such as straight head, elbow head and side angle head. It can also be divided into several types according to the shape of the pliers head, such as round and pointed. According to the type of pliers handle, it can be divided into two types: shouldered and shoulderless. According to the opening of the cutting edge, it can be divided into two types: normal crotch opening and opening; each type is divided into several specifications according to the size of the mouth width and overall length; it is used for biting and cutting decayed dead bones and repairing bones in orthopedic operations (according to the design) Different can bite and cut the skull, lamina or nasal septum, etc.). Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T1127-2006 rongeur. Supplied sterile. | II |  | 197
198 | 04-03-03 | Cervical rongeur | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1127 rongeur. It is used for biting and cutting bone tissue and repairing bone during cervical spine surgery. Supplied sterile. | II |  | 198
199 | 04-03-03 | Cervical double joint rongeur | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1127 rongeur. It is used for biting and cutting bone tissue and repairing bone during cervical spine surgery. Supplied sterile. | II |  | 199
200 | 04-03-03 | Elbow Flat Spinous Process Bone Forceps | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1127 rongeur. It is used for biting and cutting bone tissue and repairing bone during orthopedic surgery. | II |  | 200
201 | 04-03-03 | gun shaped rongeur | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T1127 rongeur. It is used for biting and cutting bone tissue and repairing bone during orthopedic surgery. | II |  | 201
202 | 04-03-04 | Nucleus forceps | Nucleus pulposus forceps are composed of forceps head, cover plate, forceps body, left forceps handle, and support spring. According to different types, there are two types: straight head and elbow head. For soft tissues such as nuclei. Supplied sterile. | II |  | 202
203 | 04-03-05 | Posterior Spinal Spreader | The working principle of this type of product is that the width of the head end of the instrument is greater than the height. When using it, it is inserted into the intervertebral space in a positive direction, and then the instrument is rotated 90 degrees to make the width direction change from horizontal to vertical, thereby expanding the intervertebral space. Usually made of stainless steel. Supplied sterile. It can be divided into straight type and curved type, and there are many sizes. reusable. Used in posterior spinal surgery to open the intervertebral space. | II |  | 203
204 | 04-03-05 | Laminar distractor | These products are usually made of stainless steel, and are used to open the intervertebral space by spreading the lamina during operations such as spinal fusion. Supplied sterile. | II |  | 204
205 | 04-04-01 | Nerve hooks for spinal surgery | These products, usually made of stainless steel, are used in spinal surgery to block nerve roots and facilitate placement of fusion cages. Supplied sterile. | II |  | 205
206 | 04-04-01 | Nerve Retractors for Spine Surgery | These products are usually made of stainless steel and are used to retract nerve roots or soft and dural sacs in spinal surgery. Supplied sterile. | II |  | 206
207 | 04-05-01 | Surgical probes for arthroscopy | It usually consists of a head, a rod or a soft catheter and a handle. The head is needle-shaped, and the work of the head is transmitted and controlled through the handle. The head is generally made of stainless steel. Operate under an endoscope during surgery to explore tissues and take samples. | II |  | 207
208 | 04-05-02 | Sterile Bone Traction Needle | It usually consists of a head, a needle body and a tail, and can be divided into two types: threaded type and polished rod type. It is generally made of stainless steel or titanium alloy. Supplied sterile. | II |  | 208
209 | 04-06-01 | arthroscopic surgical curette | It usually consists of a head, a rod or a soft catheter and a handle. The head is spoon-shaped, and the work of the head is transmitted and controlled through the handle. The head is generally made of stainless steel. It is operated under the endoscope during the operation and is used in minimally invasive orthopedic surgery to scrape the diseased tissue. | II |  | 209
210 | 04-06-02 | spatula for spinal surgery | These products are usually made of stainless steel. It is used to scrape the nucleus pulposus and fibrous ring during spinal surgery such as intervertebral fusion. Supplied sterile. | II |  | 210
211 | 04-06-02 | spinal surgery curette | These products are usually made of stainless steel. It is used to scrape the crushed intervertebral disc, remove the upper and lower endplates, etc. during posterior spinal surgery. Supplied sterile. | II |  | 211
212 | 04-06-02 | Curette for Vertebroplasty | Such products usually consist of a head and a handle. There is a handle at the proximal end and a sharp-edged, spoon-shaped concave point at the distal end, which may also be double-ended. It is generally made of stainless steel and can be reused. It is used to scrape away lesions, scars in sinus tracts, granulation tissue, and sequestrum or pathological tissue in bone cavities and potential cavities. Supplied sterile. | II |  | 212
213 | 04-07-01 | Orthopedic minimally invasive interventional bone awl | Usually consists of a guide wire, a hollow drill with a handle, and a working channel. Metal parts are generally made of stainless steel, and handles are generally made of engineering plastics. Supplied sterile. It is used in orthopedic minimally invasive interventional surgery (percutaneous vertebroplasty, kyphoplasty, etc.) to drill holes and establish working channels. | II |  | 213
214 | 04-07-02 | spinal surgery hand awl | The heads of such products are usually made of stainless steel. A hole is made in the bone for spinal surgery. Supplied sterile. | II |  | 214
215 | 04-08-01 | Bone drill for posterior spinal surgery | This type of product usually consists of a handle and a metal part. The handle is usually made of plastic and the metal part is usually made of stainless steel. Used to drill bone tunnels in posterior spinal surgery. | II |  | 215
216 | 04-08-01 | Trephine | Usually a hollow structure. It is generally made of stainless steel that meets the requirements of YY/T 0294.1 or ASTM F899. It is used for bone resection or bone extraction in spinal surgery. | II |  | 216
217 | 04-08-02 | Aseptic grinding head | Usually consists of a head, blade and handle. It is generally made of diamond or stainless steel, tungsten carbide, or emery. Supplied sterile. It is used for opening or grinding bone tissue. | II |  | 217
218 | 04-08-02 | Arthroscopy Surgical Drills | It usually consists of a head and a handle, with a cutting edge on the head. The head is generally made of stainless steel, tungsten carbide, and emery, and the handle is generally made of stainless steel, titanium or synthetic materials. It is operated under endoscope during operation, used in minimally invasive joint surgery, used for drilling and tapping. | II |  | 218
219 | 04-09-01 | Trepanation saw for intervertebral disc surgery | This type of product is a circular surgical cutting instrument with a serrated edge and is generally made of stainless steel. It is used for resection and correction of carrion, drilling, decompression or removal of broken nails during intervertebral disc surgery. | II |  | 219
220 | 04-09-01 | Sterile bone saw | A surgical cutting instrument, usually flat or wire, with a serrated edge. Available with attached handle. Generally made of stainless steel. Supplied sterile. Does not contact the central nervous system. Used for sawing bones. | II |  | 220
221 | 04-10-01 | Vertebral osteotome | Vertebral osteotome is usually made of stainless steel, and the handle can be made of other suitable materials. It is usually a hand-held manual one-piece surgical instrument with a single edge. Different designs and technical parameters are divided into several models and specifications; surgical mallets or hammers can be used to manually transmit force from the proximal end to the distal end of the instrument, for chiseling, trimming, stripping cervical vertebrae, lumbar vertebrae and other vertebrae during surgery and making the vertebrae forming. Supplied sterile. | II |  | 221
222 | 04-10-01 | Cervical Osteotome | These products are usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 91141 Orthopedic Chisel General Technical Specifications. It is used for the cutting and trimming of bone tissue during cervical spine surgery. Supplied sterile. | II |  | 222
223 | 04-10-01 | Laminar osteotome | This type of product usually consists of a handle and a cutter head, which is a beveled sharp edge. Usually made of stainless steel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 91141 Orthopedic Chisel General Technical Specifications. For cutting and trimming of lamina during spinal surgery. Supplied sterile. | II |  | 223
224 | 04-11-01 | spinal rasp | These products are usually made of stainless steel. It is used for eradicating bone fragments and correcting bones during spinal surgery. Supplied sterile. | II |  | 224
225 | 04-11-01 | bone rasp for spinal surgery | This type of product is usually made of stainless steel and is used for the treatment of bone surfaces such as endplates and vertebral body edges in posterior spinal surgery. Supplied sterile. | II |  | 225
226 | 04-11-01 | Arthroscopy Surgical Rasp | Usually consists of a file body and a handle. The file body is generally made of stainless steel and emery. Operated under the endoscope during the operation, it is used in the minimally invasive surgery of the joint to grind the tissue. | II |  | 226
227 | 04-11-02 | spinal spatula | These products are usually made of stainless steel. It is used for eradicating bone fragments and correcting bones during spinal surgery. Supplied sterile. | II |  | 227
228 | 04-12-01 | Orthopedic powered surgery equipment | It usually consists of a host, a flexible shaft, a cable, a mobile phone and a tool (or a foot switch), or a mobile phone, a tool, a battery and a control system. Used for drilling, cutting, sawing, grinding, milling, etc. of bone tissue. | II |  | 228
229 | 04-13-02 | Orthopedic External Fixator | Orthopedic external fixation brackets are usually composed of needle-bar linkage devices, rods, rings, connecting parts, etc., excluding bone needles. Generally made of aluminum alloy, stainless steel, carbon fiber, titanium alloy materials, it can be composed of multiple brackets and fixed parts, it can be single arm or ring, and it can have a universal structure. It can be based on material, design, technical parameters, and applicable parts. It is divided into several models and specifications according to different types and intended uses; it can be used to connect the bone needles in the body to achieve the therapeutic purposes of orthopedic deformity correction and fracture reduction through fixation, compression or pulling of bone ends. | II |  | 229
230 | 04-14-02 | Disposable/Sterile Bone Cement Filler | The disposable/sterile bone cement filler can be composed of a bone cement glue gun (a gun-like structure composed of a metal handle, a mechanism screw or similar structure), a detachable nozzle sleeve (the two ends of the sleeve have threaded connection structures, which can be pressed Installed on the bone cement glue gun, the sleeve can be made of metal or polymer material), bone cement preparation container, bone cement spoon, stirring rod, sleeve seat, etc. Metallic parts are usually reused, non-metallic parts are usually single use. Provided aseptically or for one-time use, it is used for mixing bone cement in orthopedic surgery and injecting (delivering) the bone cement into the bone marrow cavity to be filled. Includes Low Viscosity Bone Cement Filler Kit. | II |  | 230
231 | 04-14-02 | bone cement filling kit | Sterile bone cement filler, one-time use bone cement filler can be composed of bone cement glue gun (gun-shaped composed of metal handle, mechanism screw or similar structure), detachable nozzle sleeve (threaded at both ends of the sleeve, etc.) structure, can be installed on the bone cement glue gun, the sleeve can be made of metal or polymer material), bone cement preparation container, bone cement spoon, stirring rod, sleeve seat and so on. Metallic parts are usually reused, non-metallic parts are usually single use. Provided aseptically or for one-time use, it is used for mixing bone cement in orthopedic surgery and injecting (delivering) the bone cement into the bone marrow cavity to be filled. Includes Low Viscosity Bone Cement Filler Kit. | II |  | 231
232 | 04-14-02 | Bone cement compression plug | This type of product usually consists of a femoral end compression plug and an acetabular compression plug, which are made of polyurethane and silicon 6810 rubber respectively, and are used for pressure shaping after bone cement is injected into the medullary cavity or acetabulum. Single use, sterile packaging. | II |  | 232
233 | 04-14-03 | Bone Graft Funnel for Posterior Spinal Surgery | These products, usually made of stainless steel, are used to introduce bone graft into the grafted area during posterior spinal fusion. Supplied sterile. | II |  | 233
234 | 04-14-04 | bone sampler | It usually consists of a sleeve, a tube core and a handle. The casing and core are generally made of stainless steel. For bone tissue sampling in percutaneous kyphoplasty/percutaneous vertebroplasty. Supplied sterile. | II |  | 234
235 | 04-14-05 | Trial models for spinal surgery | These products are usually made of stainless steel and are used to evaluate the height of the intervertebral space in operations such as posterior spinal fusion and provide reference for the selection of intervertebral fusion devices. Supplied sterile. | II |  | 235
236 | 04-14-05 | Depth finder for posterior spinal surgery | These products are usually made of stainless steel and are used to measure the depth of the cavity after the removal of intervertebral tissues such as the nucleus pulposus in posterior spinal surgery. | II |  | 236
237 | 04-15-01 | reamer | These products are divided into two categories, one is composed of flexible rods, drill bits, guide needles and needle feeders; the other is composed of reaming heads, tubular assembly parts, filtering devices, sealing heads, locking clips, etc. It is used to enlarge the medullary cavity during orthopedic surgery. For use with active instruments. | II |  | 237
238 | 04-16-01 | Bone cement mold for knee joint | It is generally made of medical grade silicone and does not include components such as reinforcement handles implanted in the body. For bone cement setting of the knee joint. Supplied sterile. | II |  | 238
239 | 04-16-01 | Bone cement mold for hip joint | Composed of a reinforced handle made of stainless steel and a bone cement setting mold made of medical grade silicone rubber, it is a single-use, sterile product. For shaping temporary hemiarthroplasty prostheses in skeletally mature patients requiring secondary revision due to infection. | III |  | 239
240 | 04-16-03 | 3D printed osteotomy guide | Orthopedic surgery accessory tool for matching patient anatomy. It is made of polymer polyamide material or metal material, and is made by 3D printing. Supplied non-sterile, single use after sterilization. Intended for positioning, guidance and protection in orthopedic procedures for specific patients. | II |  | 240 | New
241 | 04-17-04 | Vertebroplasty Auxiliary Instruments | The main body of metal material instruments is usually made of stainless steel; the filling device (pressure charging device) is made of polymer materials such as polycarbonate and polyurethane; the bone cement injector can be made of polymer materials or metal materials; , expansion set, working sleeve, bone cement injection set and filling device, etc., excluding vertebral body expansion balloon catheter. It is used in percutaneous kyphoplasty/percutaneous vertebroplasty to establish a working channel, fill bone cement, and restore the anatomical structure of the vertebral body. | II |  | 241
242 | 04-17-05 | Disposable fiber ring stapler | This type of product usually consists of a stapler housing, a suturing component and a transmission component. It is used for annulus fibrosus suture after simple disc herniation and nucleus pulposus surgery. | II |  | 242
243 | 04-17-06 | Vertebral Posterior Processor | This type of product usually consists of guide wire positioning, expansion sleeve, high-precision drill, working sleeve, etc. It is used to treat posterior osteophytes and hard protrusions, and to dredge nerve root channels. | II |  | 243
244 | 04-17-13 | spinal surgery access system | The supporting surgical tool for spinal surgery is generally composed of an expandable channel, a step-by-step expansion sleeve, a expansion channel, an inner core and a depth gauge. Made of medical polymer materials. Supplied sterile. For the establishment of spinal surgery access. | II |  | 244
245 | 04-17-15 | Bone Graft Pusher for Posterior Spinal Surgery | These products, usually made of stainless steel, are used in posterior spinal fusion to push and compact the implanted bone into the grafted area. | II |  | 245
246 | 04-17-15 | cage holder for spinal surgery | This type of product is usually made of stainless steel and is used to grasp the fusion device in posterior spinal fusion to complete the implantation or removal of the fusion device. | II |  | 246
247 | 04-18-01 | Lamina stripper | These products are usually made of stainless steel. Used to strip laminar attachment tissue during spinal surgery. | II |  | 247
248 | 04-18-01 | vertebral anterior stripper | These products are usually made of stainless steel. It is used to strip or separate mucosa, periosteum and other tissues during spinal surgery. | II |  | 248
249 | 04-18-02 | Cranioplastic material forming molds | Such products usually consist of positive and negative molds for material formation (mainly made of medical silicone rubber materials), and auxiliary tools for cranioplasty. Used as a mold for cranioplasty materials. | II |  | 249
250 | 04-18-02 | Infant skull orthopedic fixer | Consists of shell, filling material/pad and fixture. Generally made of polymer materials. It is an orthopedic device used for infants between 3-18 months to prevent or treat congenital or acquired asymmetry or plagiocephaly and brachycephaly by applying a certain amount of pressure to the contact parts of the infant's head. | II |  | 250
251 | 05-04-01 | X-ray beam scanning measurement system - water phantom system | It usually consists of a water tank, a control unit, a detector, control software, cables, etc. It is used to measure the absorbed dose distribution of the ray beam in water and the kerma energy distribution in the air. The measurement results are used for data configuration and modification of the radiotherapy planning system and quality control of radiotherapy equipment. | II |  | 251
252 | 05-04-01 | X-ray beam scanning measurement system - dose measurement matrix | It usually consists of a detector array, a measurement control system, and control software. | II |  | 252
| #N/A |  | It is used to measure the dose distribution of the radiotherapy plan executed by the radiotherapy equipment, and to compare it with the data of the planning system, and the comparison results are used as the basis for the verification and modification of the planning system.
253 | 05-04-06 | Radiation therapy laser positioning system | It usually consists of multiple fixed or mobile laser lights, etc. | II |  | 253
254 | 05-04-07 | Radioactive Seed Implantation Protector | It usually consists of protective parts, accessories, etc., and is used to protect the operator and other surrounding people. | II |  | 254
255 | 05-04-08 | Brain tumor radiotherapy positioning device | Usually consists of a frame with head rest and front piece, dental pad, thermoplastic mask, vacuum pad, patient control unit, adapters, stereotaxic frame, repeat positioning check tool, etc. For immobilization, positioning and repositioning of the head during stereotactic radiation therapy. | II |  | 255
256 | 06-01-02 | Urinary X-ray machine | It usually consists of an X-ray generating device, an image display system, and a dedicated urination bed. X-ray fluoroscopy equipment. It is used for X-ray fluoroscopy and photography in urology, gynecology, and gastrointestinal tract to obtain images for clinical diagnosis. Product does not include digital tomography capabilities. | II |  | 256 | New
257 | 06-01-03 | mammography machine | It usually consists of a frame, an X-ray generating device, a breast compressor, and an image receiving device. Digital products also come with workstations and display systems. Generally, materials such as molybdenum or rhodium are used to make the target surface of the X-ray tube. Use a lower tube voltage to generate low-energy X-rays for photography. It is used to take pictures of human breast tissue and obtain images for clinical diagnosis. | II |  | 257
258 | 06-01-04 | Dental X-ray machine | It usually consists of an X-ray generating device and its supporting components. For use with an intraoral image receiver. It is used for X-ray photography of teeth to obtain images for clinical diagnosis. | II |  | 258
259 | 06-01-05 | Perspective photography X-ray machine | It is usually composed of X-ray generating device, image display system, patient support device, and some are equipped with special gastrointestinal bed, compressor, film spotting device, tomography device, etc. It is used to assist gastrointestinal diagnosis, and has both photography and X-ray equipment with fluoroscopy function. During use, the region of interest is generally locked through perspective dynamic images, and then a clear static image of the region is obtained by photography with a larger dose. It is used for routine and gastrointestinal fluoroscopic examination, and can be equipped with urography function to obtain images for clinical diagnosis. | II | Revised product description. | 259
260 | 06-01-05 | Perspective photography X-ray machine | It is usually composed of X-ray generating device, image display system, patient support device, and some are equipped with special gastrointestinal bed, compressor, spotting device, etc. It is an X-ray that is used to assist gastrointestinal examination and has both photography and perspective equipment. During use, the region of interest is generally locked through perspective dynamic images, and then a clear static image of the region is obtained by photography with a larger dose. It is used for routine and gastrointestinal fluoroscopic examination, and has the function of digital subtraction angiography of peripheral blood vessels to obtain images for clinical diagnosis. Product does not include digital tomography capabilities. | III | Revised product description and product management category description. | 260
261 | 06-01-06 | Mobile C-arm X-ray machine | It usually consists of a mobile C-arm bracket, an X-ray generator, an image intensifier TV system or a digital flat panel detector imaging system, etc. It is used for surgical fluoroscopy and photography to obtain images for clinical diagnosis. | II |  | 261
262 | 06-01-06 | Mobile C-arm X-ray machine | It usually consists of a mobile C-arm bracket, an X-ray generator, an image intensifier TV system or a digital flat panel detector imaging system, etc. It is used for surgical fluoroscopy and photography, and has the function of digital subtraction angiography of peripheral blood vessels to obtain images for clinical diagnosis. Product does not include digital tomography capabilities. | III |  | 262
263 | 06-01-07 | Photographic X-ray machine | It usually consists of an X-ray generating device and ancillary equipment for photographic X-rays. The digital system also has an X-ray detector and its imaging system. Utilizing the principle of different radiation attenuation when the X-rays emitted from the X-ray tube pass through different tissues and organs of the patient's body, the image formed by projecting the X-rays that pass through the patient and carry sufficient information onto the imaging medium is transformed into a visible image. General-purpose X-ray equipment for planar grayscale images. It is used for routine photography of patients to obtain a single image for clinical diagnosis. | II |  | 263
264 | 06-01-08 | Perspective X-ray machine | It usually consists of an X-ray generating device, a fluorescent screen or a TV system or an image receiving device such as a dynamic detector, and may have a patient support device, etc. Utilizing the principle of different radiation attenuation of different tissues and organs of the human body, through continuous loading of the X-ray source, it is a general-purpose X-ray device that converts dynamic images on the imaging medium. It is used for routine fluoroscopy of patients and obtains continuous images for clinical diagnosis. | II |  | 264
265 | 06-01-09 | X-ray Densitometer | It usually consists of an X-ray generating device, a detector, an information analysis and display system, and may also have a patient support device. According to the different degrees of absorption of X-rays by bones and tissues of different densities, the received digital signals with human body information are input into the computer for analysis to obtain the results of bone density. It is used to evaluate the bone density and mineral content of the patient's bone and adjacent tissues by measuring the X-ray attenuation of the human body for clinical diagnosis. | II |  | 265
266 | 06-01-10 | Vehicle X-ray machine | An X-ray machine mounted on a movable vehicle with fluoroscopy and/or photography capabilities. It is used under mobile conditions to carry out X-ray fluoroscopy and photographic diagnostic examinations at sites far away from hospitals. | II |  | 266
267 | 06-01-11 | Portable X-ray machine | It usually consists of X-ray source components, image receiving devices, etc. An x-ray machine that can be carried by one or more persons from one place to another during use or between uses. It is used for X-ray imaging of limbs or other body parts with less attenuation for clinical diagnosis. | II |  | 267
268 | 06-03 | X-ray generation, beam limiting device | Used for X-ray generation and beam limitation. | II |  | 268
269 | 06-04-01 | X-ray image intensifier | The image intensifier is a photoelectric vacuum tube that converts the X-ray signal carrying patient information into a visible light image. Assembled in diagnostic X-ray machines for converting X-ray images into visible light images. | II |  | 269
270 | 06-04-01 | X-ray image intensifier TV system | X-ray image intensifier TV system is usually composed of X-ray image intensifier, optical system, camera signal processing system and image display device, which converts X-ray images into corresponding visible light images. Assembled in diagnostic X-ray machines for converting X-ray images into visible light images. | II |  | 270
271 | 06-04-02 | X-ray detector | X-ray detectors (including flat panel detectors or photocoupler (CCD) detectors, etc.) use specific photoelectric conversion media to convert X-ray signals passing through the human body into digital signals. Assembled or cooperated with diagnostic X-ray machine, it is used to convert X-ray signal into digital signal. | II |  | 271
272 | 06-04-02 | Digital X-ray Imaging System | Digital X-ray imaging systems generally include X-ray detectors, image transmission, processing and display systems. Assembled or cooperated with diagnostic X-ray machine, it is used to convert X-ray signal into digital signal. | II |  | 272
273 | 06-04-03 | Image plate imaging device | Digital X-ray imaging systems generally include X-ray detectors, image transmission, processing and display systems. Assembled or cooperated with diagnostic X-ray machine, it is used to convert X-ray signal into digital signal. | II |  | 273
274 | 06-05 | X-ray tomography device | X-ray tomography device can clearly capture the image of a certain layer or several layers of tissue parallel to the longitudinal axis of the human body, but makes the images of other layers blurred; it can be designed according to design, technical parameters, It is divided into several models with different additional auxiliary functions; it can be used for tomography of the human body in conjunction with the photography flat bed. | II |  | 274
275 | 06-05-01 | X-ray fluoroscopy bed | It usually consists of a bed body, a spotting device, a remote control device and/or a near-table operation device, etc. It is used for gastrointestinal X-ray examination and used with gastrointestinal X-ray machine. | II |  | 275
276 | 06-05-02 | X-ray catheter bed | It is usually supported by a single arm, and the bed panel is generally made of carbon fiber and other materials. When combined with C-arm, U-arm and other interventional X-ray equipment for multi-directional photography, the field of vision will not be hindered. It is used for general interventional therapy and digital subtraction angiography (DSA), and is used with angiographic X-ray equipment. | II |  | 276
277 | 06-05-03 | X-ray electric photography flat bed | Patient support device. Can be electrically translated, rotated, etc. For support of patients in radiographic imaging. | II |  | 277
278 | 06-05-04 | Suspension and support device attached to X-ray diagnostic equipment | It usually consists of a base, a column, a support arm or a suspension frame, and a guide rail. Can be electrically translated, rotated, etc. It is used for suspension and support of X-ray source components, image receiving devices and other components. | II |  | 278
279 | 06-05-04 | Chest stand | Chest film stand is an auxiliary equipment for medical X-ray diagnosis. It is composed of base, column, support arm and other components. It can be electrically translated and rotated. X-ray monitor and chest X-ray point-and-shoot device, providing suitable human body X-ray position and so on. | II |  | 279
280 | 06-05-05 | X-ray contrast agent injection device | It usually consists of a jacket, a piston and a piston seal, and accessories include connecting pipes, suction pipes, etc. Used together with high pressure injector. | III |  | 280
281 | 06-05-05 | Disposable high pressure contrast syringe and accessories | It consists of a syringe, a connecting tube, a drug suction tube, etc. Intended use of the product: To support CT high-pressure injection pumps, MRI high-pressure injection pumps, DSA high-pressure injection pumps, and devices for injecting contrast agents during ultrasound examinations and cardiovascular angiography, without touching the human body. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product classification code. | 281
#N/A
282 | 06-05-05 | Contrast Media Injection Device | It usually consists of injection head, control part and so on. Clear blood vessel images are obtained by injecting contrast agents into human blood vessels. It is used together with high-pressure syringe barrels and accessories. For X-ray contrast, CT imaging, injection of contrast agents, saline, etc. | II | Combined with the original 277 items to revise the product description. | 282
283 | 06-05-05 | Contrast Media Injection Device | It usually consists of a jacket, a piston and a piston seal, and accessories include connecting pipes, suction pipes, etc. Used together with high pressure injector. The exemption does not include products with new materials, new mechanisms of action, new structures, or new functions. | III |  | 283
284 | 06-05-06 | Anti-scatter grid | It usually consists of lead strips, medium, etc. A device placed in front of the image receiving surface to reduce the scattered radiation radiated onto the image receiving surface, thereby improving the contrast of the X-ray image. Used with X-ray machines to increase the contrast of X-ray images. | II |  | 284
285 | 06-05-07 | X-ray photography cassette | Units for X-ray photographic film, etc., with grids, in different sizes for different applications. | II |  | 285
286 | 06-06 | Medical radiation protection spray | The main components of medical radiation protection spray are superoxide dismutase and its stabilizer, which exist in the form of freeze-dried powder in a brown plastic bottle (bottle A) that can be covered with a spray pump, and other excipients (curcumin, sorbitol, sorbitol Potassium acid potassium) exists in a capped white plastic bottle (B bottle) in the form of aqueous solution, which can be divided into several models and specifications according to design, technical parameters, and loading; it is used to prevent and reduce the freedom of physical and chemical factors such as medical radiation Damage to human skin and mucous membrane tissue. | II |  | 286
287 | 06-06-01 | Medical radiation protection equipment | It is usually composed of thin and uniform soft lead rubber, lead rubber plastic, lead glass or other protective material sheets containing heavy metal elements, and is used for the protection of human body during radiation therapy. | II |  | 287
288 | 06-07-01 | Ultrasound Pulse Echo Imaging Equipment | Ultrasonic pulse echo imaging equipment utilizes the principle of ultrasonic pulse echo to complete the ultrasonic system of human organ tissue imaging. It usually consists of probes (linear array, convex array, phased array, mechanical sector scan), ultrasonic transmission/reception, signal processing, and image display; it can be classified according to model, imaging property (analog and digital), probe configuration, and technology. It is divided into several models according to different parameters, additional auxiliary functions, and intended use; ultrasound imaging of abdominal organs, heart, superficial tissues and organs through body surface, transvaginal, and transrectal, excluding unconventional application methods. Only basic ultrasound B-mode imaging is included, and imaging modes/functions (such as compound imaging, harmonic imaging) generated by modifying beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment, etc. | II |  | 288
289 | 06-07-01 | Ultrasound Pulse Echo Imaging Equipment | Ultrasonic pulse echo imaging equipment utilizes the principle of ultrasonic pulse echo to complete the ultrasonic system of human organ tissue imaging. It usually consists of probes (linear array, convex array, phased array, mechanical sector scan), ultrasonic transmission/reception, signal processing, and image display; it can be classified according to model, imaging property (analog and digital), probe configuration, and technology. Different parameters, additional auxiliary functions, intended use, etc. are divided into several models; ultrasonic imaging of abdominal organs, heart, superficial tissues and organs, etc., through body surface, transvaginal, transrectal or intraoperative, excluding unconventional and contact How the circulating blood is applied. Only basic ultrasound B-mode imaging is included, and imaging modes/functions (such as compound imaging, harmonic imaging) generated by modifying beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment, etc. | III |  | 289
290 | 06-07-02 | Ultrasound Spectral Doppler Diagnostic Equipment | Ultrasonic spectral Doppler diagnostic equipment uses the principle of ultrasonic Doppler to extract information such as the direction and speed of blood flow, and displays it in the form of spectrum-time or blood flow-time after processing. It is further divided into pulse wave Doppler and continuous wave Doppler. It usually consists of probes (unit type, convex array, linear array, phased array, mechanical sector sweep), ultrasonic transmission/reception, signal processing and spectrum display, etc.; ), probe configuration, technical parameters, additional auxiliary functions, intended use, etc., are divided into several models; detect blood flow in abdominal organs, heart, superficial tissues and organs through body surface, vagina, rectum or intraoperatively , excluding unconventional applications. Only the basic Pulse Wave Doppler (PW) and Continuous Wave Doppler (CW) modes are included, and the imaging modes/functions generated by modifying the beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0767-2009 Ultrasonic Color Blood Flow Imaging System, etc. | III |  | 290
291 | 06-07-02 | Ultrasonic color blood flow imaging equipment | Ultrasonic color blood flow imaging equipment uses Doppler and autocorrelation technology to image blood flow on the basis of ultrasonic pulse echo imaging, and superimposes color-coded information on B-mode gray-scale images for real-time display. It usually consists of probes (convex array, linear array, phased array, mechanical sector scan), ultrasonic transmitting/receiving circuits, signal processing, and image display; it can be divided into several models according to model, imaging properties (analog and digital), probe configuration, technical parameters, additional auxiliary functions, and intended use; it can be used to image blood flow in abdominal organs, heart, superficial tissues and organs through body surface, transvaginal, transrectal, or intraoperatively, excluding unconventional applications. Only the basic ultrasonic color blood flow imaging mode is included, and the imaging modes/functions (such as ultrasonic blood flow vector imaging) generated by modifying beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment, YY 0767-2009 ultrasonic color blood flow imaging system, etc. | III |  | 291
292 | 06-07-02 | Ultrasound Spectral Doppler Diagnostic Equipment | Ultrasonic spectral Doppler diagnostic equipment uses the principle of ultrasonic Doppler to extract information such as the direction and speed of blood flow, and displays it in the form of spectrum-time or blood flow-time after processing. It is further divided into pulse wave Doppler and continuous wave Doppler. It usually consists of probes (unit type, convex array, linear array, phased array, mechanical sector sweep), ultrasonic transmission/reception, signal processing and spectrum display, etc.; ), probe configuration, technical parameters, additional auxiliary functions, intended use, etc., are divided into several models; the blood flow of abdominal organs, heart, superficial tissues and organs, etc. are detected through the body surface, vagina, and rectum, excluding unconventional application. Only the basic Pulse Wave Doppler (PW) and Continuous Wave Doppler (CW) modes are included, and the imaging modes/functions generated by modifying the beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0767-2009 Ultrasonic Color Blood Flow Imaging System, etc. | II |  | 292
293 | 06-07-02 | Ultrasonic color blood flow imaging equipment | Ultrasonic color blood flow imaging equipment uses Doppler and autocorrelation technology to image blood flow on the basis of ultrasonic pulse echo imaging, and superimposes color-coded information on B-mode gray-scale images for real-time display. It usually consists of probes (convex array, linear array, phased array, mechanical sector scan), ultrasonic transmitting/receiving circuits, signal processing, and image display; it can be divided into several models according to model, imaging properties (analog and digital), probe configuration, technical parameters, additional auxiliary functions, and intended use; it can perform imaging of blood flow in abdominal organs, heart, and superficial tissues and organs through body surface, transvaginal, and transrectal, excluding unconventional applications. Only the basic ultrasonic color blood flow imaging mode is included, and the imaging modes/functions (such as ultrasonic blood flow vector imaging) generated by modifying beamforming, image pre-processing, and image post-processing algorithms are not included. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment, YY 0767-2009 ultrasonic color blood flow imaging system, etc. | II |  | 293
294 | 06-08-01 | Sterilized Medical Ultrasonic Couplant | Sterile medical ultrasonic coupling agent can be composed of disinfectant (such as triclosan), propylene glycol, glycerol (glycerin), triethanolamine, carbomer (cross-linked polyacrylic resin); provided sterile. It can be divided into several models and specifications according to different formulas, technical parameters, applicable parts, and filling volumes; it is used for filling or coating between the skin-mucosa and the radiating surface of the probe (or treatment head) during ultrasonic diagnosis or treatment operations. It is an intermediary medium for transmitting sound waves, which is used to improve the ultrasonic coupling effect between the probe and the patient's skin/mucosa, and has the function of disinfecting the skin and mucous membranes. | II |  | 294
295 | 06-08-01 | Medical ultrasonic gel for cavity | The medical ultrasonic coupling agent for the cavity is composed of Bock-DP (trichlorohydroxy diphenyl ether), propylene glycol, glycerol triethanolamine, and carbomer (cross-linked polyacrylic resin); it can be provided in sterile form; coating It is connected with the head end of the intracavity ultrasound probe to improve the ultrasound coupling between the intracavity mucous membrane and the probe. | II |  | 295
296 | 06-08-03 | Water bladder for ultrasonic endoscope | The water bladder for ultrasonic endoscope is made of natural rubber latex/silica gel and polytetrafluoroethylene, etc. It is composed of a water bladder and a water bladder outer sleeve. The front end is injected with sterile saline and expanded as an ultrasonic conduction medium; it can be divided into several types according to the material, capacity, shape, and use site; it is provided in a sterile form (the connector can be non-sterile); it is used for ultrasonic endoscopes It is used for coupling between the probe and the skin/mucosa of the inspected part during inspection. | II |  | 296
297 | 06-08-04 | Ultrasound Mechanical Scanning Probe | Ultrasonic mechanical scanning probes are usually composed of ultrasonic transducers and/or ultrasonic transducer array elements, acoustic lenses, housings, cables, etc. The transducer array elements are arranged in a plane or convex array, and can be arranged according to structure, type, technology Different parameters and functions are divided into several models, which are matched with ultrasound equipment, and can be realized through intact skin and cavity mucosa (such as vagina, rectum, excluding the ultrasound part of transesophageal ultrasound probe and upper and lower gastrointestinal ultrasound endoscope). Expected functions such as information collection and tissue imaging. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment. | II |  | 297
298 | 06-08-04 | Electronic Line Array Probe | Electronic linear array probes are usually composed of ultrasonic transducers and/or ultrasonic transducer array elements, acoustic lenses, housings, cables, etc. The transducer array elements are arranged in a plane, and can be selected according to structure, type, technical parameters and functions. They are divided into several models, which are matched with ultrasound equipment, through intact skin and cavity mucous membrane (such as vagina, rectum, excluding transesophageal ultrasound probe, ultrasound part of upper and lower gastrointestinal ultrasound endoscope) to realize biological information collection, Expected features like tissue imaging. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment. | II |  | 298
299 | 06-08-04 | Electronic convex array probe | Electronic convex array probes are usually composed of ultrasonic transducers and/or ultrasonic transducer array elements, acoustic lenses, housings, cables, etc. The transducer array elements are arranged in a convex array, which can be selected according to the structure, type, technical parameters and It is divided into several models according to different functions, which are matched with ultrasound equipment, and can realize biological information collection through intact skin and mucous membrane of cavity (such as vagina and rectum, excluding transesophageal ultrasound probe and ultrasound part of upper and lower gastrointestinal ultrasound endoscope). , tissue imaging and other expected functions. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 10152-2009 B-type ultrasonic diagnostic equipment. | II |  | 299
300 | 06-08-05 | Ultrasound puncture stand | The product is used with intracavitary or extracorporeal ultrasound probes, and is usually installed on the ultrasound probe, which can be used to fix the ultrasound probe and puncture needle, and guide the puncture needle to the target position of the human body under the guidance of ultrasound. The product is provided sterile. The puncture frame should be a passive medical device, and should not contain active structures and/or functions such as navigation modules and automatic puncture. | II | Merged with the original 298 to revise the product name and product description. | 300
301 | 06-10-01 | Magnetic Resonance High Pressure Contrast Injection System | Usually consists of an operator room assembly and a scan room assembly. The scan chamber components are made of non-ferromagnetic materials and the wires are RF shielded. Accompanying syringe is not included. It is used for intravenous injection of magnetic resonance contrast agent and/or injection of commonly used flushing fluid in human blood vessels to meet the needs of magnetic resonance scanning diagnosis. | II |  | 301
302 | 06-10-02 | Magnetic Resonance Assisted Stimulation System | It is usually composed of computer, operator display, fiber optic media converter and other components used in the control room and patient interface display, power supply, optical cable, button response unit, earphone and other components used in the MRI scanning room. For use with magnetic resonance imaging systems. Used to provide visual, motor, auditory and verbal task stimulation to patients by playing text, pictures, slides, sounds, CDs and DVDs while the brain is being scanned. | II |  | 302
303 | 06-11-01 | Gamma (γ) camera | It usually consists of a rack, collimator, detector, data acquisition, image processing workstation, system software, patient support device, etc. It is used to obtain the distribution of single-photon radionuclides in the whole body or some organs of the human body to form a plane image. | II |  | 303
304 | 06-11-04 | Radionuclide Scanner | It usually consists of host computer, computer, color printer, working software and isolated power supply. For human thyroid radionuclide scanning imaging. | II |  | 304
305 | 06-13-01 | Medical Infrared Camera | Usually consists of an infrared camera, a processing system, a stand and a display. Annotate thermal images of the human body by infrared photography. | II |  | 305
306 | 06-13-02 | Infrared Breast Examiner | Infrared mammary gland examination instrument consists of probe (halogen lamp probe, laser probe), camera, host (including display, printer, computer, software), and can be divided into several models according to design, type, technical parameters, additional auxiliary functions, etc.; Infrared light inspection for breast diseases. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0324-2008 Infrared Breast Examination Instrument. | II |  | 306
307 | 06-13-04 | operating microscope | Surgical microscopes are composed of four basic parts: optical system, frame, lighting system, and electrical device. They can also include camera/camera systems and image management systems. Several models; for (non-ophthalmological) surgical microsurgical magnification, lighting, image management, etc., excluding fluorescence modules and other special spectral uses. Refer to Part 1 of the standard GB 11239.1-2005 Surgical Microscope. | II |  | 307
308 | 06-13-05 | Microcirculation Microscope | Microcirculation microscope inspection instrument can be composed of microcirculation microscope, light source (LED light source/cold light source), CCD (black and white/color), display (monitor), etc., and can include image management software, which can be composed according to structure, technical parameters, and intended use It is divided into several models according to different auxiliary functions, etc.; it is used for checking the microcirculation of the human body (such as the nail fold), so as to evaluate the microcirculation status of the human body. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0067-2007 microcirculation microscope. | II |  | 308
309 | 06-14-01 | cystoscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. It is used for observation and imaging of urethra and bladder. | II |  | 309
310 | 06-14-01 | Rectal endoscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. It is used for observation and imaging of anus and rectum. | II |  | 310
311 | 06-14-01 | fiber optic endoscope | It usually consists of an optical cable connector, a flexible tube containing a guide wire, and distal and proximal optical lenses, which can be divided into several types according to different designs and technical parameters; the flexible tube is provided in a sterile form; it is matched with an adapted light source and used Observation and imaging of nasal and sinus structures. | II |  | 311
312 | 06-14-01 | Disposable Fiber Optic Laryngoscope | The disposable fiber optic laryngoscope consists of a plastic handle and a speculum (made of plastic) containing an anti-fog CCD camera and an optical fiber wire. The other end of the handle can be connected to a monitor. The handle and the speculum are usually detachable, and can be divided into several models and specifications according to the applicable population, material, design, technical parameters, speculum size, etc.; the product is provided in a sterile form; it is used to clinically provoke the patient's epiglottis to expose the glottis, guide medical staff to accurately perform airway intubation for anesthesia or first aid, and can also be used for intraoral diagnosis and treatment. Used to observe and examine the patient's upper airway and to assist in the insertion of respiratory tubes. | II |  | 312
313 | 06-14-01 | sinus endoscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. For viewing imaging in the sinuses. | II |  | 313
314 | 06-14-01 | Laryngoscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. For imaging of the larynx. | II |  | 314
315 | 06-14-01 | fiber optic nasopharyngoscope | Flexible fiberscopes are generally composed of an optical imaging system and an illumination system. The optical imaging system consists of three parts: an objective lens, an image transfer system, and an eyepiece. The lighting system is composed of multiple bundles of optical fibers arranged in a mixed arrangement, and the imaging system is composed of multiple bundles of optical fibers arranged in a regular manner. The objective lens directly focuses and images the object on the fiber surface, and each pixel (each light guide fiber) on the fiber surface array receives the light energy of the corresponding position image, and transmits the light energy to the other end of the image transmission fiber, and all the light energy of all pixels at the image side is recombined to form the focused image of the objective lens, which is then amplified by the eyepiece or camera system for observation. Used to observe imaging during examination and surgery of the nose and throat. | II |  | 315
316 | 06-14-01 | Fiberoptic Upper Gastrointestinal Endoscopy | Flexible fiberscopes are generally composed of an optical imaging system and an illumination system. The optical imaging system consists of three parts: an objective lens, an image transfer system, and an eyepiece. The lighting system is composed of multiple bundles of optical fibers arranged in a mixed arrangement, and the imaging system is composed of multiple bundles of optical fibers arranged in a regular manner. The objective lens directly focuses and images the object on the fiber surface, and each pixel (each light guide fiber) on the fiber surface array receives the light energy of the corresponding position image, and transmits the light energy to the other end of the image transmission fiber, and all the light energy of all pixels at the image side is recombined to form the focused image of the objective lens, which is then amplified by the eyepiece or camera system for observation. For observation and imaging of the upper gastrointestinal tract (excluding the duodenum). | II |  | 316
317 | 06-14-01 | fiberoptic lower gastrointestinal endoscopy | Flexible fiberscopes are generally composed of an optical imaging system and an illumination system. The optical imaging system consists of three parts: an objective lens, an image transfer system, and an eyepiece. The lighting system is composed of multiple bundles of optical fibers arranged in a mixed arrangement, and the imaging system is composed of multiple bundles of optical fibers arranged in a regular manner. The objective lens directly focuses and images the object on the fiber surface, and each pixel (each light guide fiber) on the fiber surface array receives the light energy of the corresponding position image, and transmits the light energy to the other end of the image transmission fiber, and all the light energy of all pixels at the image side is recombined to form the focused image of the objective lens, which is then amplified by the eyepiece or camera system for observation. For observation and imaging of the lower gastrointestinal tract. | II |  | 317
318 | 06-14-01 | fiberoptic bronchoscope | Flexible fiberscopes are generally composed of an optical imaging system and an illumination system. The optical imaging system consists of three parts: an objective lens, an image transfer system, and an eyepiece. The lighting system is composed of multiple bundles of optical fibers arranged in a mixed arrangement, and the imaging system is composed of multiple bundles of optical fibers arranged in a regular manner. The objective lens directly focuses and images the object on the fiber surface, and each pixel (each light guide fiber) on the fiber surface array receives the light energy of the corresponding position image, and transmits the light energy to the other end of the image transmission fiber, and all the light energy of all pixels at the image side is recombined to form the focused image of the objective lens, which is then amplified by the eyepiece or camera system for observation. For observation and imaging of the trachea and bronchi. | II |  | 318
319 | 06-14-01 | Laparoscopic | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three major systems: the objective lens system, the image transfer system, and the eyepiece system. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. For observation and imaging of the abdominal cavity. | II |  | 319
320 | 06-14-01 | Arthroscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. It is used for observation and imaging of joint parts. | II |  | 320
321 | 06-14-01 | rigid optical thoracoscopy | Rigid optical thoracic endoscope generally consists of optical imaging system and lighting system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. For observation and imaging of the thoracic cavity. | II |  | 321
322 | 06-14-01 | Ureteroscopy | Generally semi-rigid optical endoscope. It usually consists of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is composed of multiple bundles of light guide fibers arranged in a mixed arrangement, and the light is transmitted through an external light source to provide illumination for the endoscopic image. The imaging system is composed of regularly arranged bundles of light guide fibers. The observed object is transmitted to the eyepiece through the objective lens through the image guide fiber, and then enlarged by the eyepiece or camera system for observation. The outer tube is medical stainless steel. May contain working channel for compatible surgical instruments and suction perfusion. The product is inserted into the urethra and penetrates the ureter through the bladder for observation and imaging of the ureter. | II |  | 322
323 | 06-14-02 | Prostatectomy Endoscopy | It generally consists of a rigid optical endoscope, a sheath, a manipulator, and surgical electrodes. Under the observation of the endoscope, the prostate tissue is excised using high-frequency current. | III |  | 323
324 | 06-14-03 | Electronic Laparoscopy | The product is composed of objective lens end, lens body, optical connector and video connector. The image sensor (usually CCD or CMOS) at the head end converts the received optical signal into an electrical signal, which is processed by the image processing device and observed on the display. The product is used for observation and imaging in abdominal surgery. The product does not include a built-in light source. | II |  | 324 | NEW
325 | 06-14-03 | Electronic Upper Gastrointestinal Endoscopy | The electronic upper gastrointestinal endoscope is a flexible electronic endoscope, which generally consists of a head end, a bending part, an insertion part, and parts connected to electrical and light sources. The photoelectric conversion device at the head end converts the received optical signal into an electrical signal, which is observed on the display through the camera system. Images are provided through a video monitor for observation, diagnosis, and photography of the upper gastrointestinal tract (excluding the duodenum). | II |  | 325
326 | 06-14-03 | Electronic nose and throat endoscopy | Flexible electronic endoscopes are generally composed of a head end, a bending part, an insertion part, an operating part, and parts connected to electrical and light sources, and may contain a working channel. The photoelectric conversion device at the head end converts the received optical signal into an electrical signal, which is observed on the display through the camera system. Provide images through a video monitor for the observation, diagnosis and treatment of the upper respiratory tract such as the nose and throat. | II |  | 326
327 | 06-14-03 | electronic lower gastrointestinal endoscopy | Flexible electronic endoscopes are generally composed of a head end, a bending part, an insertion part, an operating part, and an electrical and light source connection part, and generally contain a working channel. The photoelectric conversion device at the head end converts the received optical signal into an electrical signal, which is observed on the display through the image processing system. Images are provided by a video monitor for observation, diagnosis and photography of the lower gastrointestinal tract (excluding the small intestine). | II |  | 327
328 | 06-14-03 | electronic cystonephroscope | Flexible electronic endoscopes are generally composed of a head end, a bending part, an insertion part, an operating part, and an electrical and light source connection part, and may contain a working channel. The photoelectric conversion device at the head end converts the received optical signal into an electrical signal, which is observed on the display through the image processing system. Provide images through video monitor for observation, diagnosis, photography and treatment of urethra, bladder and renal pelvis. | II |  | 328
329 | 06-15-01 | Cold light source for endoscopy | It usually consists of a cold light source host and a power cord, and can include a light guide. The bulb types include LED lights, xenon lamps, and halogen lamps. The cold light source host is connected to the endoscope through the light guide to provide illumination for endoscopic inspection or surgery. Does not contain special spectrum. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1081-2011 Cold light source for medical endoscope and endoscope function supply device. | II |  | 329
330 | 06-15-01 | strobe light | The stroboscopic light source is generally composed of a host, a light source, a microphone, and a foot switch. It is powered by a network power supply or a battery, and is usually used together with a laryngoscope and a monitor. It can be divided into several models according to the design, technical parameters and auxiliary functions. Strobe and normal light are provided for throat examination. | II |  | 330
331 | 06-15-02 | Endoscopic camera system | It is composed of a host, a camera, an adapter and a cable. It is used in conjunction with an optical endoscope and a monitor to convert the optical signal collected by the endoscope into an electronic signal and transmit it to the monitor for imaging. | II |  | 331
332 | 06-15-03 | Electronic endoscope image processor | The product is generally composed of a host and accessories. The accessories include power cords, interface cables, data keyboards, etc., and are used in conjunction with electronic endoscopes and monitors to process images collected by electronic endoscopes and send them to the monitors. | II |  | 332
333 | 06-15-04 | Insufflation machine | It consists of a host, a gas filter and connecting pipelines, and is used to inject carbon dioxide gas into the abdominal cavity during laparoscopic minimally invasive surgery, to establish and maintain an insufflation, and to provide space for surgical operations and vision. | II |  | 333
334 | 06-15-05 | Irrigation Aspirator for Endoscope | It consists of a main unit, connecting wire, foot switch, flushing circuit and suction circuit; this product is used for flushing and/or suction in endoscopic surgery. | II |  | 334
335 | 06-15-06 | uterine compression expander | The uterine pressurization and expansion device is composed of a pressurizer host, a pipeline and an inflation needle. It needs to have an overpressure alarm and an automatic deflation function; it can be divided into several models according to different designs, technical parameters, auxiliary functions, and intended uses; for gynecology During endoscopy and surgery, the uterus is pressurized and inflated to create a visible space. | II |  | 335
336 | 06-16-05 | endoscope tip cap | The tip caps for disposable endoscopes are made of silicone rubber, etc., and can be divided into several types according to different materials, technical parameters, and applicable endoscopes; they are provided in a sterile state; they are installed on the tip of the endoscope to maintain a proper Endoscopic view. | II |  | 336
337 | 06-16-06 | Biopsy bag for endoscope | The product can be composed of conveying device, opening device, tying rope, storage bag and other components. For single use, sterile product. For removal of tissue and/or biopsy samples from the surgical field during endoscopic procedures. | II | Merged with the original 336 to revise the product name and product description. | 337
338 | 06-16-07 | Disposable Endoscopic Cannula | Usually consists of an open end and a sleeve. It is generally made of polymer materials such as polyethylene. Supplied sterile. Used with the endoscope to isolate the endoscope from the patient during surgery or inspection to prevent cross-infection. | II |  | 338
339 | 06-16-07 | endoscopic bite | Auxiliary instruments for patients opening during surgery or examination, usually made of polyethylene and other polymer materials. Single use, provided sterile. It is used to maintain the patient's opening state during oral surgery or examination, and prevent unexpected damage to the bite protection device. | II |  | 339
340 | 06-18-01 | Image display processing workstation | It usually consists of a dedicated diagnostic display device, storage and processing system, software, etc. With medical imaging equipment, it is used to display, process, transmit and store digital diagnostic images. | II |  | 340
341 | 06-18-02 | Medical film digitizing scanner | It usually consists of a film upper plate, a film lower plate, a film transmission mechanism, a photocoupler (CCD), a digital signal processor, etc. It is used for scanning medical film and outputting digitized images in BMP, DICOM and other formats. | II |  | 341
342 | 07 | Intra-abdominal pressure monitoring kit | Including abdominal pressure monitoring and urine collection two-part device. The abdominal pressure monitoring device consists of a connector, a connector cap, a pipeline, a pipeline regulator, a guide clip, and a filtering device. The precise urine collection device consists of an introduction tube, a measurement chamber, a collection bag, an outlet valve and accessories (suspension belt/lanyard/hook). When in use, it is connected to the excretion port of the urinary catheter to measure the hydrostatic pressure in the bladder, thereby monitoring the intra-abdominal pressure, and accurately measuring and collecting urine. | II |  | 342
343 | 07-01-01 | spatula | Usually made of wood or other materials. Supplied sterile. Used to depress tongue during examination. | II |  | 343
344 | 07-01-02 | electronic stethoscope | It usually consists of a pickup, a signal processing module and earphones. Performs (frequency) non-linear amplification of collected sounds. Used to collect and amplify sounds from the heart, lungs, arteries, veins and other internal organs. | II |  | 344
345 | 07-02-01 | Gas measuring equipment | It is clinically used to continuously measure parameters such as the concentration, partial pressure, flow rate or volume of oxygen, carbon dioxide and other gases in breathing gas. | II | Revised product name and product description. | 345
346 | 07-02-03 | Lung Function Test Equipment | It usually consists of a host, a flow rate sensor and a nose clip. It is used for respiratory medicine, chest department, occupational disease prevention institutions, hospital physical examination, etc., to measure lung capacity, maximum ventilation and challenge test before and after medication. | II |  | 346
347 | 07-02-04 | expiratory flow meter | It usually consists of a housing, a slider, a reed and a mouthpiece. It is used to measure the maximum expiratory flow of patients with asthma and chronic obstructive pulmonary disease, measure the degree of restriction of exhalation of patients, and monitor the condition of asthma. | II |  | 347
348 | 07-02-05 | breathing pressure gauge | It usually consists of a host, mouthpiece, exhalation pressure test valve, inspiratory pressure test valve, bacterial filter, etc. It is used to measure the maximum oral inspiratory pressure, the maximum oral expiratory pressure and the nasal cavity inspiratory pressure. | II |  | 348
349 | 07-02-07 | breath analyzer | It usually consists of a host, gas sensors such as nitric oxide or carbon monoxide, respiratory filter devices or gas collection devices and other accessories. It is used to detect the concentration of nitric oxide or carbon monoxide in exhaled air. | II |  | 349 | NEW
350 | 07-03-01 | Sports load test test system | It usually consists of a host, accessories, and a motion unit. The host part usually includes a signal input part, an amplification circuit, a control circuit, a display part, a recording part, and a power supply part; the accessories usually include electrodes and cables; the motion unit is composed of a setting unit, an indicating unit, and a motion part that provide different strengths. The exercise unit can provide exercise with different loads. The host and accessories can monitor the ECG signal of the subject during exercise, process the signal and display it in real time. It is used for real-time detection of ECG changes in the patient's exercise state for clinical diagnosis. | II |  | 350
351 | 07-03-01 | ECG machine | Before the heart contracts mechanically, it first generates electrical excitation. The tiny current generated by the myocardial excitation can be transmitted to the body surface through the body tissue, causing different potentials to be generated in different parts of the body surface. The visible record of this cardiac action potential is the electrocardiogram. The electrocardiograph uses two electrodes placed on the body surface and connected to the device through the patient cable to record and process the potential difference between two points on the body surface in sequence according to the time sequence of cardiac excitation. Electrocardiographs are generally desktop or portable, consisting of a host, patient cables, and electrodes. Electrodes can also be divided into two types: reusable and disposable. The recording method can adopt thermal pen type or hot line array recording method, etc. Some products have signal input or signal output ports, which can be divided into several models according to different technical parameters and functions. The electrical signal is input into the computer for processing (such as high-frequency signal processing, etc.), but does not include auxiliary analysis and diagnosis functions. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1139-2013 ECG diagnostic equipment; single-channel and multi-channel electrocardiogram machines, YY 0782-2010 medical electrical equipment Part 2-51: recording and analysis type single-channel and special requirements for safety and basic performance of multi-channel ECG machines, etc. | II |  | 351
352 | 07-03-01 | Dynamic ECG recorder | A dynamic ECG recorder usually consists of a mobile recorder and lead wires, and may also include data management software and accessories (lead wires, chargers, USB cables, etc.), which can be classified according to the number of leads, technical parameters, intended use, Additional auxiliary functions are divided into several models; it is used for 24-hour dynamic ECG monitoring of the human body. Does not include automatic analysis diagnostics. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0885-2013 Medical electrical equipment Part 2: Special requirements for safety and basic performance of dynamic electrocardiogram systems. | II |  | 352
353 | 07-03-01 | Fetal ECG machine | The fetal electrocardiograph is mainly composed of a host, lead wires, and electrodes, and may include data management software, printing facilities, and accessories; it can be divided into several models according to the number of leads, technical parameters, intended use, and additional auxiliary functions; It is used for electrical inspection and monitoring, but does not include auxiliary analysis and diagnosis functions. | II |  | 353
354 | 07-03-03 | Mercury Sphygmomanometer | Mercury sphygmomanometers can have desktop, vertical and hanging types, basic parameters: measuring range: 0-40kPa (0-300mmHg), children's sphygmomanometer 0-20kPa (0-150mmHg); The division value is 0.5 kPa (2 mmHg); the allowable error of indication is ±0.5kPa (±3.75 mmHg); the cuff can be divided into multiple specifications according to the size of the airbag; the parameters for children can be adjusted; it is used for measuring human blood pressure. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 3053-1993 Sphygmomanometer and sphygmomanometer. | II |  | 354
355 | 07-03-03 | Mechanical Elastic Element Sphygmomanometer | The mechanical elastic element type blood pressure monitor adopts double scales, the measuring range for adults is 0-40kPa (0-300mmHg), the minimum graduation values are 0.5kPa and 2mmHg respectively, and the allowable error is ±0.5kPa (±3.75mmHg). Different airbag sizes are divided into multiple specifications; parameters for children can be adjusted, and can be divided into several models according to different types, technical parameters and applicable objects; it is used for measuring human blood pressure. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 3053-1993 Sphygmomanometer and sphygmomanometer. | II |  | 355
356 | 07-03-03 | Non-invasive automatic blood pressure monitor | The non-invasive automatic measuring blood pressure monitor is mainly composed of a host, an inflatable tube, a blood pressure cuff and a power supply. | II |  | 356
| #N/A |  | This product uses the oscillometric method or other similar methods to non-invasively measure human arterial blood pressure.
| #N/A |  | The product should comply with YY 0670-2008 non-invasive automatic blood pressure monitor, and the manufacturer should provide a clinical verification report on the overall effectiveness of blood pressure measurement.
357 | 07-03-04 | forehead thermometer | It usually consists of an infrared temperature sensor, a probe cover, a display unit, a power supply circuit, and a measurement circuit. Use infrared temperature sensing method to measure temperature display or data output. Clinical validation data for clinical accuracy should be provided. | II |  | 357
358 | 07-03-04 | Body temperature measuring equipment (passive) | It usually consists of a glass tube, a temperature-sensitive bulb, mercury or other temperature-sensitive liquid, and a scale. The temperature is measured by the principle of thermal expansion and contraction of mercury or other liquids. It is used for clinical measurement of patient's body temperature. It is usually placed in the mouth, armpit, and anus of the human body for measurement. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 358
359 | 07-03-04 | Medical Electronic Thermometer | Medical electronic thermometer is an electronic device for intermittent contact monitoring of patient's body temperature. It can be composed of plastic shell, circuit board, temperature measurement components, display screen, power supply, etc.; it can be divided into several types according to design, technical parameters, additional auxiliary functions, and intended use. Models for measuring body temperature or for women to monitor ovulation cycles, etc., do not include predictive mode, and do not include medical infrared thermometers. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 21416-2008 Medical Electronic Thermometer. | II |  | 359
360 | 07-03-04 | Ear cavity medical infrared thermometer | Ear cavity medical infrared thermometers can be composed of plastic shells, circuit boards, temperature measurement components, display screens, power supplies, isolation films, etc.; they can be divided into several models according to design, technical parameters, additional auxiliary functions, and intended uses; through heat radiation Display the temperature of the ear cavity of the measured human body. Clinical accuracy and clinical repeatability reports are required. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 21417.1-2008 Medical Infrared Thermometer Part 1: Ear cavity type. | II |  | 360
361 | 07-03-05 | Medical pulse oximeter | Medical pulse oximeters are usually composed of a host (including control, data processing/display modules) and blood oxygen probes connected to it, which can be divided into several models according to design, type, technical parameters, additional auxiliary functions, etc.; by establishing the relationship between optical radiation and blood oxygen saturation of human tissue, it is used for clinical non-invasive estimation and monitoring of human arterial blood oxygen saturation and pulse rate. A human assessment report on blood oxygen accuracy is required. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0784-2010 Medical Electrical Equipment - Special Requirements for Basic Safety and Main Performance of Medical Pulse Oximeter Equipment. | II |  | 361
362 | 07-03-06 | EEG machine | An EEG machine can be composed of an amplifier input box, a computer, a monitor, special software, and electrodes, and can also include a flashlight, a flash controller, and a printer. It can be divided into several models according to design, type, technical parameters, and additional auxiliary functions. It is used for human brain electrophysiological signal detection, processing, display, and storage. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB9706.26-2005 Medical Electrical Equipment Part 2-26: Special Requirements for the Safety of EEG Machines. | II |  | 362
363 | 07-03-06 | EEG topographer | The EEG topography instrument is mainly composed of a host, monitor, amplifier, printer, keyboard, mouse, etc. It can be divided into several models according to design, type, technical parameters, and additional auxiliary functions; it is used for quantitative processing and analysis of human EEG physiological signal data, excluding automatic diagnosis. | II |  | 363
364 | 07-03-06 | EMG machine | The electromyography machine is mainly composed of a host, stimulating electrodes, and may also include a computer, special data management software, and a printer. It can be divided into several models according to design, type, technical parameters, and additional auxiliary functions. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0896-2013 Medical Electrical Equipment Part 2: Special Requirements for the Safety of Myoelectric and Induced Response Equipment. | II |  | 364
365 | 07-03-09 | Transcutaneous Jaundice Tester | The transcutaneous jaundice tester is hand-held and consists of an optical test unit, display unit, power supply/battery, and charging accessories. It can be divided into several models according to the working principle, design, technical parameters, applicable parts, applicable objects, and auxiliary functions. It is mainly used for the measurement of neonatal percutaneous jaundice through specific wavelength light reflection. | II |  | 365
366 | 07-03-11 | Nasal Resistance Meter | It usually consists of a host, a pressure tube and a flow sensor. It is used to analyze nasal airway resistance by detecting nasal gas flow parameters. | II |  | 366
367 | 07-04-01 | multi-parameter monitor | A multi-parameter monitor is mainly composed of a host and accessories. The accessories can include ECG lead cables, blood pressure cuffs, blood oxygen probes, body temperature probes, respiratory gas measurement components, etc., which can be different according to design, type, technical parameters, additional auxiliary functions, etc. Divided into several models; for the patient's ECG, non-invasive blood pressure, pulse, blood oxygen saturation, body temperature, respiration, breathing gas monitoring. A human assessment report on the overall effectiveness of the non-invasive blood pressure system and blood oxygen accuracy is required. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1079-2008 ECG monitor, YY 0670-2008 non-invasive automatic blood pressure monitor, YY 0784-2010 medical electrical equipment - medical pulse oximeter equipment basic safety And the special requirements for main performance, YY 0785-2010 clinical thermometer - performance requirements for electronic thermometers for continuous measurement, YY 0667-2008 medical electrical equipment - special requirements for safety and basic performance of automatic circulation non-invasive blood pressure monitoring equipment, etc. | II | Revised product description. | 367
368 | 07-04-01 | Maternal/Fetal Multiparameter Monitor | The mother/fetal multi-parameter monitor is mainly composed of a host and accessories. The accessories can include ECG lead cables, blood pressure cuffs, blood oxygen probes, body temperature probes, ultrasound probes, uterine contraction pressure probes, marking handles, etc., which can be customized according to the design and type. It is divided into several models according to different technical parameters and additional auxiliary functions; it is used for monitoring the mother's ECG, non-invasive blood pressure, pulse, blood oxygen saturation, body temperature, respiration, uterine contraction pressure, fetal heart rate and fetal movement. A human assessment report on the overall effectiveness of the non-invasive blood pressure system and blood oxygen accuracy is required. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1079-2008 ECG monitor, YY 0670-2008 non-invasive automatic blood pressure monitor, YY 0784-2010 medical electrical equipment - basic safety of medical pulse oximeter equipment And main performance special requirements, YY 0785-2010 clinical thermometer - performance requirements for electronic thermometers for continuous measurement, etc. | II |  | 368
369 | 07-05-01 | Audiometer | It usually consists of power supply, sound card, power amplifier, controller, electro-acoustic transducer, operating software and patient responder. An instrument used to measure the individual's sensitivity to various frequencies. By comparing it with normal hearing, the hearing loss of the subject can be determined. | II |  | 369
370 | 07-05-02 | Electroglottograph | It usually consists of a host, glottis electrodes and audio output lines. When the vocal cords vibrate, changes in the contact impedance of the vocal cords cause changes in the modulation current, forming an electroglottogram. It is used to detect changes in glottal tissue impedance and vocal cord contact area, reflecting the characteristics of the glottis closure stage in each cycle of vocal cord vibration and the trajectory of each cycle of vocal cord vibration. | II |  | 370
371 | 07-05-04 | Middle Ear Analyzer | It is composed of a host, probe, earphone, etc. It is used to test the acoustic impedance and static pressure of the middle ear, and is suitable for the diagnosis of middle ear lesions. | II |  | 371
372 | 07-06-03 | gamma ray detector | Usually consists of solid state scintillation detectors and electronics. Used to find the distribution of radioactivity in the body and can be used in surgery. | II |  | 372 | NEW
373 | 07-07-01 | Ultrasound Doppler Blood Flow Analyzer | Ultrasonic Doppler blood flow analyzer uses the principle of ultrasonic Doppler frequency shift to detect and measure the movement information of non-fetal blood flow. It usually consists of a probe (generally a unit probe), ultrasonic transmitting/receiving circuit, signal processing and display, etc.; it can be configured according to model, imaging property (analog and digital), probe configuration, technical parameters, additional auxiliary functions, software processing, It is divided into several models according to the intended use; it is mainly used for blood flow measurement of transcranial, cervical and peripheral blood vessels, and cannot be used in surgery, and does not include analysis and diagnosis functions. Reference or applicable standards: YY 0593-2015 Transcranial Ultrasonic Doppler Blood Flow Analyzer, etc. | II |  | 373
374 | 07-07-01 | Ultrasound Doppler Cord Blood Flow Analyzer | Ultrasonic Doppler umbilical cord blood flow analyzer consists of a host and an ultrasonic Doppler probe, which can be divided into several models and functions according to design, technical parameters, auxiliary functions, etc. Blood flow detection/monitoring. | III |  | 374
375 | 07-07-02 | Ultrasound Bone Densitometer | Ultrasonic bone densitometer uses the difference of ultrasonic conduction velocity and the attenuation of amplitude to reflect the bone mineral content, bone structure and bone strength of the human body, by measuring the ultrasonic velocity (SOS) and/or broadband ultrasonic Attenuation (BUA), reflecting bone density status. The scope of application does not include the diagnostic conclusions of related diseases such as osteoporosis and bone loss. It mainly consists of ultrasonic transmitting circuit, receiving circuit, signal processing display part and recording part; it can be divided into several models according to design, technical parameters, additional auxiliary functions, etc.; it is suitable for people of corresponding age groups according to the distribution of clinical database population; The body surface is used for the measurement of human bone density. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0774-2010 Ultrasonic Bone Densitometer, etc. | II |  | 375
376 | 07-08-03 | Multi-parameter central monitoring system | There are two types of multi-parameter central monitoring systems: wired and wireless. The wired monitoring system is composed of a central host and a bedside multi-parameter monitor, and the wireless monitoring system is composed of a central host, a transmitter, and a bedside multi-parameter monitor; it can be divided into several models according to design, technical parameters, and auxiliary functions; For medical institutions to perform centralized monitoring and management of various physiological parameters of patients. A human assessment report on the overall effectiveness of the non-invasive blood pressure system and blood oxygen accuracy is required. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1079-2008 ECG monitor, YY 0670-2008 non-invasive automatic blood pressure monitor, YY 0784-2010 medical electrical equipment - basic safety of medical pulse oximeter equipment And main performance special requirements, YY 0785-2010 clinical thermometer - performance requirements for electronic thermometers for continuous measurement, etc. | II |  | 376
377 | 07-09-01 | Electronic urine meter | The electronic urine volume meter consists of a host (including a display terminal), supporting urine bags, brackets, etc. (excluding catheters); the instrument can be divided into several models according to technical parameters and additional functions; by measuring the urine in the urine bag The weight and volume of the fluid were used to calculate the parameters related to urine volume. | II |  | 377
378 | 07-09-03 | sleep apnea monitoring device | A sleep apnea monitor usually consists of a recorder, EEG electrodes, eye movement electrodes, myoelectric electrodes, chest/abdominal breathing probes, body position sensors, nasal airflow tubes, pulse oximeter probes, and ECG electrodes. Physiological parameters during sleep were recorded for analysis. It is used to record various physiological parameters during sleep, and to analyze and diagnose sleep disorders, sleep-disordered breathing, sleep apnea, and hypoventilation syndrome. | II |  | 378
379 | 07-09-04 | Balance Test System | It usually consists of measurement platform, auxiliary bracket, balance test software, etc. The patient's balance is assessed while standing or sitting on the test platform. Used to test human balance ability. | II |  | 379
380 | 07-09-08 | pH detector | It usually consists of a catheter with a pH sensor and a data recording device. It is used to monitor physiological parameters such as the pH value of the patient's stomach and esophagus, and to examine and diagnose gastroesophageal reflux disease. | II |  | 380
381 | 07-10-03 | Disposable ECG electrodes | Disposable electrocardiogram electrode is an accessory of electrocardiogram acquisition equipment, composed of sensing elements and electrolyte, with or without connecting wires. The electrode can be composed of base lining material, conductive paste, electrode buckle, etc. The base lining material can be made of air-permeable paper, spunlace cloth, non-woven fabric, foam paper, cotton cloth or PE coated with medical pressure-sensitive adhesive, and can be made according to its Different shapes are divided into round, elliptical, square, etc., and are divided into several specifications according to their size; with the instrument, it is used for ECG detection and monitoring. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0196-2005 disposable ECG electrodes. | II |  | 381
382 | 07-10-03 | Disposable non-invasive EEG electrodes | Disposable non-invasive EEG electrodes are composed of electrode sheets (conductive plastic and metal plating), wires and equipment connectors. One end can be connected to EEG diagnosis/monitoring equipment, and the other end can be connected to human head skin for collecting and transmitting EEG. For biosignals. A single patient can be used repeatedly in a short period of time. | II |  | 382
383 | 07-10-03 | body surface reference electrode | It is generally composed of electrodes (adhesive patches, sensors) and connecting wires, and can be divided into round, oval, square, etc. according to its shape. Non-sterile product, single use. Used in cardiac electrophysiological examination, it should be used in conjunction with corresponding electrophysiological mapping equipment or catheters to transmit the position of the catheter. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T0196-2005 disposable ECG electrodes. | II |  | 383
384 | 07-10-04 | Non-invasive medical sensors | Including temperature, pressure, gravity, angle and other medical sensors. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0781-2010 blood pressure sensor, etc. Those with special requirements will be dealt with according to special requirements. | II |  | 384
385 | 07-10-05 | Medical Conductive Paste | Medical conductive paste is usually a water-based polymer gel, which contains a biocompatible and stable electrolyte, and is used for ECG, EEG, EMG and other examinations and electrotherapy electrode coupling. | II | Revised product description. | 385
386 | 08-02-02 | anesthesia puncture needle | It usually consists of a bushing seat, a needle seat, a needle tube, a protective sleeve, and a bushing. The needle tube is generally made of stainless steel; the lining core is generally made of stainless steel or plastic material. It is used to puncture the human body and inject drugs. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 386
387 | 08-02-02 | Disposable anesthesia puncture needle | Performance, structure, composition, use, etc. belong to the products that can be fully covered by YY 0321.2 single-use anesthesia needles. The product is generally composed of a liner seat, a needle seat, a needle tube and a liner, etc., and is used for puncture and drug injection during anesthesia. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 387
388 | 08-02-02 | Disposable filter for anesthesia | Performance, structure, composition, use, etc. belong to the products that can be fully covered by YY 0321.3 single-use anesthesia filter. Disposable filters for anesthesia are generally composed of upper cover, filter membrane, base, inner conical joint protective sleeve (optional), outer conical joint protective sleeve (optional), and are divided into liquid medicine filter and air filter. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 388
389 | 08-03-05 | Artificial resuscitator (simple respirator) | It usually consists of an intake valve, a compression unit (such as an airbag) and a patient valve, and is generally equipped with accessories such as an air storage bag and a breathing mask. It is a resuscitation device that ventilates the patient's lungs by the operator depressing a compression unit (such as an air bag) on the device. It is used to provide lung ventilation when the power supply and air supply are incomplete and emergency situations are used to provide artificial respiration for patients with sudden dyspnea or respiratory failure. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 389
390 | 08-04-01 | Medical molecular sieve oxygen system | It usually consists of air compression system, air source purification system, air tank, medical molecular sieve adsorption separation system, product gas tank, control system, monitoring and alarm system, etc. A gas source system for producing oxygen-enriched air (93% oxygen) by using the principle of molecular sieve pressure swing adsorption to enrich oxygen from the air by separating nitrogen in the atmosphere. It is used to produce oxygen-enriched air (93% oxygen), provide gas source to other oxygen-using medical devices through the medical gas pipeline system, and supply oxygen to patients according to its clinical application range. | II | Revised product description. | 390
391 | 08-04-02 | Small molecular sieve oxygen generator | It usually consists of an air compression pump, a medical molecular sieve adsorption separation system, an oxygen tank, an output flow control display device, an oxygen concentration monitoring device, a timing device and an alarm system. A device for producing oxygen-enriched air (93% oxygen) by separating nitrogen in the atmosphere and enriching oxygen from the air by using the principle of molecular sieve pressure swing adsorption. It is used to produce oxygen-enriched air (93% oxygen) and supply oxygen to patients according to its clinical application range. | II | Revised product description. | 391
392 | 08-04-03 | Medical Membrane Separation Oxygen System | Using the principle of membrane separation technology, oxygen is enriched from the air to obtain oxygen-enriched air with an oxygen content of more than 90%. It usually consists of air compression and pretreatment equipment, medical membrane separation oxygen generator, control and monitoring instruments, pipeline accessories, etc. It is established and used in medical institutions to provide gas sources to other oxygen-using medical devices through the medical gas pipeline system, and to supply oxygen to patients according to its clinical application range. | II |  | 392
393 | 08-04-04 | Small medical membrane separation oxygen generator | Using the principle of membrane separation technology, oxygen is enriched from the air to obtain oxygen-enriched air with an oxygen content of more than 90%. It usually consists of air compression and pretreatment equipment, medical membrane separation oxygen generator, control and monitoring instruments, pipeline accessories, etc. It can be composed of the main unit of the oxygen generator and accessories (humidifier, oxygen inhalation tube/mask, etc.), and has the function of prompting the oxygen concentration to be too low. It can be divided into different models according to the oxygen production per unit time, additional components, and functions. For patients requiring high-concentration supplemental oxygen as prescribed. | II |  | 393
394 | 08-04-05 | oxygen generator | It usually consists of a cylinder and an upper cover, and can be equipped with accessories such as an oxygen suction tube. Using water as the raw material, the oxygen generator is chemically reacted by using water electrolysis to generate medical oxygen and provide patients with oxygen for clinical breathing. | II |  | 394
395 | 08-05-01 | anesthesia vaporizer | Anesthesia vaporizers have different implementation methods such as mechanical adjustment mechanical control type, mechanical adjustment electronic control type and electronic adjustment electronic control type. The anesthetic agents that can be used include sevoflurane, isoflurane, desflurane, etc. The product is generally composed of a concentration control device, a dosing device (including tank filling adapter, etc.) and an anesthetic liquid level indicator. An anesthetic vapor flow containing a certain percentage concentration is formed at the output port through the concentration control device. The anesthesia vaporizer is a component of the anesthesia machine, which is used to provide a controlled concentration of anesthetic vapor. | III |  | 395
396 | 08-05-02 | Medical respiratory humidifier | Medical respiratory humidifiers are usually used in conjunction with ventilators or oxygen inhalers, and generally consist of a temperature control system, a humidification chamber and a heating device, or only an air inlet, a water storage bottle, a humidification chamber and an air outlet. A device or device used to increase the humidity level of a medical gas delivered to a patient. Used to humidify the breathing gas delivered to the patient. | II |  | 396
397 | 08-05-03 | respiratory filter | The respiratory system filter can be composed of four parts: the upper cover, the filter medium, the lower cover, and the protective cap. The filter housing can be made of ABS and other materials, and the filter medium can be made of polypropylene composite materials. Divided into several models and specifications according to different sizes and sizes, it is provided in a sterile form and matched with anesthesia breathing equipment and pulmonary function instruments to reduce the number of particulate matter (including microorganisms) inhaled or exhaled by patients. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0753 Respiratory system filter for anesthesia and respiration. | II |  | 397
398 | 08-05-04 | Disposable HME/Filter | The product can be composed of shell, filter element, heat and moisture exchange medium and interface. Single use, clean or sterile packaging. This product is suitable for patients who have established an artificial airway to perform heat and moisture exchange and/or filter particulates on the gas they breathe. Product recommendation standard YY/T 0735 Anesthesia and respiratory equipment Heat and moisture exchangers for humidifying human respiratory gases. | II |  | 398
399 | 08-05-05 | Breathing circuit aids | It is usually used to realize the connection between airway products, or to assist in the insertion of accessories such as airways. Accessories for airway connection, sampling, introduction, etc. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 399
400 | 08-05-05 | Cannula introducer | The intubation introducer can be made up of components such as catheter, hardening sleeve, joint. The product is sterilized and disposable. The intubation introducer is used to assist in inserting the tracheal tube into the patient's trachea when the patient's glottis is not fully exposed. | II |  | 400
401 | 08-05-05 | Catheters and connectors for lung ventilation function tester | Usually the product contains one catheter, one connecting device and one operation manual. The sterile device is connected with the lung lobe ventilation function tester and used through the working channel of the bronchoscope. Product is intended for use during diagnostic bronchoscopy in adult patients with chronic obstructive pulmonary disease (COPD) and emphysema by a bronchoscopist in a bronchoscopy laboratory. | II | Add | 401
402 | 08-05-06 | Anesthesia Laryngoscope | It usually consists of a handle, a peeping piece, an internal power supply and a light source for lighting, and may have a video display function. It is an auxiliary device used during endotracheal intubation. | II |  | 402
403 | 08-05-06 | Anesthesia video laryngoscope | The anesthesia video laryngoscope consists of a handle, a camera and a peeping piece, and the other end of the handle can be connected to a monitor. The speculum can be made of metal for repeated use, or it can be made of polymer material for one-time use; the handle and the speculum are usually detachable, and can be classified according to the applicable crowd, material, design, technical parameters, and speculum size. There are several models and specifications; it is used to clinically provoke the epiglottis of patients to expose the glottis, guide medical staff to accurately perform airway intubation for anesthesia or first aid, and can also be used for oral examination and treatment. | II |  | 403
404 | 08-05-06 | Disposable anesthesia laryngoscope | The disposable anesthesia laryngoscope can be composed of a peeping piece and a lens barrel. The peeping piece can be divided into two types: direct lighting and optical fiber lighting; the peeping piece is made of polyvinyl chloride, and the lens barrel should be made of high-density polyethylene resin; It is divided into several models and specifications according to different objects, types, and sizes; the peeping piece is provided in a sterile form, and it is used with an anesthesia throat handle for clinical throat examination, first aid, and anesthesia operations to assist intubation. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0499-2004 General technical conditions for anesthesia laryngoscopes. | II |  | 404
405 | 08-05-06 | Disposable anesthesia peep | It is generally made of medical polymer materials (such as: PC); the product consists of three parts: a lock, a sleeve and a tongue depressor. It is used to clinically provoke the epiglottis of patients to expose the glottis, guide medical staff to accurately perform airway intubation for anesthesia or first aid, and can also be used for oral examination and treatment. | II |  | 405
406 | 08-05-07 | Mesh atomizer | The net atomizer can be composed of the main engine (control part), atomization device, liquid container, nozzle/mask, etc. It can be divided into several models according to the structure, composition, technical parameters, additional auxiliary functions, etc., through piezoelectric ceramic vibration Drive the movement of the atomizing device, squeeze out the liquid medicine from the hole, and atomize the medicine for inhalation treatment. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0109-2003 Ultrasonic Nebulizer. | II |  | 406 | NEW
407 | 08-05-07 | Atomization equipment / atomization device | Atomization is carried out by means of ultrasonic waves, built-in electric pumps, and external air sources. A device or device used to convert a liquid into an aerosol. | II |  | 407
408 | 08-05-07 | Compression atomizer | The compression atomizer is composed of a reciprocating piston micro-compressor, a connecting hose, a mist manual switch, a nozzle, a hose, and a mask; it can be divided into several models according to different technical parameters and additional functions; the air mist produced by compressed gas Aerosolized drugs and delivered to the respiratory tract for inhalation therapy of respiratory nebulized drugs. | II |  | 408
409 | 08-05-07 | Medical Ultrasonic Nebulizer | The medical ultrasonic nebulizer can be composed of the host (including the atomization and control part) and the atomization delivery accessories. It can be divided into several models according to the structure, composition, technical parameters, and additional auxiliary functions. The atomized medicine is used for treatment. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0109-2003 Ultrasonic Nebulizer. | II |  | 409
410 | 08-05-07 | Disposable Airflow Nebulizer | In the respiratory device system, one end is connected to oxygen, and the pressure of oxygen is used to atomize the medicine placed in the nebulizer, and the generated nebulized gas is inhaled into the mouth of the patient through the other end of the nebulizer. | II |  | 410
411 | 08-05-07 | dosing spray bottle | Products are generally made of medical polyvinyl chloride, silicone rubber, polypropylene, polyacetate and other materials. It is divided into mask type and mouth type, and can be composed of dosing spray bottle, medicine cup, micro-inhaler and other components. Used to aerosolize medication and administer it into the patient's mouth or nose. | II |  | 411
412 | 08-05-08 | anesthesia reservoir | Usually composed of breathing bags and connectors, it is an elastic container for storing gas in the anesthesia breathing system of the anesthesia machine. | II |  | 412
413 | 08-05-09 | Disposable anesthesia waste gas adsorber | The main structure of the disposable anesthetic gas adsorber is mainly composed of a packaging cover, a gas distribution head, a filter plate, an upper cover, and a lower seat, and is made of plastics that comply with GB/T 21460.2-2008; - For activated carbon specified in Table 3 of 2008 (excluding particle size indicators), the particle size indicators should meet the requirements in Table 4. The filling amount of activated carbon is not less than 80g, and the average adsorption rate is ≥90%. Product is supplied non-sterile. It is used for clinical adsorption of residual anesthetic gas discharged from anesthesia pipeline. | II |  | 413
414 | 08-05-10 | Anesthetic gas purification delivery and collection system | It is usually composed of a connection port, a gas storage tank, etc., and is used to transmit and collect anesthesia waste gas and perform centralized purification treatment. For the delivery of exhaled and/or expelled excess anesthetic gases to a suitable vent. | II |  | 414
415 | 08-06-01 | Disposable Anesthesia Catheters and Connectors | Performance, structure, composition, use, etc. adopt the applicable parts of the following reference standards, such as YY/T0321.1-2022 . The product is generally composed of catheters and catheter connectors, which are matched with disposable anesthesia needles for clinical delivery of anesthetics. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III | Revised product description. | 415
416 | 08-06-02 | breathing circuit | Usually it is a corrugated tube with a "herringbone" or "one" structure, and some pipelines can be stretched axially. The pipeline with a herringbone structure is composed of an inhalation branch and an exhalation branch, and is generally made of plastic or silicone rubber. into. The breathing circuit with heating function also includes heating wire and power adapter. Disposable or reusable. Or a bellows or hose that is usually in the shape of a "one" and may also include a mouthpiece, which is generally made of plastic or silicone rubber. Supplied sterile. | II |  | 416
417 | 08-06-02 | Breathing circuits for anesthesia machines and ventilators | The breathing pipelines for anesthesia machines and ventilators are made of suitable materials such as polyvinyl chloride, and are divided into two types: anesthesia machine pipelines and ventilator pipelines. Type joints, bellows, and swivel elbows. The two types of pipelines can be combined. Air bags and masks are optional. They can be classified according to material, design, combination, object of use (adult/children), technical parameters, and optional accessories. Available in several models and sizes; provided in sterile form; used for tubing connections to anesthesia machines, ventilator interfaces, and oxygen masks. The product performance index adopts the applicable parts of the following reference standards, such as: YY 0461-2003 breathing circuit for anesthesia machine and ventilator. | II |  | 417
418 | 08-06-03 | Endotracheal Tube / Endotracheal Tube | The common intubation head has one or two cuffs. After the cuffs are inflated, they can fix the intubation tube and seal the airway, or they can be without cuffs. The intubation tube body is usually made of polymer material, and a steel wire coil is embedded in the tube body to improve radial strength and axial flexibility. Part of the intubation tube body is made of anti-laser material or layered to resist laser irradiation. Intubation for nasal/oral or percutaneous insertion into the patient's trachea. One end is connected to the anesthesia ventilator through the breathing tube to maintain the patient's breathing. Supplied sterile, single use. Or usually consists of a base plate, an inner casing, an outer casing and a tube core. | II |  | 418
419 | 08-06-03 | Disposable endotracheal tube | Disposable tracheal intubation usually consists of a tracheal intubation tube body, an inflatable tube, a cuff, a joint, and an inflatable valve; the tube body, inflatable tube, and cuff of the tracheal intubation can be made of polyethylene and other plastics, silicone rubber, etc. The joints and inflation valve can be made of polyethylene and other materials; endotracheal intubation can be divided into different types such as ordinary with/without cuff and reinforced with/without cuff, and each type can be used according to the nominal The inner diameter and cuff are divided into several specifications; the product should be provided in a sterile form; it is used for establishing an artificial airway during clinical anesthesia or first aid. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0337-2002 Endotracheal Tube. Excludes bronchial intubation. | II |  | 419
420 | 08-06-03 | Disposable tracheostomy tube | Disposable tracheostomy tube consists of tracheostomy tube body (patient end with rounded tip, bevel and fixed balloon), standard male conical connector (machine end), integrated indicator balloon/valve combination and fixation Among them, the tube body, air tube and cuff of the tracheotomy tube can be made of polyethylene plastic, rubber and other suitable materials, and the joint and inflation valve can be made of polyethylene and other materials; the tracheotomy tube It can be divided into different types such as ordinary with/without cuff intubation and reinforced with/without cuff intubation. Each type can be divided into several specifications according to the applicable population, nominal inner diameter, and cuff capacity; Provided in a sterile form; used for clinical emergency tracheal section incision and establishment of artificial airway. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0338-2002 Tracheostomy Tube. | II |  | 420
421 | 08-06-03 | bronchial intubation | The product can be composed of joints, tube body, tracheal cavity/cuff/indicating balloon, bronchial cavity/cuff/indicating balloon, inflatable tube/valve, guide wire and other components. The product is sterilized and disposable. For mainstem bronchial intubation, allowing selective inflation/aspiration, suction, and bronchoscopy of either lung. | II |  | 421
422 | 08-06-03 | Double-lumen endobronchial intubation | A bronchial double-lumen tube can be composed of a bronchial double-lumen tube with a probe and a connector. The product is sterilized and disposable. The bronchial double-lumen tube is placed in the main bronchus and can be used for selective inflation or deflation, suction, unilateral lung ventilation, or for use during bronchoscopy. | II |  | 422
423 | 08-06-03 | airway exchange catheter | Airway exchange catheters can be composed of catheters, connectors and other components. This product is sterilized and disposable. For exchanging endotracheal tubes when needed. | II |  | 423
424 | 08-06-05 | laryngeal mask | It usually consists of a cuff, an inflatable tube, a laryngeal mask tube, a machine end, a joint, and an indicator balloon. After being inserted into the throat, the cuff is inflated to block the oral cavity and esophagus, while keeping the patient's trachea open. Generally provided sterile, single use. | II |  | 424
425 | 08-06-05 | Disposable medical laryngeal mask | Disposable medical laryngeal mask can be made of silicone rubber, etc. It is composed of cuff, inflatable tube, laryngeal mask intubation, machine end, joint, indicator balloon, etc. The laryngeal mask intubation can be metal reinforced, and can be selected according to material, Different designs, technical parameters, and auxiliary functions are divided into several models and specifications; they are used for general anesthesia or cardiopulmonary resuscitation to establish short-term artificial airways. | II |  | 425
426 | 08-06-06 | Oropharyngeal/nasopharyngeal airway | Disposable oral/nasopharyngeal airway is made of polymer materials, and it is a breathing tube with a flanged end. It can be divided into several models and specifications according to material, design, and insertion site diameter/length; it can be sterile Provided in the form of oral / nasopharyngeal airway for the prevention of airway obstruction caused by falling back of the tongue. | II |  | 426
427 | 08-06-07 | bronchial occluder | It usually consists of catheters, catheter adapters, balloons, balloon filling connectors, multi-way airway conversion connectors and other devices. The balloon is inflated for insertion of a bronchial tube to temporarily occlude the bronchi. It is inserted into the patient's bronchus in operations requiring one-lung ventilation to block the airway of the left lung or the airway of the right lung. | II |  | 427
428 | 08-06-08 | nasal oxygen tube | It usually consists of oxygen inlet interface, oxygen hose, adjustment ring, nasal plug (or mask), etc. The nasal oxygen tube is connected to the oxygen delivery system for the patient to inhale oxygen. Single use. It is used for direct delivery or humidification of oxygen between the oxygen source and the oxygen receiver during oxygen inhalation. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 428
429 | 08-06-09 | Suction catheter for airway | The suction catheter for respiratory tract usually consists of a catheter with end holes and a plugged connector; the catheter can be made of polyvinyl chloride, natural latex, silicone rubber, rubber material, and the plugged connector can be made of PVC, silicone rubber or MD type accordingly Made of polyvinyl chloride; catheters are divided into several models and specifications according to different materials and diameters, and the products should be provided in sterile form. It can be connected with a vacuum negative pressure system or equipment for suctioning sputum from the respiratory tract. Does not include multilumen airway suction catheters without port holes. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0339-2009 suction catheter for respiratory tract. Also known as disposable suction tube. | II |  | 429
430 | 08-06-10 | breathing mask | It usually consists of a headband, a forehead cushion, a mask frame, an interface cover, an elastic band and a silicone sealing cover, which directly contacts the patient and ventilates through the patient's mouth and nasal cavity. | II |  | 430
431 | 08-06-11 | Face masks, masks, nasal masks for continuous positive pressure ventilation | It usually consists of a headband, a forehead cushion, a mask frame, an interface cover, an elastic band and a silicone sealing cover, which directly contacts the patient and ventilates through the patient's mouth and nasal cavity. | II |  | 431
432 | 08-06-11 | Face masks, masks, nasal masks for continuous positive pressure ventilation | It usually consists of a nasal mask, a mask or an oronasal mask body, a fixed headband, an air circuit interface, an exhaust port, an anti-suffocation valve, etc.; an interface connection device that directly contacts the patient's face and ventilates through the nasal cavity and/or oral cavity. It is used for non-invasive ventilation support such as chronic respiratory insufficiency, improved ventilation and sleep therapy. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 432
433 | 08-06-13 | anesthesia mask | It usually consists of an interface, an airbag, and a cover. The material that forms mask has plastic etc. Disposable or reusable. | II |  | 433
434 | 08-06-13 | Disposable anesthesia mask | Disposable anesthesia masks can be made of polymer materials such as silicone rubber, and can be composed of face cups, air bags, valve core components (valve core valve body, valve core, sealing ring, spring and bottom cover) and fixing rings, which can be selected according to the material It is divided into several models and specifications according to different types, designs, technical parameters, intended uses, and auxiliary functions; it is used for clinical anesthesia and breathing. | II |  | 434
435 | 08-06-14 | oxygen mask | It usually consists of a mask and a connecting tube. Made of medical polymer materials. It is used to deliver oxygen to hypoxic patients as a channel for oxygen to enter the patient's body. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II | Add | 435
436 | 08-07-01 | Medical Air Compressor | It is used in conjunction with ventilators to produce medical compressed air and provide compressed air sources for ventilators or similar breathing ventilation systems. | II | Add | 436
437 | 08-07-02 | medical gas mixer | It is usually composed of a host (gas pressure balance processing module, gas proportion distribution control module, gas differential pressure or lack of pressure alarm module), gas outlet, oxygen inlet, air inlet, oxygen concentration adjustment device, etc. | II | Revised product name and product description. | 437
| #N/A |  | It is used to mix the input medical oxygen and air according to the set concentration, and adjust and control the oxygen concentration and flow rate of the air-oxygen mixed gas.
438 | 08-07-03 | medical oxygen supply | Medical oxygen supply consists of flowmeter, pressure gauge, pressure reducer, trolley, etc., for medical emergency or medical scene oxygen supply. | II |  | 438
439 | 08-07-03 | Portable Medical Oxygen Provider | It usually consists of an aerosol valve, an oxygen mask or a nasal oxygen tube (without oxygen). For the prevention of mild hypoxia. | II | Revised product description. | 439
440 | 08-07-04 | Medical Air Compressor Units | The medical air compressor unit is used in the medical gas conduit system, which can be composed of an air tank, a connecting device and an electric control box; it produces compressed gas for diagnosis and treatment. | II | Revised product classification code, product name and product description. | 440
441 | 08-07-04 | Medical gas pipeline system - oxygen supply system | Medical gas pipeline system - the oxygen supply system usually consists of a central oxygen supply station (which can be an oxygen concentrator, a medical oxygen station or a combination of the two), a set of monitoring and alarm systems, and a pipeline distribution system with several terminals. The central oxygen supply station is divided into several models depending on the type. The system has functions such as undervoltage and other sound and light alarms; it is used for centralized oxygen supply in medical institutions. The main product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0801-2010 Medical gas pipeline system terminal, YY/T 0799-2010 Medical gas low-pressure hose assembly, YY 1468-2016 Oxygen for medical gas pipeline system Concentrator gas supply system, YY/T 0187 General technical specifications for medical center oxygen supply system. | II |  | 441
442 | 08-07-05 | Medical Gas Manifold | It usually consists of power supply device, gas valve, flow control valve, gas deviation controller, etc., excluding gas cylinders. | II |  | 442
443 | 08-07-06 | Medical Gas Alarm System | The medical gas alarm system can be composed of a variety of multi-point gas sensor devices and display transmission systems. The monitoring points are transmitted to the central gas management system through a wired connection. The central management system monitors the medical gas monitoring points in a centralized way. Alarm; it can be used to monitor various gases such as oxygen and anesthetic gases in the hospital gas supply system. | II |  | 443
444 | 09 | Inverted nipple retractor | The inverted nipple suction device is composed of a cylindrical tube, an inner core, etc., and uses the principle of negative pressure to suck and pull the nipple in place; it is used for treating inverted nipples. | II |  | 444
445 | 09 | Microwave Hyperthermia Silicone Therapy Catheter | Microwave hyperthermia silicone treatment catheters are made of silicone, usually composed of valves, cavity heads, multi-lumen tubes and balloons, without electrodes; they can be divided into several models and Specifications; provided in a sterile form; used for drug infusion and microwave therapy with drugs and microwave therapy devices respectively. | II |  | 445
446 | 09-01-03 | Medium frequency electrotherapy products | The product consists of a host (signal generation and control device), electrodes, wires and other accessories. It is a method of treatment and rehabilitation by applying alternating current (including sine wave, pulse wave and modulated wave, etc.) with a frequency of 1kHz-100kHz. The intended use should reflect the clinical indications and therapeutic effects. Such as: This product has anti-inflammatory and analgesic effects on frozen shoulder. Product performance indicators, clinical therapeutic effects and indications should be within the scope of the "Technical Guidelines for Registration of Medium Frequency Electrotherapy Products". | II |  | 446
447 | 09-01-05 | Electrodes for nerve and muscle stimulators | Electrodes for nerve and muscle stimulators are composed of electrode sheets (conductive plastic and metal plating), wires and equipment connectors, which can be divided into several models and specifications according to different materials, technical parameters, and sizes; one end can be connected to nerve and muscle stimulators One end is connected to the intact skin of the human body; the product can be provided in a sterile form, and can be reused by a single patient within a certain period of time; it is used for non-invasive transmission of nerve and muscle stimulator stimulation electrical signals to the human body. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0868-2011 Electrodes for nerve and muscle stimulators. | II |  | 447
448 | 09-02-01 | Computer constant temperature electric wax therapy instrument | It usually consists of a host, a wax melting device, a temperature control device, a temperature detection device, wax, etc. A device that uses heated and melted paraffin and beeswax as a heat conductor to transmit heat energy to the body to achieve therapeutic effects. | II |  | 448
449 | 09-02-01 | warming blanket | A warming blanket generally consists of a host, a blanket and a body temperature sensor. In the clinical use environment of medical institutions, by controlling the temperature of the blanket, it has the function of physically heating the human body outside the body to achieve the purpose of assisting in regulating the temperature of the human body. It can be divided into several models according to design, technical parameters, auxiliary functions, etc. | II |  | 449
450 | 09-02-01 | Hot compress (bag) | Usually contains heat-generating materials and is encapsulated in medical non-woven fabrics or other medical materials. There is no temperature protection device, and it is directly applied to the affected area when used, and the heat is transferred to the affected area by conduction, which is used to promote local blood circulation, assist in anti-inflammatory, detumescence and pain relief. | II |  | 450
451 | 09-02-03 | Medical temperature control blanket | It usually consists of a temperature adjustment device, a temperature control circuit, a control mechanism and an application part. In the clinical use environment of medical institutions, by controlling the temperature of the application part, the physical heating and/or cooling function of the human body can be realized outside the body, so as to achieve the purpose of assisting in regulating the temperature of the human body. | II |  | 451
452 | 09-02-03 | cooling blanket | The cooling blanket generally consists of a host, a blanket and a body temperature sensor. In the clinical environment of medical institutions, by controlling the temperature of the blanket, it has the function of physically cooling the human body outside the body to achieve the purpose of assisting in regulating the temperature of the human body. It can be divided into several models according to design, technical parameters, auxiliary functions, etc. | II |  | 452
453 | 09-03-06 | Newborn Blue Light Therapy Apparatus | The neonatal blue light therapy instrument consists of a blue light lamp, a control box, a support rod, a base and its accessories. No other confounding light sources. For the treatment of neonatal jaundice. YY 0669-2008 Medical electrical equipment Part 2: Special requirements for the safety of infant phototherapy equipment are applicable to this product. | II |  | 453
454 | 09-03-08 | light guide nose | Used in conjunction with phototherapy equipment, its composition and principle depend on the type and function of the phototherapy equipment. It is mainly used to connect or wrap the end of the light guide element and insert it into the nasal cavity for easy fixation and avoiding cross-infection. | II |  | 454
455 | 09-04-01 | electric nasal aspirator | The electric nasal aspirator consists of a suction head, a host, and a battery. When in use, the suction head touches the outer edge of the child's nostrils, and the host generates suction. Use the principle of negative pressure to suck away children's nasal secretions. | II |  | 455
456 | 09-04-01 | Vibrating tapping expectorant machine | The vibrating knocking phlegm machine is mainly composed of a host, a flexible transmission shaft and a variety of knocking heads. The speed, frequency, intensity, tapping mode and working time of the tapping head are adjustable. Different designs, technical parameters, applicable objects, etc. are divided into several models. Vibration and percussion can improve the blood circulation of the patient's lungs and assist in the discharge of respiratory secretions. | II |  | 456
457 | 09-04-02 | medical elastic stockings | The product meets the requirements of YY/T 0851 anti-thrombotic socks or YY/T 08531 medical compression socks for varicose veins. It is used to promote venous blood flow, prevent deep vein thrombosis, and relieve symptoms of varicose veins. | II |  | 457
458 | 09-04-03 | Electric cervical and lumbar traction device | The electric cervical and lumbar traction device is composed of traction power system, control system and auxiliary components. It can be divided into several models according to different design types, technical parameters, applicable parts, and intended uses. It can be used for traction treatment of cervical spine indications. Quick pull device not included. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0697-2016 Electric cervical and lumbar traction therapy equipment . | II |  | 458
459 | 09-04-03 | manual traction bed | It is usually composed of a traction host that generates and adjusts mechanical force and a rope that transmits force, and also includes accessories such as a bed (chair) that carries the patient and matching patient fixing belts. The traction host is a manual structure, and the patient fixing belt is tied to the patient's pillow, jaw, chest, hip or limbs, etc., and the traction force is transmitted to the patient through skin friction, which can provide horizontal cervical and lumbar traction, or vertical cervical spine traction. | II |  | 459
460 | 09-04-04 | traction device | It is fixed on the neck and waist through inflatable airbags or devices with curved surfaces, so that the cervical spine and lumbar spine can maintain a certain angle and produce a stretching effect on the surrounding muscles, which is used to relax the muscles around the spine and relieve intervertebral pressure. | II |  | 460
461 | 09-08-01 | medical air pressurized chamber | It is generally composed of a cabin (shell, hatch, observation window, safety valve, etc.), supply and exhaust system, oxygen supply and exhaust system, air conditioning system, electrical system, water spray system, and console. The pressurized medium is air, and the maximum working pressure is not more than 0.3MPa oxygen chamber. According to the number of people treated in the cabin, it is divided into single-person oxygen chamber and multi-person oxygen chamber. Expose the patient to a hyperbaric oxygen environment and make him breathe hyperbaric oxygen to achieve the purpose of treating the disease. It is used for adult patients who need supplementary oxygen supply therapy according to the doctor's prescription, and is used for adjuvant treatment of ischemic and hypoxic diseases. Does not include software cabins, special-shaped cabins, and infant oxygen cabins. | III |  | 461
462 | 09-08-01 | medical pressurized oxygen chamber | It generally consists of a cabin, an oxygen supply and exhaust system, an air conditioning system and a control system. The pressurized medium is medical oxygen, the maximum working pressure is not more than 0.2MPa, and the number of people entering the warehouse is 1. It is intended for adult patients who need high-concentration supplemental oxygen supply according to the prescription, and is used for adjuvant treatment of ischemic and hypoxic diseases. Expose the patient to a hyperbaric oxygen environment and make him breathe hyperbaric oxygen to achieve the purpose of treating the disease. It is used for adult patients who need supplementary oxygen supply according to the doctor's prescription, and it is used for adjuvant treatment of ischemic and hypoxic diseases. Does not include software cabins, special-shaped cabins, and infant oxygen cabins. | III |  | 462
463 | 09-08-04 | Medical bath therapy machine | Medical immersion therapy machine is composed of host, patient bath bed, water supply/drainage system, temperature control unit, etc. It can also include patient transfer hanger and other auxiliary immersion functions, which can be divided according to design, technical parameters, auxiliary functions, etc. for several models. It is used for bath treatment of scalded patients. | II |  | 463
464 | 09-08-05 | intestinal hydrotherapy machine | It usually consists of a host, a temperature control device, a pressure (flow) control system, a liquid tank, a peristaltic pump, a liquid injection pipe, a liquid discharge pipe, etc. During the treatment, the fluid is perfused into the intestinal tract, and the fluid is drained out of the body through a drainage tube. Used in healthcare settings for colon cleansing. | II |  | 464 | NEW
465 | 10 | blood line connector | The blood circuit connector (joint for short) belongs to the cardiopulmonary bypass system and is used for blood delivery, observation and connection in the cardiopulmonary bypass system of extracorporeal circulation surgery. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 465
466 | 10-01 | Medical Vacuum Freeze Dryer | The medical vacuum freeze dryer freezes the plasma and other freeze-dried substances into a solid state, and then sublimates the water in it from the solid state to a gaseous state to remove the water and preserve the plasma. It is mainly composed of refrigeration system, vacuum system, heating system, control system and so on. | II |  | 466
467 | 10-01-06 | blood thawing equipment | It usually consists of a heating water tank, a thawing tank, a circulation pipeline, a water inlet pipe, a drain pipe, and a control box. It is clinically used for heating and thawing plasma or blood. | II |  | 467
468 | 10-01-06 | Frozen plasma thaw box | The frozen plasma thawing box is composed of a heating water tank, a thawing tank, a circulation pipeline, a water inlet pipe, a drain pipe, and an electrical control box. There are several models with different functions, which are used for heating and thawing clinical plasma. | II |  | 468
469 | 10-02-03 | Disposable arteriovenous puncture needle/device (internal fistula needle) | It usually consists of a puncture needle tube, a hose, etc. Generally made of austenitic stainless steel, polyvinyl chloride and other materials. Supplied sterile, single use. Used in conjunction with blood component collection machines (such as centrifugal type, rotating membrane type) or hemodialysis machines, it is used to collect blood from human veins or arteries, and return the processed blood or blood components to the human body. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 469
470 | 10-02-04 | Traditional disposable blood transfusion set with needle | Refer to GB 8369.1 "Disposable Blood Transfusion Set Part 1: Gravity Blood Transfusion Type" or GB 8369.2 "Disposable Blood Transfusion Set Part 2: Pressure Blood Transfusion Equipment" for performance, structure, composition, and use, and GB 18671 "Disposable Blood Transfusion Set " "Intravenous Infusion Needles" can fully cover traditional disposable infusion sets whose raw materials meet the requirements of GB15593 and other relevant standards, and are used for intravenous blood transfusion. Infusion patches are available (in a single package). The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product description. | 470
471 | 10-02-07 | Disposable tubing for frozen red blood cell washer | It usually consists of main pipeline (including silicone rubber pump tube, large drip pot, connecting tube), return blood vessel and waste liquid collection part (including waste liquid pipeline and waste liquid bag). Supplied sterile, single use. Tubing used for washing and separating erythrocytes and protection solution. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 471
472 | 10-03-05 | Hemodialysis water treatment equipment | It usually consists of tank filter, activated carbon filter, softener, precision filter, reverse osmosis device, power unit, disinfection device, monitoring device and delivery pipeline, and is used to prepare water for hemodialysis. | II |  | 472
473 | 10-03-05 | Hemodialyzer multiplexing machine | It usually consists of a control system, a monitoring system and a water system, and is used in conjunction with a special disinfectant solution for flushing, cleaning, testing, and filling of a special disinfectant solution for reusable dialyzers. | II |  | 473
474 | 10-03-05 | Electric dialysis chair | It usually consists of a seat, backrest, footrest, sliding footrest, pillow, armrest, lockable casters, push handle, and controller. It is used to electrically adjust the position of the back cushion, seat cushion, and foot pad to facilitate dialysis treatment for patients. | II |  | 474
475 | 10-03-05 | Blood flow monitoring system for hemodialysis | It usually consists of a host (electronic flowmeter), flow/dilution sensor, etc. | III |  | 475
476 | 10-03-07 | Peritoneal dialysis bag warmer | It usually consists of a heating plate, power supply and cable connector, and is used for warming the peritoneal dialysis bag before use. | II |  | 476
477 | 10-04 | peritoneal dialysis accessories | Includes Silicone Catheter Extension, Catheter Adapter, Retainer Stand, and Disc Retainer. Catheter extensions and catheter adapters are single-use sterile products. Used as an adjunct to peritoneal dialysis treatment. The basic principle, scope of application, performance and composition are the same as those of the products already on the market. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 477
478 | 10-04 | peritoneal dialysis external tube | It consists of a tube body, joints, locking tube joints and protective caps, etc., excluding Class III medical device components such as iodine solution protective caps. Supplied sterile. It is used in connection with the external part of the peritoneal dialysis tubing during peritoneal dialysis. The basic principle, scope of application, performance and composition are the same as those of the products already on the market. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 478
479 | 10-04-01 | Continuous renal replacement therapy tubing | It usually consists of arterial lines, venous lines, replacement fluid lines, etc. The components do not include hemodialysis filters, dialyzers, etc. The product is provided sterile for single use. Used in conjunction with continuous renal replacement dedicated dialyzer or hemodialysis filter, continuous renal replacement dedicated dialysis equipment, used in continuous renal replacement therapy, assuming the function of blood access". The exemption does not include the use of new materials and new functions Mechanism, new function products. | III |  | 479 | NEW
480 | 10-04-01 | Disposable hemodialysis tubing | It usually consists of arterial lines, venous lines, replacement fluid tubes and other necessary accessories, and does not include any vascular access devices (such as dialysis catheters, arteriovenous fistula puncture needles, etc.) and hemodialyzers, hemodialysis filters and Hemoperfusion etc. Supplied sterile, single use. Used in conjunction with hemodialyzer, hemodialysis filter and hemoperfusion, it is used in hemodialysis, hemodialysis filtration, hemoperfusion series hemodialysis treatment, and undertakes the function of blood access. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 480
481 | 10-04-01 | Disposable blood warmer tubing | It usually consists of pipelines, conduits, joints, protective caps, etc. It can be used with warming device in CRRT treatment to warm blood immediately. | III |  | 481
482 | 10-04-01 | hemodialysis concentrate | Usually consists of A agent and B agent. Agent A is usually composed of sodium chloride, potassium chloride, calcium chloride, magnesium chloride, sodium acetate/acetic acid/glacial acetic acid, and agent B is composed of sodium bicarbonate. Except for new ingredients, new formulations (supply status, mixing ratio, final solute concentration), new functions. | III | Add | 482 | NEW
483 | 10-04-03 | Dialysate filter | Usually consists of hollow fiber, sealant, housing, housing cover and gasket. Utilizing the function of the hollow fiber membrane, it is used to remove endotoxin, bacteria and insoluble particles in the dialysate. | III |  | 483
484 | 10-04-03 | Dialysis machine disinfectant | It usually consists of citric acid or glacial acetic acid and water. For cleaning and disinfection of dialysis machines. Raw materials are recommended to meet the requirements of the Pharmacopoeia. The exemption does not include products with new formulations, new mechanisms of action, or new functions. | III |  | 484
485 | 10-04-03 | Disposable rehydration tubing/replacement tubing | It usually consists of pipelines, joints, protective sleeves and clamps. Supplied sterile, single use. It is used as a pipeline for replenishing replacement fluid during hemodiafiltration and hemofiltration. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 485
486 | 10-04-04 | Disposable peritoneal dialysis tubing probe | Generally made of stainless steel. Supplied sterile, single use. Used to facilitate the introduction of acute and chronic peritoneal dialysis catheters. | II |  | 486
487 | 10-04-04 | Disposable peritoneal dialysis connector | Usually consists of a connector and a screw lock cap. Supplied sterile, single use. It is used for the connection of peritoneal dialysis catheter and external extension tube or peritoneal dialysis external short tube. | II |  | 487
488 | 10-04-04 | peritoneal dialysis external short tube | It usually consists of a tip protection cap, a switch, a sleeve, a tubing, a peritoneal dialysis tube connection port, a pull-ring cap, and a peritoneal dialysis fluid connection port. Supplied sterile, single use. It is used for aseptic connection and separation with peritoneal dialysis patient-side tubing (or titanium connector) and peritoneal dialysis fluid-side tubing. | II |  | 488
489 | 10-04-04 | peritoneal dialysis tubing | It usually consists of pipelines, connection ports, protective caps, etc. Generally made of polymer materials. Supplied sterile, single use. It is used to establish therapeutic access for peritoneal dialysis in patients with renal failure. | II |  | 489
490 | 10-04-04 | iodine protective cap | It usually consists of microcaps, sponges soaked in povidone-iodine solution, outer packaging and other components. Supplied sterile, single use. It is used to protect the connection between the convex interface of the peritoneal dialysis bag and the external pipeline. | II |  | 490
491 | 10-04-04 | Disposable peritoneal dialysis drain | It usually consists of a protective cap, connection interface, three-way perfusion tube, drainage tube and waste liquid collection bag. Supplied sterile, single use. | II |  | 491
492 | 10-04-04 | Disposable peritoneal dialysis tubing | It usually consists of cassettes, tubing racks, joint assembly assemblies, tubing, casting ports, clips (such as O-clamps), Y-connectors, pull-ring end protection caps, inner pull-ring caps, pull-ring caps, etc. Supplied sterile, single use. | II |  | 492
493 | 10-05-02 | Extracorporeal circulation continuous blood gas monitoring system | It usually consists of a host, a venous probe, an arterial probe, a sensor, a sample pool, etc. The measured data is transmitted to the host computer of the blood gas monitoring system through each sensor, so as to realize the monitoring of various parameters during extracorporeal circulation. Clinical verification data on the accuracy of blood gas monitoring should be provided. | III |  | 493
494 | 10-05-03 | Cardiopulmonary bypass equipment - heat exchange equipment | It usually consists of a water circulation system, a refrigeration system, and a heating system. It is used in conjunction with cardiopulmonary bypass equipment for temperature regulation of extracorporeal circulation in cardiac surgery. | II |  | 494
495 | 10-05-05 | Warmer | It is usually composed of a control device, a heating device, and a monitoring device, and is used for instant heating of the blood or liquid circuit in blood transfusion, infusion, hemodialysis and extracorporeal circulation to maintain the patient's body temperature. | II |  | 495
496 | 10-06-02 | venous blood reservoir | It usually consists of a top cover, a blood bank, a filter, a bleeding mouth, etc. Supplied sterile, single use. In clinical open-heart surgery extracorporeal circulation, it is used with a membrane oxygenator to store, defoam, and filter extracorporeal blood. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 496 | NEW
497 | 10-06-03 | Cardiopulmonary bypass blood filter | It is usually composed of the overall component shell, filter element, base, tee joint, emergency bypass pipe, matching component tee switch, pressure measuring tube assembly, exhaust pipe, etc. Supplied sterile, single use. The product refers to the mandatory industry standard YY 0580. It is used in open heart surgery with extracorporeal circulation to remove tiny solids and gas emboli in the blood. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 497
498 | 10-06-05 | Disposable cardioplegia infuser | It usually consists of a bubble trap, a temperature changer, connecting pipes and joints. Supplied sterile, single use. It is used to infuse myocardial cardioplegia into the coronary arteries to protect the myocardium during open-heart surgery with extracorporeal circulation. The product refers to the mandatory industry standard YY 0485. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 498
499 | 10-06-06 | Disposable cardioplegia perfusion tube | The perfusion tube is usually composed of tube body, joint, stainless steel needle core (optional), tube clamp and other components. Supplied sterile, single use. It is suitable for perfusing cardioplegia into the heart during cardiopulmonary bypass surgery, or connecting the patient's artery with the blood cardioplegia infusion device. | III |  | 499
500 | 10-06-06 | arterial cannulation, venous cannulation | It usually consists of suction head, handle, connector, cannula and other components. Supplied sterile, single use. Equipped with a cardiopulmonary bypass system, it is used for drainage or perfusion of blood during open-heart surgery performed by extracorporeal circulation. The product refers to the mandatory industry standard YY 0948. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 500
501 | 10-06-06 | Left heart drainage tube, right heart suction tube | It consists of three parts: suction head, tube body and joint. The product refers to YY 1271 regulation. It is used for exhausting the left heart in cardiovascular surgery, suctioning decompression or reducing the load of the left heart, and suctioning fluids such as blood in the field of cardiac surgery. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 501
502 | 10-06-06 | Cardiopulmonary bypass tubing Extracorporeal circulation tubing | It is usually composed of pipelines, joints, pump tubes and other components, and complies with the provisions of YY 1048. Supplied sterile, single use. It is used to provide a blood bypass channel during extracorporeal circulation or incomplete cardiopulmonary function support surgery. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 502
503 | 10-06-07 | centrifugal pump head | Usually consists of housing, vanes, shaft, bushing and magnets. Supplied sterile, single use. Used with a centrifugal pump for extracorporeal circulation or circulatory assistance in cardiac surgery. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 503
504 | 13-01-01 | metal locking plate | Made of stainless steel, pure titanium, and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, ISO 5832-3, and ISO 5832-11, through conventional machining, heat treatment and Surface treatment processes (such as passivation, electrolytic polishing, anodic oxidation, etc.), do not include products with new structural designs, and do not include innovative processes such as 3D printing. It is suitable for internal fixation of limb fractures. | III | Revised product description. | 504 | NEW
505 | 13-01-01 | Metal Locking Bone Screws | The products are made of stainless steel conforming to GB 4234.1, ISO 5832-1 standards and pure titanium, Ti6Al4V, Ti6Al4VELI, Ti6Al7Nb titanium alloy materials conforming to GB/T 13810, ISO 5832-2, ISO 5832-3, ISO 5832-11 standards. It is made through conventional machining, heat treatment and surface treatment processes (such as passivation, electrolytic polishing, anodic oxidation, etc.), excluding products with new structural designs, and excluding innovative processes such as 3D printing. It is suitable for internal fixation of limb fractures. | III | Revised product description. | 505 | NEW
506 | 13-01-01 | Metal Bone Screws (Nonlocking) | Refer to YY 0018 for the structure of this type of product. The product is made of stainless steel conforming to GB 4234.1 and ISO 5832-1 standards and pure titanium and Ti6Al4V conforming to GB/T 13810, ISO 5832-2, ISO 5832-3 and ISO 5832-11 standards. , Ti6Al4VELI, Ti6Al7Nb titanium alloy materials are made through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It can be used alone and is suitable for internal fixation of limb fractures. | III | Revised product description. | 506
507 | 13-01-01 | Cannulated Bone Screws (Nonlocking) | This type of product is made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, and ISO 5832-3. )production. It can be used alone and is suitable for internal fixation of limb fractures. | III | Revised product description. | 507
508 | 13-01-01 | Anatomical Metal Plates (Non-Locking) | Refer to YY 0017 for the structure of this type of product. It is made of stainless steel, pure titanium, and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, and ISO 5832-3. Made by heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for internal fixation of metaphyseal fractures of limbs. | III | Revised product description. | 508
509 | 13-01-01 | Straight Metal Plates (Non-Locking) | Refer to YY 0017 for the structure of this type of product. It is made of stainless steel, pure titanium, and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, and ISO 5832-3. Made by heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for internal fixation of long bone shaft fractures of limbs. | III | Revised product description. | 509
510 | 13-01-01 | Angled Metal Plates (Non-Locking) | Refer to YY 0017 for the structure of this type of product. It is made of stainless steel, pure titanium, and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, and ISO 5832-3. Made by heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for internal fixation of long bone shaft fractures of limbs. | III | Revised product description. | 510
511 | 13-01-01 | Staples | Refer to ISO 882 for the structure of this type of product. It is made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, and ISO 5832-3. 3D printing and other innovative processes). It is suitable for internal fixation of limb fractures. | III | Revised product description. | 511
512 | 13-01-01 | Bone plug | This type of product is composed of screw and nut, made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-3, through conventional machining, heat treatment and surface treatment processes (not Including innovative processes such as 3D printing). It is suitable for internal fixation of limb fractures. | III | Revised product description. | 512
513 | 13-01-01 | metal femoral neck screw | Refer to YY 0346 for the structure of this type of product, which is made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, and ISO 5832-3, and undergoes conventional machining, heat treatment and surface treatment processes (excluding 3D printing and other innovative processes). It is suitable for internal fixation of femoral neck fracture. | III | Revised product description. | 513
514 | 13-01-01 | Fixing nails for external fixation frame | This type of product is made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, and ISO 5832-3. )production. Cooperating with the external fixation bracket, it is suitable for partial implantation into the human body for fixation during the reduction of limb fractures. | III | Revised product description. | 514
515 | 13-01-01 | Acetabular screw | This type of product is made of stainless steel, pure titanium, titanium alloy, forged cobalt, etc. Chromium molybdenum materials are made through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for internal fixation of biological acetabular prosthesis. | III | Revised product description. | 515
516 | 13-01-01 | Bone defect filling block fixation screw | This type of product is made of stainless steel, pure titanium, titanium alloy, forged cobalt, etc. Chromium molybdenum materials are made through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for the fixation of metal bone defect filling blocks in hip joints and knee joints. | III | Revised product description. | 516
517 | 13-01-01 | Locking Ankle Plating System | This type of product is composed of locking bone plate and locking bone screw. The locking plate is made of stainless steel, Pure titanium and titanium alloy materials, locking bone screws are made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-3, ISO 5832-11, through conventional machining and heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for internal fixation of ankle fractures. | III | Revised product description. | 517
518 | 13-01-01 | Wrist Locking Plate System | These products consist of locking plates and locking bone screws. The locking bone plate is made of stainless steel, pure titanium, and titanium alloy that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, ISO 5832-3, and ISO 5832-11. The locking bone screw is made of GB Stainless steel and titanium alloy materials specified in 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-3, ISO 5832-11 standards, through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing) production. Suitable for internal fixation of wrist fractures | III | Revised product description. | 518
519 | 13-01-01 | Rib Plating System | This type of product consists of rib plate and screws. Made of pure titanium and titanium alloy materials that meet the standards of GB/T 13810, ISO 5832-2, ISO 5832-3, and ISO 5832-11, through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing) )production. It is suitable for internal fixation of rib fractures. | III |  | 519
520 | 13-01-04 | metal interlocking intramedullary nail | Made of stainless steel, pure titanium, and titanium alloy materials that comply with standards such as GB 4234.1 , GB/T 13810 , ISO 5832-1 , ISO 5832-2 , ISO 5832-3 , and ISO 5832-11 , through conventional machining, heat treatment and Made by surface treatment process (excluding innovative processes such as 3D printing). It is usually composed of main nails and / or accessories, and the nail holes can contain ultra-high molecular weight polyethylene that meets the requirements of GB/T 19701 or PEEK that meets the requirements of YY/T 0660 and other materials with mature clinical application history in this type of products made accessories. It is suitable for intramedullary cavity fixation of limb fractures. | III | Revised product description. | 520
521 | 13-01-04 | metal intramedullary nail | Refer to YY 0019 for the structure of this type of product, excluding intramedullary nails with locks. Made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, and ISO 5832-3, through conventional machining processes, heat treatment and surface treatment processes (excluding new technologies such as 3D printing and new process); suitable for internal fixation of limb fractures. Implanted into the bone marrow cavity and fixed by screws and other connections | III | Revised product description. | 521 | NEW
522 | 13-01-05 | flexible wire | Refer to YY/T 0816 for the structure of this type of product. It is made of stainless steel, pure titanium, and high-nitrogen stainless steel materials that meet the standards of YY 0605.9, ISO 5832-1, and ISO 5832-2. Printing and other new technologies and processes) are made. Suitable for bone fixation or bone joint . | III |  | 522
523 | 13-01-05 | Metal Cables and Cables | Refer to YY/T 0812 for the structure of this type of product, which may contain accessories used in conjunction with it (such as locking pins, restraints , locks, etc.). It is made of stainless steel, pure titanium, high-nitrogen stainless steel, titanium alloy materials in accordance with GB/T 13810, YY 0605.9, ISO 5832-1, ISO 5832-2 standards, 22Cr-13Ni-5Mn stainless steel materials in accordance with ASTMF1314 standards, and ISO5832 The forged cobalt-chromium-tungsten-nickel alloy material stipulated in the -5/YY 0605.5 standard is made through conventional machining and heat treatment processes (excluding innovative processes such as 3D printing). It is suitable for banding and internal fixation of limb fractures. | III | Revised product description. | 523
524 | 13-01-06 | metal spicule | This kind of products include Kirschner wires and Steiner wires, see YY 0345 for the structure. It is made of stainless steel and titanium alloy materials conforming to GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-3 standards, and forged cobalt-chromium-tungsten-nickel materials conforming to ISO 5832-5 standards through conventional machining, heat treatment and Surface treatment process (excluding innovative processes such as 3D printing). It is suitable for partial implantation into the human body for traction during the reduction of limb fractures, and external fixation brackets for limb fracture fixation, or implanted alone for internal fixation of limb fractures. | III | Revised product description. | 524
525 | 13-02-01 | non-absorbable threaded anchors | This type of product consists of anchors and sutures, and can be equipped with suture needles. The anchors are made of stainless steel, titanium alloy and polyether ether ketone materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-3, ISO 5832-11, and YY/T 0660, through conventional mechanical processing , injection molding, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). Sutures can be made of non-absorbable materials such as polyethylene, polypropylene, polyester, and polyamide, and the performance indicators refer to the applicable parts of the YY0167 standard. For suture needles, refer to the applicable parts specified in the YY/T 0043 standard. The product can be pre-installed with an inserter, and the part of the inserter in contact with the human body is usually made of stainless steel that meets the requirements of YY/T 0294, ASTM F899 or YY/T 0726. It is suitable for the connection and fixation of bone and soft tissue. | III | Revised product description. | 525
526 | 13-02-01 | nonabsorbable ligament fixation screws | This type of product is made of stainless steel and titanium alloy materials that meet the standards of GB 4234.1 , GB/T 13810 , ISO 5832-1 , and ISO 5832-3 . Through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing) ) or made of polyetheretherketone ( PEEK ) materials in accordance with YY/T 0660 . For the repair and reconstruction of ligaments. | III | Revised product description. | 526
527 | 13-02-01 | Titanium plate with loop | Products consist of titanium plates, coils and /or sutures. The titanium plate should be made of Ti6Al4V and Ti6Al4VELI titanium alloys that meet the standards of GB/T 13810 and ISO 5832-3, and the coils and sutures are made of polymer materials, such as polyethylene terephthalate, polypropylene, polyethylene , polyester material. It is suitable for the fixation of ligament or tendon and bone in orthopedic reconstruction. | III |  | 527
528 | 13-02-01 | Meniscus Suture System | The meniscus suture system generally consists of implants and auxiliary tools. The implants include anchors (rods) and sutures. The auxiliary tools include inserters, etc., which are not implanted in the body. The fixed anchor (rod) is made of polyetheretherketone (PEEK) material that meets the requirements of YY/T 0660 standard, and the suture is made of ultra-high molecular weight polyethylene (or main) material, and the performance index meets the requirements of YY0167 standard Applicable section. The inserter is made of a handle and a needle bar made of stainless steel, and the part of the inserter in contact with the human body is usually made of stainless steel that meets the requirements of YY/T 0294, ASTM F899 or YY/T 0726. It is suitable for meniscus suture operation under endoscope. | III | Add | 528 | NEW
529 | 13-03-01 | Laminar plate system | This type of product consists of a fixing plate and screws. The fixing plate is made of pure titanium and titanium alloy materials that meet the standards of GB/T 13810, ISO 5832-2, and ISO 5832-3. The screws are made of materials that meet the standards of GB/T 13810, ISO 5832-3 Titanium alloy materials specified in the standard are made through conventional machining, heat treatment and surface treatment processes (excluding innovative processes such as 3D printing). It is suitable for laminoplasty after expansion and decompression of the spinal canal. | III |  | 529
530 | 13-05-04 | stuffing sticks/blocks | The product is made of pure tantalum that complies with the ISO 13782 standard (excluding products prepared by innovative processes such as 3D printing). The product is used alone for bone filling of extremities. | III |  | 530
531 | 13-06-01 | Single/multi-component prefabricated cranioplasty plates and fasteners | Usually composed of plates and fasteners, generally made of pure titanium, titanium alloy, cobalt-chromium alloy and other metal materials and non-absorbable polymer materials, in which the plates are fixed by fastening devices (such as screws, wires or other components) bit. It is suitable for the repair of skull defects. | III | Add | 531
532 | 13-06-03 | head nail | The product is usually composed of a main body, a base, an O-ring, and a rubber cap assembly, of which only the main body is in contact with the human body. The main body should be made of stainless steel specified in GB 4234.1, ISO 5832-1, ASTM F899 or Ti6Al4V, Ti6Al4VELI titanium alloy specified in ISO 5832-3, GB/T 13810. Immobilize the patient's head and neck in a specific position during major revision surgery. This product is a non-implant, used only during the operation, in short-term contact with the human body, and removed immediately after the operation. | II | Revised product description. | 532
533 | 13-06-03 | Craniomaxillofacial Plate System | Made of stainless steel, pure titanium, and titanium alloy materials that meet the standards of GB 4234.1, GB/T 13810, ISO 5832-1, ISO 5832-2, and ISO 5832-3, through conventional machining, heat treatment, and surface treatment processes (excluding 3D printing and other innovative processes), excluding products with new structural designs. It is suitable for internal fixation of craniofacial fractures. Except for those who repair, cover, or fill bone defects or holes. | III |  | 533 | NEW
534 | 13-06-05 | Aneurysm clip | It is generally made of metal materials such as titanium alloy. A device used to occlude the neck or sac of an intracranial aneurysm. It is used for permanent occlusion treatment of intracranial aneurysm neck or tumor body, and can also be used for permanent occlusion treatment of arteriovenous malformation small vessels or arteriovenous fistula. The product is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 534
535 | 13-09-04 | Surgical Hernia Repair Mesh | External abdominal hernia repair patch: a patch implanted in the body (such as plain film, mesh plug, etc.), made of polypropylene, polyethylene terephthalate or polyvinylidene fluoride, with some absorbable materials, used for repairing hernia and abdominal wall defects outside the abdominal cavity; intra-abdominal hernia repair patch: the abdominal wall side is made of polypropylene, polyethylene terephthalate or polyvinylidene fluoride, and the abdominal side is made of materials that can reduce adhesions, used for intra-abdominal hernia and abdominal wall repair. Wall defects are repaired. The above two types of patches are equivalent to products that have been approved for domestic registration, and the composition materials are mature, and materials that can reduce adhesion have been applied in similar products on the market. New structural design, new mechanism of action, and drug-device combination products are excluded. Excludes products that are ultimately fully absorbed and products that are repaired with autologous tissue reconstruction. | III |  | 535
536 | 13-09-06 | Biliary/pancreatic duct/ureteral stent | Stents are tubular structures made of non-degradable/absorbable polymeric materials (such as polyethylene, polyurethane, etc.), can have components such as radiopaque markers, can be coated, and can contain delivery components. It is used to implant the cavity under the endoscope, expand the stenosis of the implanted site, and perform drainage. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III |  | 536
537 | 13-09-06 | non-vascular self-expanding metallic stent | Stents are made of metals with shape memory properties (such as Nitinol, etc.), can have radiopaque markers, can be with or without coating, and can contain delivery components. It is used to dilate stenosis in non-vascular parts such as esophagus, biliary tract, airway, and intestinal tract. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III |  | 537
538 | 13-09-10 | tissue expander | It generally consists of a housing, a catheter, an injection seat (or an equivalent component), and a connector (if any). The main material of the shell is silicone rubber. The principle of action is that it can be filled with saline to increase its own volume to expand the tissue. It is used to obtain autologous skin tissue to solve the shortage of skin donor area. It is used during surgery or removed within 6 months of implantation. It is equivalent to products that have been approved for domestic registration. Excluding new structural design and drug-device combination products. | III |  | 538
539 | 13-04-01 | Centralizer | This type of product is made of PMMA and ultra-high molecular weight polyethylene materials that meet the requirements of GB/T19701.2. It is suitable for keeping the distal end of the femoral stem in the middle during cemented hip replacement surgery, blocking the bone cement and ensuring the uniform distribution of the bone cement around the prosthesis. | III | Revised classification codes. | 539
540 | 13-04-01 | distal plug | This type of product is made of PMMA and ultra-high molecular weight polyethylene materials that meet the requirements of GB/T19701.2. It is suitable for blocking bone cement in cemented hip replacement. | III | Revised classification codes. | 540
541 | 13-04-01 | Hole plug for acetabular cup | The product is made of Ti6Al4V titanium alloy that meets the standards of GB/T 13810 and ISO 5832-3. It is used with the acetabular cup to fill the hole of the acetabular cup. | III | Revised classification codes. | 541
542 | 13-04-01 | medullary plug | The product is made of ultra-high molecular weight polyethylene material that complies with GB/T 19701.2 and ISO 5834-2 standards. It is used for bone cement type prosthesis implantation, and its function is to prevent bone cement from overflowing into the limited area of the medullary cavity. The applicable parts are the medullary cavity of femur, tibia and humerus. | III | Revised classification codes. | 542
#N/A
#N/A
#N/A
543 | 13-11-03 | sternal ligature | The product is made of Ti6Al4V, Ti6Al4VELI, Ti6Al7Nb titanium alloy materials and stainless steel materials that meet the standards of GB/T 13810, ISO 5832-3, ISO 5832-11, and ISO 5832-1. The product is usually composed of two parts, a ligature and a threading hook, and may contain a lock. The threading hook is not implanted in the human body, and is only used to assist the threading and installation of the ligature, and is discarded after installation. For sternal fixation during thoracotomy. Products with completely new structural designs are not included. | III | Revised product description. | 543 | NEW
544 | 14 | Medical masks (non-surgical) | Medical masks (non-surgical) are usually made of non-woven fabrics or non-woven fabric composite materials, which can have a two-layer or three-layer structure, and can have plastic nose clips. The mask belt can be elastic or non-elastic, and has the ability to filter particles and bacteria. It can be worn by various clinical personnel during non-invasive operations, covering the user's mouth, nose and jaw, and providing a certain physical barrier to prevent the direct penetration of pathogenic microorganisms and particulate matter. The exemption does not include products with new features such as anti-virus. | II | Revised product description. | 544
545 | 14 | self-locking syringe | It is generally composed of a jacket, a conical joint, a core rod, a collar, a piston, and a sealing "O" ring. It is expected to be used for interventional therapy, radiography surgery, or pre-radiotherapy inspection balloon or suction. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 545
546 | 14-01 | Disposable bioprotein glue dispenser | This product can be composed of a dispenser, a double-chamber liquid pusher, a bracket, a tee, a spray head and a spray needle. This product is sterilized and disposable. It is used to prepare and mix protein glue components and then spray it on the surface of human wounds. | II |  | 546
547 | 14-01-01 | Injection pump | A syringe pump product is a device that controls the flow of fluid into a patient's body through one or more single-action syringes or similar containers (for example, by pushing a plunger to empty the barrel), the infusion rate is set by the operator, and the unit is indicated by the device The flow rate within a certain period of time; it is generally composed of pump housing, motor drive system, input system, storage system, control system, display system, sensor monitoring system and alarm system. It is divided into several models according to design, technical parameters, intended use, additional auxiliary functions, etc.; it is matched with specific infusion equipment, and is used for long-term constant drug delivery speed and precise dosage infusion of arteries and veins. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.27-2005 Medical Electrical Equipment Part 2-24: Special Safety Requirements for Infusion Pumps and Infusion Controllers. | II |  | 547
548 | 14-01-02 | Disposable sterile syringe | Refer to GB 15810 "Single-use Sterile Syringe" for performance, structure, composition, and use, etc. Products that can be fully covered are generally composed of a syringe jacket, piston, sealing ring, core rod, pusher, and cone, and are used for pumping liquids. Or a manual syringe for injection immediately after filling the liquid. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 548
549 | 14-01-02 | Disposable sterile syringe with needle | Refer to GB 15810 Disposable Sterile Syringes and GB 15811 Disposable Sterile Injection Needles for performance, structure, composition, use, etc. Products that can be fully covered are the single-use sterile syringes in the published exemption catalog in conjunction with the published exemption catalog Single-use sterile injection needles. It is generally composed of a syringe jacket, piston, sealing ring, core rod, pusher, tapered needle seat, connecting part, needle tube, and sheath, and is used for intradermal, subcutaneous, intramuscular, and intravenous injection or extraction of medicinal liquid. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 549
550 | 14-01-02 | Disposable low resistance syringe with needle | Refer to the products that can be fully covered by YY/T0909 "Single-Use Low Resistance Syringes" for performance, structure, composition, and use. The product is generally composed of a jacket, a piston cap, a piston, a core rod and an injection needle (optional), supplied aseptically, and is used for aspiration of liquid or injection immediately after injecting the liquid. | III |  | 550
| #N/A |  | The exemption does not include products with new materials, new mechanisms of action, or new functions.
551 | 14-01-02 | Disposable sterile insulin syringe | Refer to YY 0497 "Sterile Insulin Syringes for Disposable Use" for performance, structure, composition, use, etc. Products that can be fully covered are mainly composed of a jacket, a core rod, a rubber piston, an injection needle, and a sheath/end cap. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 551
552 | 14-01-02 | Disposable Pump Syringes | It is generally composed of a syringe jacket, piston, sealing ring, core rod, pusher, and cone, with or without a needle. For aspiration of liquid or injection immediately after infusion, used with a syringe pump. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 552
553 | 14-01-04 | Pen Insulin Syringes | The pen-type insulin syringe consists of a dose adjustment button, a pen body, a release button, a lead screw, a cartridge holder, and a pen cap, excluding drugs and injection needles, and all parts are not in contact with liquid medicine; they must be compatible with specific insulin and injection needles or standard Injection needle matching for subcutaneous injection of insulin. | II |  | 553
554 | 14-01-04 | pen injector | Usually consists of a pen cap, cartridge holder, screw rod, pen body, dose adjustment plug, and injection button; active pen injectors usually also contain other electronic components with auxiliary functions. Used in conjunction with a pen core and/or a needle, the liquid medicine can penetrate the surface of the skin or mucous membrane through pressure and be delivered into the body. For the injection of liquid medicine (such as insulin). | II |  | 554
555 | 14-01-05 | All Glass Syringes | The all-glass syringe is made of borosilicate-aluminum glass material, which consists of a cone head, a jacket and a core. It can be divided into two types: the middle head type and the partial head type. Each type is divided into several specifications according to the nominal capacity; the injection needle is installed , for the human body for subcutaneous, intramuscular, intravenous injection of liquid medicine or extraction of liquid and so on. It can also be reused after being cleaned and sterilized. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1001.1 "Glass Syringes Part 1: All Glass Syringes". | II |  | 555
556 | 14-01-05 | Blue Core All Glass Syringe | The blue-core all-glass syringe is made of borosilicate-aluminum glass material, consisting of a cone head, a jacket and a core (blue solid glass), usually a middle head type, divided into several specifications according to the nominal capacity; when the injection needle is installed, It can be used for biochemical tests and subcutaneous tests on human body. It can be used for vaccine injection, oral anesthesia, and also for injection of other medicinal liquids. The product performance index adopts the applicable parts of the following reference standards, such as: YY 1001.2 "Glass Syringe Part 2: Blue Core Full Glass Syringe". | II |  | 556
557 | 14-01-06 | metal injection needle | Metal injection needles are usually made of copper-based chrome-plated (needle seat) and high-grade stainless steel (needle tube), generally composed of a needle seat and a needle tube, matched with a syringe, and used for subcutaneous, intradermal, muscle, oral and other injections It can be used repeatedly for medicine, vaccine, anesthetic or intravenous infusion and blood transfusion. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0282 "Injection Needles". | II |  | 557
558 | 14-01-06 | Disposable sterile injection needles | Refer to GB 15811 "Single-Use Sterile Injection Needles" for performance, structure, composition, use, etc. Products that can be fully covered. It is composed of a needle base, a connecting part, a needle tube, and a sheath, and is used for intradermal, subcutaneous, muscle, and vein of human body. Inject or withdraw the liquid medicine. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 558
559 | 14-01-06 | Needles for single-use injection pens | It consists of a sealing film, a needle tube, a needle seat, an outer sheath, and an inner sheath. The product is sterile and pyrogen-free. Used in conjunction with the injection pen to inject drugs into the human body subcutaneously. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 559
560 | 14-01-06 | Injection needle (Ⅱ) | It usually consists of a needle tube, a needle hub and a sheath. The needle tube and needle seat are generally made of stainless steel. Supplied non-sterile and reusable. It is used for intradermal, subcutaneous, submucosal, intramuscular injection or extraction of liquid in human body. The exemption does not include products with new materials, new mechanisms of action, or new functions. | II |  | 560
561 | 14-01-06 | Injection needle (Ⅲ) | It usually consists of a needle tube, a needle seat and a sheath, and can be equipped with a self-destruct device. The needle tube is generally made of stainless steel, and the needle seat is generally made of polymer material. Supplied sterile. It is used for injecting or extracting liquids into the skin, subcutaneous, submucosal of the digestive tract, muscle, and vein. The exemption does not include products with new materials, new mechanisms of action, new structures, or new functions. | III |  | 561
562 | 14-01-07 | Syringe Auxiliary Pusher | It is generally made of metal materials, polymer materials, etc. No contact with the injection liquid. Used with syringes, etc., for auxiliary injection of syringes. The exemption does not include products with new materials, new mechanisms of action, or new functions. | II |  | 562
563 | 14-01-08 | piercing instrument | It usually consists of a puncture needle, a puncture device, and a protective sheath. | II |  | 563
564 | 14-01-08 | metal thoracentesis needle | Metal thoracentesis needles can be made of copper-based chrome-plated (needle seat, etc.) and high-grade stainless steel (needle tube), and can be composed of a needle seat (with a standard Luer connector), a tee, a positioning device, and a graduated needle tube. The outer diameter is 1.2mm or 1.6mm, and the length of the needle tube can be 65mm-65mm or longer; matched with the syringe, it is used for thoracentesis, channel establishment, and extraction of pleural effusion, etc., and can be used repeatedly. | II |  | 564
565 | 14-01-08 | cricothyropuncture needle | The cricothyroid puncture needle can be made of metal materials such as stainless steel, and is composed of a cannula with a stop structure and a matching tapered puncture needle. It can be divided into several models and specifications according to different materials, designs, and technical parameters; Thyroid puncture, establishment of artificial airway, relief of upper airway obstruction. | II |  | 565
566 | 14-01-08 | Disposable paracentesis needle | The composition is limited to the paracentesis needle composed of a needle tube, a needle seat and a connecting tube, and is used for performing peritoneal puncture on a human body. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 566
567 | 14-01-08 | Puncture needles, trocars (for lumbar, blood vessel, ventricle puncture) | It usually consists of a puncture needle, a puncture device, and a protective sheath. It is used to puncture the lumbar spine, blood vessels, and ventricles to collect human samples, inject drugs and gases, etc., or as a channel for other instruments to enter the body. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 567
568 | 14-01-08 | metal lumbar puncture needle | Metal lumbar puncture needles can be made of copper-based chrome-plated (needle seat, bushing core) and high-grade stainless steel (needle tube, bushing core), which can be made of bushing core seat, positioning pin, needle seat (with standard Luer connector), bushing core , needle tube, usually the nominal outer diameter of the needle tube is 0.4mm-1.6mm, the length of the needle tube is 25mm-200mm, the wall of the needle tube can be divided into two types: normal wall and thin wall; matching with the syringe, it is used for human subarachnoid space block (Lumbar anesthesia) for puncture, injection of medicine or extraction of cerebrospinal fluid, reusable. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1148-2009 lumbar puncture needle. | III |  | 568
569 | 14-01-08 | Epidural puncture indicator balloon | It usually consists of a cavity, a one-way valve, an indicating membrane, an O-ring, etc. Provides a clear, visual indication of the needle tip reaching the epidural space during epidural anesthesia puncture. | II |  | 569
570 | 14-01-09 | biopsy needle | It usually consists of a needle seat, a core needle, an inner needle tube, an inner and outer needle positioning sheath, an outer needle tube, a protective sleeve, and a tissue collection device. The needle tube is generally made of stainless steel. | II |  | 570
571 | 14-01-09 | Disposable Biopsy Needle | The composition is limited to biopsy needles including puncture needles and needle cores. The materials meet the requirements of GB 18457. Pathological tissues are obtained by directly puncturing human tissues and organs, excluding those used in conjunction with endoscope systems, negative pressure suction devices, etc. or requiring ultrasound Biopsy needles operated under the guidance of imaging equipment such as , X-ray, etc. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 571
572 | 14-01-09 | biopsy needle | It consists of needle tube, needle core, handle and other components. It is used to obtain bone marrow biopsy of patients, or under B-ultrasound or X-ray monitoring, percutaneous puncture for cytological biopsy or other soft tissue biopsy of solid organs or tumors. Single use. The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. | II |  | 572
| #N/A |  | The exemption does not include: the use of new materials; products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application.
573 | 14-01-09 | coaxial biopsy needle | It consists of a stainless steel puncture needle with a needle seat (i.e., the outer cannula), a needle core with a needle seat, and a plastic ring (used to cover the stainless steel to mark the puncture depth), and is used to obtain soft tissue biopsy specimens, not for bone biopsy. The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. | II |  | 573
| #N/A |  | The exemption does not include: the use of new materials; products containing special coatings such as polymers, drugs, and biological products, and products with special structural functions and scope of application.
574 | 14-01-10 | biopsy gun | It usually consists of ejection, release and braking parts. Supplied non-sterile and reusable. When assembled with a biopsy needle, it is used to obtain a specimen from human tissue for biopsy. | II |  | 574
575 | 14-02-01 | Infusion Pump | The infusion pump is mainly composed of a control system, a motor drive unit, a peristaltic extrusion mechanism, a detection device, an alarm device, an input and display device, a housing and its supporting structure, and software components. Different functions are divided into several models; it is matched with the infusion set, does not touch the infusion liquid, and is used by the hospital to provide patients with intravenous infusion of medicinal liquid in an adjustable manner. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.27-2005 Medical Electrical Equipment Part 2-24: Special Safety Requirements for Infusion Pumps and Infusion Controllers. | II |  | 575
576 | 14-02-00 | Liquid reservoir for portable electric infusion pump | It usually consists of a liquid reservoir, a pipeline, a shut-off clip, an inner conical locking joint and a protective cap. Does not contain liquid medicine. Supplied sterile, single use. It is used in conjunction with specific analgesic pumps, chemotherapy pumps and other portable electric infusion pumps to infuse medicinal solutions to patients. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised classification code, product name, product description and product management category description. | 576
577 | 14-02-02 | Infusion Information Collection System | The infusion information collection system consists of a mobile stand with a communication interface and power supply, and integrated software; it supplies power to the infusion pump/syringe pump, and communicates data with the specified model of the infusion pump/syringe pump through the communication interface; collects data from the infusion pump/syringe pump The data is transmitted to the central workstation through the wired/wireless network, and the alarm information is prompted. | II/III | Revised product management category description. | 577
578 | 14-02-03 | Infusion Assist Electronics | Usually used in conjunction with an infusion set, it is an electronic instrument that can realize functions such as flow control, heating, and alarm during the infusion process. | II |  | 578
579 | 14-02-04 | Disposable Infusion Pump | It usually consists of an elastic medicine storage bag (without medicine), a dosing device, an extension tube and a flow rate controller. Use mechanical elasticity as kinetic energy to provide power for the pump body. Supplied sterile, single use. It is used for patient self-control to adjust the flow of medicinal liquid injected into the body (vein, subcutaneous, epidural space). The exemption does not include products using new materials, new mechanisms of action, or new functions. | III | Add | 579 | NEW
580 | 14-02-05 | Disposable light-proof infusion set with needle/disposable light-proof infusion set | Performance, structure, composition, use, etc. refer to GB 8368 "Gravity Infusion Type of Disposable Infusion Set", YY 0286.3-2017 "Special Infusion Set Part 3: Disposable Light-proof Infusion Set", GB 18671 "Disposable Intravenous Infusion Set" Infusion Needle" is a traditional disposable infusion set that can be fully covered, and is used with infusion containers and infusion needles for intravenous infusion. It can be equipped with infusion stickers, and the cork piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 580
581 | 14-02-05 | Traditional Disposable Infusion Set | Refer to GB 8368 "Gravity Infusion Type of Disposable Infusion Sets" for performance, structure, composition and use. The traditional disposable infusion set can be fully covered. It consists of a bottle stopper piercer and a protective cover, an air inlet with an air filter and a stopper, a liquid channel, a dropper, a drip funnel, a liquid medicine filter, a pipeline, a flow regulator, an injection part, an external cone joint and a protective cover. It is used with an infusion container and an infusion needle for intravenous infusion. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 581
582 | 14-02-05 | Traditional disposable infusion set with needle | The performance, structure, composition and use of the traditional disposable infusion set comply with GB 8368 "Gravity Infusion Type of Disposable Infusion Set" and GB 18671 "Disposable Intravenous Infusion Needle". It is used with infusion container and infusion needle. For intravenous infusion. It can be equipped with infusion stickers, and the cork piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 582
583 | 14-02-05 | Disposable precision filter infusion set with needle | Performance, structure, composition, use, etc. conform to YY0286.1 "Special Infusion Set Part 1: Disposable Microporous Filter Infusion Set", GB 18671 "Disposable Intravenous Infusion Needle" traditional disposable infusion set, and The infusion container and the infusion needle are used for intravenous infusion. It can be equipped with infusion stickers, and the cork piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product description. | 583
584 | 14-02-05 | Disposable vial and bag infusion set with needle | Performance, structure, composition, use, etc. conform to YY0286.5 "Special Infusion Sets Part 5: Disposable Bottle-type and Bag-type Infusion Sets", GB 18671 "Disposable Intravenous Infusion Needles" traditional disposable infusion device for intravenous infusion. It can be equipped with infusion patches (located in a single package), and the cork piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product description. | 584
585 | 14-02-05 | Disposable burette infusion set with needle | Performance, structure, composition, use, etc. conform to YY0286.2 "Special Infusion Set Part 2: Disposable Burette Infusion Set Gravity Infusion Type", GB 18671 "Disposable Intravenous Infusion Needle" traditional disposable infusion set , for intravenous infusion. It can be equipped with infusion stickers, and the cork piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 585
586 | 14-02-05 | Disposable flow setting fine-tuning infusion set with needle | The performance, structure, composition, and usage conform to YY0286.6 "Special Infusion Set Part 6: Special Infusion Set Part 6: Disposable Infusion Set with Fine-Tuning Flow Setting", GB 18671 "Disposable Intravenous Infusion Needle" The traditional disposable infusion set can be equipped with an infusion patch, and the cork piercer can be double-ended or multi-ended. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 586
587 | 14-02-05 | Disposable pressure infusion set | The performance, structure, composition and use of traditional disposable infusion sets comply with YY/T 0286.4 "Special Infusion Sets Part 4: Infusion Sets for Disposable Pressure Infusion Equipment". It is used with infusion containers and infusion needles. For intravenous infusion, it can be equipped with infusion stickers, and the bottle stopper piercer can be double-headed or multi-headed. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product description. | 587
588 | 14-02-05 | Infusion set | It usually consists of Luer conical connector, pipeline, drip funnel, flow regulator, cork piercer, liquid medicine filter, etc. Some infusion sets have air intake parts for air filters and liquid medicine injection parts. Tubing is generally made of polyvinyl chloride or other materials. Its design enables it to infuse the drug solution in the infusion container into the vein through the venipuncture device under the action of gravity or pressure. Supplied sterile, single use. For intravenous infusion of liquid medicine. The exemption does not include products with new materials, new mechanisms of action, new structures, or new functions. | III |  | 588
589 | 14-02-05 | Disposable pump infusion set | It usually consists of Luer cone connector, pipeline, dripping funnel, flow regulator, cork piercer, liquid medicine filter, etc. Some infusion sets are equipped with air intake parts and liquid medicine injection parts of air filters. Tubing is generally made of polyvinyl chloride or other materials. Used in conjunction with the infusion pump, the medicinal solution in the infusion container is infused into the vein through the venipuncture device. Supplied sterile, for intravenous infusion. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 589
590 | 14-02-06 | Disposable IV needle | Products conforming to GB 18671 "Disposable Intravenous Infusion Needles" in terms of performance, structure, composition, and use are generally composed of needle tubes, needle handles, hoses, connecting seats and protective caps, and are used in conjunction with infusion sets and blood transfusion sets. Establish peripheral venous access. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 590
591 | 14-02-06 | Disposable needles for implantable drug delivery devices | Products that conform to YY 0881 "Special Needles for Disposable Implantable Drug Delivery Devices" in terms of performance, structure, composition, and use. Injection drug delivery device and infusion set are used together. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 591
592 | 14-02-07 | Disposable intravenous indwelling needle | Products that conform to YY/T 1282 "Disposable Intravenous Catheters" in terms of performance, structure, composition, and use. Generally, it consists of two parts: a catheter assembly and a needle assembly. The catheter assembly is an assembly consisting of a catheter, catheter hub, extension tube (if any) and any integral connector. The needle assembly is composed of a needle tube and a needle hub and/or other integrated An assembly consisting of accessories. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III | Revised product description. | 592
593 | 14-02-07 | arterial catheter | The arterial indwelling needle is composed of a flow control valve, a needle handle, a catheter seat, a catheter, a protective sleeve, a silicone tube, an exhaust connector, a needle tube, a stainless steel ball, a guide sleeve and a catheter sleeve. Supplied in sterile form; the product is intended for use in arterial pressure monitoring and continuous arterial blood gas monitoring. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 593
594 | 14-02-07 | single-use thoracentesis needle | The composition is limited to a thoracentesis needle composed of a needle tube, a needle seat and a connecting tube, and is used for thoracentesis on a human body. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 594
595 | 14-02-08 | Disposable pressure extension tube | The disposable pressure extension tube is mainly composed of inner and outer conical Luer connectors, extension tube and/or three-way valve, and is connected with a catheter entering human blood vessels when used. The intended use is to connect pipelines and equipment in interventional operations, provide an infusion channel for medical fluids and/or be used for invasive blood pressure monitoring. Raw materials should meet GB/T 15593 and other relevant standards. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III | Revised product description. | 595
596 | 14-02-08 | Infusion connection line | It usually consists of an infusion line, at least one male conical luer connector and one female conical luer connector. Tubing is generally made of polyvinyl chloride or other materials. Supplied sterile, single use. It is connected with other infusion devices through the Luer connector to realize the infusion of the drug solution after the pipeline connection between the two devices. The exemption does not include products with new materials, new mechanisms of action, new structures, or new functions. | III |  | 596
597 | 14-02-08 | Infusion connection line | It usually consists of an infusion line, an outer conical luer connector and an inner conical luer connector. Single use. It is usually composed of inner and outer conical Luer connectors, extension tubes and/or three-way valves. When used, it is connected to catheters entering human blood vessels or used to connect equipment, provide infusion channels for medical fluids, and/or be used for invasive blood pressure monitoring. Including normal pressure and pressure infusion. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 597
#N/A
598 | 14-02-09 | Connections and accessories for infusion | Needle-free connectors: products generally consist of interfaces, connectors and tubing (optional). It is a non-puncture infusion connector, used in conjunction with intravascular indwelling catheters, through which to infuse medicinal fluid into blood vessels or draw blood. Three-way (valve): generally consists of a shell, a protective sleeve, a valve, a rotary handle, a flag plug, and a nut. This product is used in connection with infusion equipment, and is used to connect in series to the infusion system to add one or more specific functions, such as liquid circuit switch, backflow prevention, filtration, etc. Heparin caps: Refer to YY 0581 Heparin Caps for Infusion for performance, structure, composition and use, which are mainly used in the form of puncture. The medicine is injected into the blood vessel. Connecting plate: The product is generally composed of joints, connecting main boards, switches (valves), and hoses, which are used as connecting channels to establish multiple channels to provide liquid medicine or contrast agent injection. Filter: usually consists of a filter membrane, a silica gel disc, a Luer seat, a Luer connector and a protective cap. It is used to be connected with a warm liquid instrument and a warm liquid pipeline. During gravity infusion, it eliminates the microporous bubbles generated when the liquid is heated, and does not include the function of filtering out particles or bacteria. The product performance indicators adopt the applicable parts of the following standards, such as: GB8368 "Gravity Infusion Type of Disposable Infusion Set"; YY 0585.3 "Disposable Liquid Circuit and Accessories for Pressure Infusion Equipment Part 3: Filter". The above products The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 598
599 | 14-02-09 | Connections and accessories for infusion | Usually with at least one luer conical connector, generally made of polymer materials. Supplied sterile, single use. It is used to be connected in series to the infusion system to add one or more functions, such as liquid circuit switch, backflow prevention, filtration, etc. Common such as: needleless connectors, three-way (valve), heparin caps, connecting plates, filters, etc. The exemptions for the above products do not include products with new materials, new mechanisms of action, or new functions. | III |  | 599
600 | 14-02-11 | Intravenous nutrition infusion bags, infusion bags for pumps | It usually consists of a bag body, a liquid medicine inlet and an infusion set socket, and is generally made of polymer materials. Supplied sterile, single use. Used in medical institutions, filled with nutrient solution or medicinal solution before use, and then connected with infusion sets and intravenous devices (such as central venous catheters) for infusion into the body. The exemption does not include products with new materials, new mechanisms of action, new structures, or new functions. | III |  | 600
601 | 14-02-11 | Disposable intravenous nutrition infusion bag | Refer to YY 0611 "Disposable Intravenous Nutrition Infusion Bag" for performance, structure, composition, and use. It is usually composed of a liquid storage bag, a drug solution inlet, and an infusion set socket. Some infusion bags have infusion lines. It is generally made of polymer materials, provided sterile, and disposable. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 601
602 | 14-02-11 | Disposable pump infusion bag | It usually consists of a bag body, a liquid medicine inlet and an infusion set socket, and is generally made of polymer materials. Used in medical institutions, filled with liquid medicine before use, and then connected to an infusion pump for infusion into the body. Supplied sterile, single use. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 602
603 | 14-02-12 | Transfer and dispensing equipment for liquid medicine | It usually consists of a cork piercer, connecting pipelines, etc., and can be equipped with a filter. When it is connected with two containers, it can realize the transfer, dissolving, dispensing or multiple infusion (multi-bottle infusion) between different medicaments. The product is supplied sterile. Including: Disposable medicine dispensing syringe: It can be composed of protective cap, medicine dispensing needle, jacket, core rod and piston. It can be divided into several specifications according to the nominal capacity. It can be equipped with side hole or inclined plane medicine dispensing needle for extracting or preparing medicine liquid. use. Disposable drug dispensing needle: It consists of a needle seat, a needle tube and a protective cap, which can be divided into side holes and slopes according to different types, and is matched with a disposable drug dispensing syringe for extracting and preparing medicinal liquid. Liquid medicine transfer device: the transfer device can have an outer cover, and can also be connected with an air filter or liquid medicine filter; the piercer can be made of stainless steel or plastic, and other components can be made of plastic such as polyvinyl chloride. | II |  | 603
604 | 14-03-01 | Enteral Nutrition Pump | The enteral nutrition pump is mainly composed of the host, power adapter, built-in battery and other parts, and can be divided into several models according to the working principle, design, technical parameters, additional auxiliary functions, etc. No contact with the infusion liquid, it is used by the hospital to infuse the patient's gastrointestinal tract with nutrient solution in an adjustable way, and it is only suitable for enteral nutrition infusion. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 9706.1-2007 Medical Electrical Equipment Part 1: General Safety Requirements. | II |  | 604
605 | 14-03-03 | Hypodermic Infusion Set for Disposable Insulin Pumps | It usually consists of a hypodermic needle, a fixation device, tubing, connectors, etc. Supplied sterile, single use. Used in conjunction with specific insulin pumps and liquid reservoirs for insulin pumps to provide continuous subcutaneous infusion of insulin to diabetic patients. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 605
606 | 14-03-04 | Disposable insulin pump reservoir | It usually consists of core rod, piston, jacket and so on. Does not contain insulin. Supplied sterile, single use. It is used in conjunction with specific insulin pumps and insulin pump infusion sets to continuously infuse insulin into diabetic patients. The exemption does not include products with new materials, new mechanisms of action, or new functions. | III |  | 606
607 | 14-03-05 | enteral feeder | It usually consists of a cork piercer, a pipeline, a drip funnel, and an enteral nutrition catheter connection joint. The pipeline is generally made of polyvinyl chloride or other polymer materials, and can be connected to an enteral nutrition container and an enteral nutrition tube. Some enteral nutrition devices also have pump tubes, so that they can infuse nutrient solution into the intestines under the action of the enteral nutrition pump; some enteral nutrition devices are connected with enteral nutrition bags to deliver nutrient solution to the gastrointestinal tract. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 607
608 | 14-03-06 | enteral nutrition bag | It usually consists of a nutritional preparation inlet, a bag body and a puncture interface of an enteral nutrition device. The bag body is generally made of polyvinyl chloride or other polymer materials. Contains no nutrients. It is used to deliver nutrition to the patient's stomach by connecting a nasogastric tube or gastric tube. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 608
609 | 14-04-01 | Electronic Pneumatic Tourniquet | It is composed of a tourniquet host, a trachea and sensors, etc. It is used to temporarily block the blood supply of the limbs of the patient during the operation, and provides a bloodless surgical field of view for the operation. | II |  | 609
610 | 14-04-02 | passive hemostat | Usually consists of a compression device and a fixation device. The compression device can be divided into compression pads, compression airbags, etc., and the types of fixing devices are divided into nuts, fixing plates, fixing belts, adhesive stickers, etc. Supplied sterile. | II |  | 610
611 | 14-04-03 | passive tourniquet | It is usually divided into binding type, buckle type, and elastic band sticking type, which pressurize and compress blood vessels to achieve the purpose of hemostasis. It is used for clinical venous blood drawing, hemostatic tying of venipuncture, and hemostasis treatment at the scene of an accident and battlefield rescue of bleeding limbs. Or used in extremity surgery to drive blood from extremities. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 611
612 | 14-04-03 | Femoral artery pneumatic compression device | Femoral artery pneumatic compression device is composed of sterile pneumatic pressure dome arch, inflation pipeline with pressure gauge, inflation source, straps and other accessories, which can be divided into several types according to design, technical parameters, auxiliary functions, etc.; Compression of femoral artery or vein after vascular catheterization for hemostasis or ultrasound-guided compression repair of femoral artery pseudoaneurysm. | II |  | 612
613 | 14-05-01 | percutaneous enteral nutrition catheter | Usually a catheter made of materials such as silicone rubber or polyurethane, with a fixation device at the tip. The function of the balloon is to insert the catheter into the stomach or jejunum percutaneously, and then inflate the balloon to fix the catheter. The residence time in the body is greater than or equal to 30 days. Supplied sterile, single use. It is used to deliver nutrients to patients who cannot receive nasoenteral nutrition. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 613
614 | 14-05-02 | Nasoenteral nutrition catheter | Usually consists of tubing made of polyvinyl chloride or other material and a fitting that cannot be connected to the male luer connector of the IV set. It can be inserted into the patient's stomach or duodenum through the nose to give enteral nutrition or medicine through it. Supplied sterile, single use. It is used to introduce nutrient solution to the gastrointestinal tract through the nose. Part of it can realize other auxiliary functions such as flushing. | II |  | 614
615 | 14-05-03 | urinary catheter | Generally made of polymer materials. Part of the head end is fixed with a balloon. The tip can be inserted into the bladder and catheterized outside the body. Some catheters are coated with a lubricating coating (with no drug, silver salt or antibacterial component coating), which is easy to insert after soaking and reduces the pain of intubation. Supplied sterile, single use. It is used to lead the urine in the patient's bladder out of the body through the urethra and into the urine collection container. All Class II products under 14-05-03 | II |  | 615
616 | 14-05-04 | rectal tube (anal tube) | Usually consists of conduits and fittings. It is generally made of medical polyvinyl chloride plastic. The tip of the catheter is divided into two types, holed and non-holed, with one or more side holes on both sides. Supplied sterile. For bowel cleansing (rinsing, emptying or priming) or for venting. | II |  | 616
617 | 14-05-05 | Ureteral stent | Usually a tubular structure made of silicone rubber, polyurethane, or other polymer, with looped bends at one or both ends. Placed between the kidney and bladder. It is used to treat ureteral blockage and stricture by supporting and draining the human ureter. The exemption does not include products using new materials, new mechanisms of action, or new functions. The residence time in the body is less than 30 days. | II |  | 617
618 | 14-05-05 | Ureteral stent | Usually a tubular structure made of silicone rubber, polyurethane, or other polymer, with looped bends at one or both ends. Placed between the kidney and bladder. It is used to treat ureteral blockage and stricture by supporting and draining the human ureter. The exemption does not include products using new materials, new mechanisms of action, or new functions. The residence time in the body is greater than or equal to 30 days. | III |  | 618
619 | 14-05-06 | Disposable biliary drainage tube | The disposable biliary drainage tube can be made of suitable materials such as silica gel, and can be composed of a spiral drainage end, a tube body and a joint; it can be divided into several models and specifications according to different materials, designs and technical parameters; the product is in sterile form Provided; for bile drainage after endoscopic biliary stone surgery or cholecystectomy. | II |  | 619
620 | 14-05-06 | Disposable surface drainage ball/tube | Disposable body surface drainage ball/tube can be made of silicone rubber, latex and other suitable materials, and can be designed into various shapes. The tube body has a side hole, which can be connected with negative pressure ball. It is divided into several models and specifications according to different parts; it is provided in a sterile form; it is used for drainage of body surface wounds or incisions. | II |  | 620
621 | 14-05-06 | Disposable ureteral drainage tube | Disposable ureteral drainage tubes can be made of suitable materials such as polyurethane, usually straight, with side holes in the tube body; they can be divided into several specifications according to tube diameters, lengths, etc., and provided in sterile form; for ureteral support and drainage use. | II |  | 621
622 | 14-05-06 | Nasopancreatic/biliary drainage tube | It usually consists of drainage tube, connecting tube and nasal conversion tube and other components. Made of polymer material. The product is sterilized and disposable. The nasopancreatic/biliary drainage tube is designed to be inserted directly into the pancreatic duct/biliary tract through the endoscope to drain pancreatic juice/biliary. | II |  | 622
623 | 14-00-00 | Disposable chest drainage tube and puncture set | The product uses medical silicone rubber / polyurethane as raw materials and is used for drainage of pleural effusion. The puncture kit can be composed of a dilator, a puncture needle, a guide wire, an introducer sheath, and accessories to establish a subcutaneous artificial channel and assist in the placement of a drainage tube. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II | Revised classification codes, product names and product descriptions. | 623
624 | 14-00-00 | Disposable abdominal drainage tube and puncture set | The product is made of medical silicone rubber / polyurethane, etc., and is used for the drainage of peritoneal effusion. The puncture kit can be composed of a dilator, a puncture needle, a guide wire, an introducer sheath, and accessories to establish a subcutaneous artificial channel and assist in the placement of a drainage tube. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II | Revised classification codes, product names and product descriptions. | 624
625 | 16-07-12 | lacrimal duct / lacrimal drainage tube | It usually consists of a silicone tube, a silicone spear, a dilator, and a perforated plug. The lacrimal duct is made of silicone rubber and is equipped with accessories such as special surgical traction hooks or guide wires. It is used for lacrimal duct support and implantation treatment after probing of lacrimal duct obstruction, dacryocystitis nasodacrimal sac anastomosis, and lacrimal canaliculus rupture anastomosis. It is equivalent to products that have been approved for domestic registration. The product composition materials are mature, except for new structural design, new mechanism of action, and drug-device combination products. | III | Revised classification codes. | 625
626 | 14-05-07 | ureteral dilation catheter | The product can be composed of balloons, catheters, catheter adapters and other components. The product is sterilized and disposable. Transluminal dilation for ureteral strictures, or for ureteral dilation prior to ureteroscopy or stone procedures. | II |  | 626
627 | 14-05-07 | sinus balloon catheter | It usually consists of a balloon catheter, an insertion probe, and an irrigation tube. This product is sterilized and disposable. A device used to dilate sinus ostia and sinus cavities during diagnostic and therapeutic procedures; it can also be used to irrigate within target sinuses to complete therapeutic procedures and facilitate diagnostic procedures. | II |  | 627
628 | 14-05-07 | Percutaneous transhepatic biliary drainage access kit | For percutaneous transhepatic biliary drainage, it can be composed of a dilator, puncture needle, guide wire, catheter sheath and accessories, etc., to establish a subcutaneous artificial channel and assist the placement of the drainage tube. | II | Add | 628 | NEW
629 | 14-05-08 | Contrast Catheter | Usually consists of a catheter and a catheter adapter. Supplied sterile, single use. With the cooperation of the guiding device, the catheter is inserted into a certain part of the body (non-vascular tissue) to inject contrast agent to the targeted site. | II |  | 629
630 | 14-05-09 | pressure catheter | It is usually composed of a catheter, a catheter seat, etc., and is often inserted into a certain part of the body with the cooperation of a guiding device for measuring physiological pressure. Supplied sterile, single-use for measuring physiological pressure at a target site (nonvascular tissue). | II |  | 630
631 | 14-06-01 | Extracranial drainage collection device | It usually consists of accessories such as extracorporeal drainage tubing, droppers, and collection containers. For cerebrospinal fluid puncture external drainage, the collection device should also have an intracranial pressure scale (the external drainage collection device for craniocerebral puncture other than external ventricular drainage may not include the intracranial pressure scale, which can be selected according to actual needs). Supplied sterile, single use. If intracranial pressure can be measured, clinical verification data on the accuracy of intracranial pressure measurement should be provided. The exemption does not include devices inserted into a patient's ventricles, devices that deliver drugs to the ventricles, and pharmaceutical ingredients. | II | Revised product description. | 631
632 | 14-06-02 | chest drainage device | It usually consists of at least a chest drainage tube and an effusion cavity, which can be divided into water-sealed type, dry-sealed valve type, etc., and some have suction ports to achieve suction and drainage. Drainage devices with autologous blood reinfusion are not included. From the production process, it can be divided into injection molding and blow molding. Chest drainage catheters inserted into the patient's chest are not included. Supplied sterile, single use. Matched with chest drainage catheter, it is used for patients who need closed drainage after pneumothorax, pleural effusion and surgery. | II |  | 632
633 | 14-06-03 | Negative pressure diverter and components | It usually consists of a body, a spring, a regulator group, a connecting joint, a connecting pipe, a stop clip, a protective cap, a drainage bag, a negative pressure pump or a manual negative pressure source (negative pressure ball). Supplied sterile. Drainage catheters inserted into the body are not included. When used for clinical negative pressure drainage, it is connected with the drainage catheter inserted into the body to act as a negative pressure conduction medium and/or guide and collect drainage fluid. | II |  | 633
634 | 14-06-03 | Disposable negative pressure drainage device | It usually consists of a body, a spring, a regulator group, a connecting joint, a connecting pipe, a stop clip, a protective cap, a drainage bag, a negative pressure pump or a manual negative pressure source (negative pressure ball). Supplied sterile. Drainage catheters inserted into the body are not included. When used for clinical negative pressure drainage, it is connected to the drainage catheter inserted into the body to act as a negative pressure conduction medium and/or guide and collect drainage fluid. | II |  | 634
635 | 14-06-04 | Vacuum Negative Pressure Machine | It usually consists of vacuum pumps, pipelines, joints, etc. The vacuum pump is used to generate negative pressure in the system pipeline to achieve continuous or intermittent wound drainage or provide a negative pressure environment to assist wound closure. Used to remove secretions and necrotic tissue from cavities or wounds. | II |  | 635
636 | 14-06-06 | Negative pressure drainage sealing membrane | Usually, a polyurethane film is evenly coated with a synthetic adhesive substance to form the adhesive surface of the closed film, and the adhesive surface is covered with a protective layer. When performing negative pressure drainage on the wound surface, the negative pressure drainage sponge is placed on the wound surface, and the negative pressure drainage sealing film is attached to form a closed negative pressure drainage cavity under the support of the sponge and the sealing effect of the sealing film. Under the action of the external suction source, the drainage of the wound surface is realized. Supplied sterile, single use. | II | Add | 636 | NEW
637 | 14-06-07 | Medical electric suction device | It usually consists of a negative pressure indicator, a filter, a collection container and an anti-overflow device, and is driven by mains power or batteries. Negative pressure can be provided by rotary vane pumps, piston pumps, diaphragm pumps and electromagnetic pumps, and the gas, liquid and/or solid in the human body can be drawn out by negative pressure. It can be used in medical places, at home, in the field or during transportation. It can absorb secretions in the human respiratory tract; gas and liquid in the stomach and chest cavity; exudates from body surface wounds and wounds; amniotic fluid; blood (washing fluid), exudates, and waste fluids during and after operations. Does not include suction devices used only for gynecological abortion, dental suction, ophthalmic suction or endoscopic surgery. | II |  | 637
638 | 14-06-07 | Medical negative pressure suction device | The medical negative pressure suction device is mainly composed of a negative pressure system, a pressure display, a main control circuit, and a liquid storage container. It usually does not include a suction tube. It can be divided into several models according to the principle of negative pressure generation, design, technical parameters, intended use, and additional auxiliary functions. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0636.1-2008 Medical Suction Equipment Part 1: Safety Requirements for Electric Suction Equipment. | II |  | 638
639 | 14-06-08 | Suction device powered by negative pressure source or pressure source | Usually consists of negative pressure indicator, filter, collection container and overflow prevention device, and may contain electrical control components. It is connected to a negative pressure source, driven by a central suction system negative pressure pipeline or compressed gas, and can also be driven by its own gas cylinder. It is also necessary to be equipped with a negative pressure regulator to connect with the negative pressure pipeline. Gases, liquids and/or solids are drawn out of the human body using negative pressure. It can be used in medical places, at home, in the field or during transportation. It can absorb secretions in the human respiratory tract; gas and liquid in the stomach and chest cavity; exudates from body surface wounds and wounds; amniotic fluid; blood (washing fluid), exudates, and waste fluids during and after operations. Does not include suction devices used only for gynecological abortion, dental suction, ophthalmic suction or endoscopic surgery. | II |  | 639
640 | 14-06-09 | Manually driven suction devices for medical use | Usually consists of negative pressure indicator, filter, collection container and overflow prevention device. Including pedal, manual or both, also can be used in combination with electrical equipment, using negative pressure to extract gases, liquids and/or solids from the human body. It can be used in medical places, at home, in the field or during transportation. It can absorb secretions in the human respiratory tract; gas and liquid in the stomach and chest cavity; exudates from body surface wounds and wounds; amniotic fluid; blood (washing fluid), exudates, and waste fluids during and after operations. Does not include suction devices used only for gynecological abortion, dental suction, ophthalmic suction or endoscopic surgery. The exemption does not include products using new materials, new mechanisms of action, or new functions. | II |  | 640
641 | 14-06-10 | Medical vacuum negative pressure unit | Generally composed of vacuum pumps, vacuum tanks, pipelines, valves, control cabinets and vacuum instruments, they are used in medical gas pipeline systems to provide negative pressure suction sources. | II | Merged with the original 632 to revise the product name and product description. | 641
642 | 14-06-11 | Extracorporeal drainage, suction tube | Usually consists of hoses and connectors. It can be connected between the drainage catheter and the drainage device to form a closed drainage system. Do not directly contact the human body. Supplied sterile, single use. After matching with suitable equipment, it is used for drainage and suction of blood and waste liquid during and after surgery. | II |  | 642
643 | 14-06-12 | Sterile drainage bag (container), sterile urine collection bag (container) | Usually a bag-like collection container. It is connected with the drainage catheter through the extracorporeal tubing to form a closed drainage system. Supplied sterile, single use. | II |  | 643
644 | 14-06-12 | Sterile drainage bag/urine bag | This product can be composed of collection bag, input tube, discharge valve, anti-reflux sheet, drainage tube connector and other components. It is connected with the drainage catheter through the extracorporeal tubing to form a closed drainage system. It can have anti-reflux function, and the waste liquid will not flow back into the human body. The raw materials are generally polymer materials such as medical PVC plastic and polypropylene. Sterilized product, single use. Used to collect secretions and body fluids during patient treatment. | II |  | 644
645 | 14-06-12 | Drainage bags (containers)/collection bags (containers), stool management devices | Usually a bag-like collection container. It is connected with the drainage catheter through the extracorporeal tubing to form a closed drainage system. Supplied sterile, single use. It is used for the collection of one-time drainage fluid (blood, gastric juice, etc.), secretions (sputum, flushing fluid, etc.) and human excrement of patients in clinical departments of hospitals and during or after surgery. | II |  | 645
646 | 14-06-13 | Sterile balloon inflation pressure pump | It usually consists of injection system, pressure gauge, connecting pipeline and so on. Supplied sterile. | II |  | 646
647 | 14-07-01 | Aseptic irrigator | There are different structures according to different intended uses. Before use, it can be filled with irrigating liquid, or connected with related irrigating equipment or appliances to form an irrigating system, which can be used for rinsing the patient's rinsing site. Supplied sterile. | II |  | 647
648 | 14-07-01 | Electric gastric lavage machine | Electric gastric lavage machine is mainly composed of pressure pump, control unit, flushing and suction conversion device, filter bottle, hose, flushing liquid and sewage container, etc., and can be divided into several types according to pressure pump, control mode, technical parameters, additional auxiliary functions, etc. model. It is used by medical units for rescue poisoning, food poisoning and gastric lavage before operation. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1105-2008 Electric Gastric Lavage Machine. | II |  | 648
649 | 14-07-01 | Nasal irrigation device | Nasal irrigation device is made of plastic, consisting of bottle body, bottle cap and outlet pipe (with flow control valve), excluding irrigation liquid; it can be divided into several models and Specifications; can be provided in sterile form; for nasal irrigation. | II |  | 649
650 | 14-07-01 | Rinse Warming Kit | It is used for heating the irrigation fluid in the operating room. | II |  | 650
651 | 14-07-01 | Electric washer | It usually consists of a host, a heating module, and a power module. Does not contain rinse solution. It is used for flushing the natural orifice (excluding the vagina), flushing tissues during and after surgery, and can also warm the flushing solution. | II |  | 651
652 | 14-07-02 | sterile enema | It is usually composed of soft bags, clips, connecting pipes, joints and other parts, and is generally made of glass, polymer and other materials. The soft bag is used for enema solution, and the connector can be connected with the rectal catheter. Contents not included. Supplied sterile. | II |  | 652
653 | 14-07-03 | sterile applicator | It usually consists of a nozzle, a cavity and an inhaler connection port. Made of polymer and metal materials. Devices for inhalation administration of aerosol medicaments may also be included. Does not contain drugs. | II |  | 653
654 | 14-09-01 | Surgical Woven Cloth Dressing | Corresponding to Class II products (except hemostatic gauze) in 14-09-01 of the "Medical Device Classification Catalog". Product performance indicators comply with applicable parts of relevant standards (including amendments). Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 654
655 | 14-09-01 | Developable gauze | Absorbent cotton gauze with X-ray detectable barium thread (sheet), sterile product for single use. As a surgical dressing, it has the function of X-ray detection. | II |  | 655
656 | 14-09-02 | Surgical Nonwoven Dressings | Corresponding to Class II products in 14-09-02 of the "Medical Device Classification Catalog". It is usually a dressing made of non-woven compresses, which is used for clinical wound protection and moisture absorption. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 656
657 | 14-09-03 | Surgical Sponge Dressing | It is usually a spongy dressing made of polymer materials. Supplied sterile, single use. It is used to absorb the exudate in the body during the operation, and to support the organs and tissues during the operation. It is also used to stop bleeding by packing and compressing the cavity (such as the nasal cavity). | II |  | 657
658 | 14-10-01 | Wound dressing | Corresponding to the wound dressings exemplified in 14-10-01 in the "Medical Device Classification Catalog". Usually consists of a rubberized base, an absorbent pad and a peelable protective layer. Absorbent pads are generally made of materials that can absorb exudate, such as cotton fibers and non-woven fabrics. Absorbent pads can be used alone or secured with bandages or tape, etc. Contains ingredients that cannot be absorbed by the body. Supplied sterile, single use. For the care of non-chronic wounds on the body surface. Exemption does not include: (1) Indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostatic, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin (2) products claiming to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, and patients with wound sepsis; (3) products containing active Products with ingredients: such as pharmaceuticals/pharmaceutical active ingredients, biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 658
659 | 14-10-01 | stickers | Corresponding to Class II products in 14-10-01 of the "Medical Device Classification Catalog". Composed of backing and adhesive, etc., sterile products for one-time use. For the care of the puncture site of a puncture device (eg catheter) and fixation of the puncture device or for secondary fixation of dressings. The exemption does not include: new structural design, new mechanism of action, new intended use, containing active ingredients, etc. | II |  | 659
660 | 14-10-02 | Bandage | Corresponding to Class II products in 14-10-02 of the "Medical Device Classification Catalog". Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 660
661 | 14-10-04 | gel dressing | Usually a gel-forming or amorphous gel dressing consisting of a gel-forming substance and water, which may contain buffer salts. The ingredients contained in it cannot be absorbed by the human body and have no pharmacological effect. Supplied sterile. It is used to absorb wound exudate or drain water to the wound, and to cover non-chronic wounds such as sutured wounds after surgery. | II |  | 661
662 | 14-10-05 | hydrocolloid dressing | Corresponding to Class II hydrocolloid dressings, hydrocolloid applicators, and hydrocolloid applicators listed in 14-10-05 of the "Medical Device Classification Catalog". Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 662
663 | 14-10-05 | hydrocolloid dressing | Corresponding to the example of Class III hydrocolloid dressing or hydrocolloid application in 14-10-05 of the "Medical Device Classification Catalog". Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 663
664 | 14-10-06 | alginate dressing | Composed of alginate fibers, the intended use is limited to class II alginate dressings that cover non-chronic wounds on the body surface and absorb wound exudate. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 664
665 | 14-10-06 | Hydrophilic fiber dressing | Mainly composed of carboxymethylcellulose sodium or ethylsulfonate cellulose, the intended use is limited to Class II hydrophilic fiber dressings covering non-chronic wounds on the body surface and absorbing exudate. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 665
666 | 14-10-06 | alginate dressing | Composed of alginate fibers, the intended use is limited to class III alginate dressings and alginate filling strips that cover wounds and absorb wound exudates. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 666
667 | 14-10-06 | Hydrophilic fiber dressing | Mainly composed of carboxymethylcellulose sodium or ethylsulfonate cellulose, the intended use is limited to Class III hydrophilic fiber dressings that cover wounds and absorb exudates. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 667
668 | 14-10-07 | polyurethane foam dressing | It is mainly composed of polyurethane foam, and its intended use is limited to Class II polyurethane foam dressings used to cover non-chronic wounds on the body surface and absorb wound exudate. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 668
669 | 14-10-07 | polyurethane foam dressing | It is mainly composed of polyurethane foam, and its intended use is limited to Class III polyurethane foam dressings used to cover wounds and absorb wound exudates. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 669
670 | 14-10-08 | Liquid and cream dressings | Usually a solution or ointment (excluding gels) containing ingredients that do not have pharmacological effects and cannot be absorbed by the body. Supplied sterile. By forming a protective layer on the wound surface, it acts as a physical barrier. It is used for the care of small wounds, abrasions, cuts and other non-chronic wounds and surrounding skin. | II |  | 670
671 | 14-10-09 | Vaseline/paraffin gauze (oil gauze) | Made of medical absorbent gauze, petroleum jelly, paraffin oil, intended use is limited to petroleum jelly/paraffin gauze or gauze strips for wound drainage, covering, packing. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 671
672 | 14-10-09 | Hydrocolloid oil yarn | Mainly composed of petrolatum, hydrocolloid particles (sodium carboxymethylcellulose) and reticulated polyester fibers. Used for wound drainage, covering and stuffing. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | III |  | 672
673 | 14-11-01 | First aid hemostatic bandage, tape application bandage | Corresponding to the Class II bandage exemplified in 14-11-01 in the "Catalogue of Medical Device Classification". Rolls, tubes, and triangular materials that are usually processed by textiles. Some have properties such as elastic or self-adhesive. It is used for the care of non-chronic wounds on the body surface, or to provide binding force on wound dressings or limbs, so as to bandage and fix them. Excluding new structural design, new mechanism of action, new intended use, containing active ingredients, etc. | II |  | 673
674 | 14-12 | Ostomy bag (including chassis) | This product includes enterostomy bags (open or closed) and urostomy bags. The enterostomy bag is mainly composed of an outer bag, a chassis and a filter sheet. The urostomy bag is mainly composed of a multi-chamber outer bag (with anti-reflux structure), a chassis and a drainage valve. Supplied sterile for single use. For the management of ostomy discharge. | II |  | 674
675 | 14-13 | Disposable sterile clean clothes | Disposable sterile clean clothes are generally made of non-woven fabric or film-coated non-woven composite material with certain protective properties. It is an elastic closure, and the waist closure can be elastic closure, drawstring closure or buckle; it can be divided into gown type, one-piece type, split type, etc.; provided aseptically, for personnel entering the surgical clean area to reduce the source of infection caused by personnel The resulting air microbial contamination in the clean area prevents postoperative wound infection in patients. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0506 Surgical drapes, surgical gowns and clean clothes for patients, medical staff and instruments. | II |  | 675
676 | 14-13-01 | Disposable sterile surgical drapes | Disposable sterile surgical drapes are generally made of non-woven fabric or film-coated non-woven composite material with certain protective properties, and are provided in a sterile form to cover the surface of the patient's body to reduce the source of infection in non-surgical parts such as the patient's skin Move to the surgical site to prevent postoperative wound infection of patients, and also have the function of collecting surgical dirt due to the design. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0506 Surgical drapes, surgical gowns and clean clothes for patients, medical staff and instruments. | II |  | 676
677 | 14-13-01 | surgical drapes | Surgical drapes can be made of non-woven or film-coated non-woven composite materials with certain protective properties. They can be divided into several models and specifications according to different design types and sizes. They are usually provided in sterile form for covering the surgical skin of patients. Around the wound, reduce the infection source of the patient's surgical wound skin to the surgical wound site, and prevent the patient from postoperative wound infection. | II |  | 677
678 | 14-13-01 | Sterile protective cover | It is usually composed of an open end and a sleeve body, generally made of polyethylene and other polymer materials, and provided aseptically. It is used to cover the instrument table, operating table, display screen, etc. in the surgical operating room, or to isolate the instrument (such as gastroscope) from the patient, so as to prevent the doctor during the operation from touching the above-mentioned parts, and then contacting the patient's wound during the operation to cause infection. | II |  | 678
679 | 14-13-02 | Disposable sterile surgical film | Disposable sterile surgical membranes are generally made of natural rubber latex (medical pressure-sensitive adhesive may not be applied) or polyurethane or polyethylene substrates coated with hypoallergenic pressure-sensitive adhesives, with release paper on top, and are provided in a sterile form ; It is used to simplify the preoperative skin care operation and prevent contact and migration surgical wound infection. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0852 Disposable sterile surgical film. | II |  | 679
680 | 14-13-03 | Disposable sterile surgical gloves | Disposable sterilized surgical gloves are made of natural rubber latex or synthetic rubber and other materials, provided in a sterile form, and used for protection during clinical surgical operations. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 7543 Disposable sterile rubber surgical gloves. | II |  | 680
681 | 14-13-04 | Disposable medical surgical mask | Disposable medical surgical masks are usually made of non-woven or non-woven composite materials, which can have a three-layer structure, have a plastic nose clip, and the mask belt can be elastic or non-elastic, with small airflow resistance and synthetic blood Blocking, filtering particles and bacteria, flame retardant and other characteristics; for clinical medical personnel to wear during invasive operations, covering the user's mouth, nose and jaw, providing physical protection to prevent direct penetration of pathogenic microorganisms, body fluids, particles, etc. barrier. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0469 Medical surgical masks. The exemption does not include products with new features such as anti-virus. | II | Revised product description. | 681
682 | 14-13-05 | reusable surgical gown | Operating theater garments, usually consisting of a base material and a water-blocking layer. The base material is generally made of non-woven fabric or woven fabric, and the water-blocking layer is a water-blocking material. Supplied non-sterile for reusable use. According to the barrier ability of key areas, surgical gowns are divided into two types: standard type and high-performance type. | II |  | 682
683 | 14-13-05 | Disposable surgical gown | Disposable surgical gowns are generally made of non-woven fabrics or film-coated non-woven composite materials with certain protective properties; the surgical gowns can be composed of collars, body, sleeves, or jackets, trousers, cuffs, Ankle openings are elastic closures, waist closures can be elastic closures, drawstring closures or buckles; front collars, elbows, etc. can be designed to enhance protection, provided in a sterile form, for surgical personnel to wear, and reduce the spread of infection sources to patients' surgical wounds To prevent postoperative wound infection; surgical gowns that prevent liquid penetration can also reduce the risk of transmission of infection sources carried in blood or body fluids to surgical personnel. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0506 Surgical drapes, surgical gowns and clean clothes for patients, medical staff and instruments. | II |  | 683
684 | 14-14-01 | Medical protective mask | Medical protective masks are made of suitable materials with small airflow resistance, synthetic blood barrier, specific strength, filtration efficiency, surface moisture resistance and flame retardancy. They are self-priming and consist of a mask body (including nose clip) and a mask belt The composition can be divided into 3 levels according to the different filtration efficiency. It is used for filtering particulate matter in the air, blocking droplets, blood, body fluids, secretions, etc. in the medical working environment. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 19083 Technical requirements for medical protective masks. The exemption does not include products with new features such as anti-virus. | II |  | 684
685 | 14-14-02 | Medical Disposable Protective Clothing | Medical disposable protective clothing is made of suitable materials with specific liquid barrier function, specific strength, filtration efficiency, flame retardant performance, antistatic and static attenuation functions. It consists of a hooded jacket and trousers. The cuffs and ankles are elastic closures. The face closure and waist closure can adopt elastic closures, drawstring closures or buckles; they can be divided into two types: one-piece type or split type, for clinical medical staff to come into contact with potentially infectious patients' blood, body fluids, and secretions during work. objects, etc. to provide barrier and protection. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 19082 Technical Requirements for Medical Disposable Protective Clothing. | II |  | 685
686 | 14-14-04 | hand protection | It is usually made of materials such as polyvinyl chloride, rubber or stainless steel. Have sufficient strength and barrier properties. Supplied sterile, single use. It is used to check or palpate the patient's condition by wearing it on the doctor's hand or finger, or to prevent the doctor's hand from being bitten. | II |  | 686
687 | 14-15-01 | Baby phototherapy goggles | Baby phototherapy protective goggles can be composed of goggles and goggles fixed; the goggles can be made of high-quality non-woven fabrics/blue-black physical composite fabrics, etc., and the fixed parts can be made of materials such as elastic bands or soft Velcro; for babies when they are treated with blue light For eye protection. | II |  | 687
688 | 14-15-02 | Eye patch | Corresponding to Class II products in 14-15-02 of the "Medical Device Classification Catalog". It can be composed of base layer (non-woven fabric, etc.), medical polyvinyl alcohol hydrogel and release paper; it can be divided into several models and specifications according to different shapes and sizes; Provide a relatively closed humid environment to prevent exposure keratitis. The exemption does not include products containing active ingredients such as pharmaceuticals/pharmaceutical active ingredients, biological products/biologically active ingredients, silver, disinfectants, etc. | II |  | 688
689 | 14-15-03 | nasal hemostat | It usually consists of an inflation valve, an airway tube, and an air bag, provided sterile and for one-time use. The inflated air bag compresses and suppresses nasal bleeding. | II | Add | 689 | NEW
690 | 14-15-04 | seawater nasal spray | It usually consists of a sprayer and a spray liquid. The spray liquid is composed of purified water and sea water. It is used for nasal cleaning of patients with acute and chronic rhinitis, allergic rhinitis, nasal polyps, sinusitis and other nasal diseases. It is also used for nasal cleaning after rhinitis surgery and chemotherapy. | II |  | 690
691 | 14-15-05 | sterile pad | It is usually made by laminating or sewing non-woven fabric and plastic film. Supplied sterile, single use. Hygienic care articles for hospital beds or examination beds | II |  | 691
692 | 14-16 | Colored grid for marking skin injection sites | Consists of tattoo paper and vinyl acetate polymer film printed with cosmetic grade ink on it. The ink consists of food grade pigments, ethyl cellulose and castor oil. The mechanism of action and intended use of the product are: to transfer multiple color grids composed of colored dots formed by the ink matrix on intact skin, and to provide doctors with a reference for injection intervals (grids at intervals of 1 square centimeter). | II |  | 692
693 | 14-16-01 | Clean screen | It usually consists of a box, a fan unit, a filter, an electrical controller, etc. It is used to purify the local ambient air of the patient's surgical site during the operation to prevent infection. | II |  | 693
694 | 14-16-03 | Prefilled Catheter Irrigator | It consists of 0.9% sodium chloride injection, jacket, core rod, and rubber piston. For gaps between different drug treatments, seal and flush the end of the catheter. | III | Add | 694 | NEW
695 | 14-16-05 | Nasal strips (without medicine) | Nasal strips (without medicine) fix the elastic plastic sheet on the bridge of the nose with superglue and base cloth to dilate nostrils and relieve nasal congestion. | II |  | 695
696 | 14-16-05 | upper airway dilator | By expanding the nasal cavity, upper and lower jaw or correcting it, it is only used for adjuvant treatment of claimed obstructive apnea, for dilating nostrils and relieving nasal congestion. The contact time with the human body is less than 30 days. | II |  | 696
697 | 14-16-05 | nasal lubricant | It usually consists of polysorbate 80, glycerin, sodium chloride, sodium edetate, potassium sorbate, and purified water. Reduces upper airway resistance by lubricating and keeping mucous membranes moist to improve respiratory obstruction. | II |  | 697
698 | 14-16-06 | sterile bite | Auxiliary instruments for patients to open their mouths during surgery or examinations are generally made of polymer materials such as polyethylene. Supplied sterile. Used to maintain the patient's mouth during transoral procedures or examinations, to prevent unintended occlusion, or to facilitate insertion and securing of endotracheal tubes. | II |  | 698
699 | 14-16-08 | Body surface device fixation device | It is usually a device that can provide fixation specifically for the use of a certain type or type of equipment. contact with the wound. Supplied sterile, single use. Used to fix medical devices during use | II |  | 699
700 | 14-16-09 | Medical sterile liquid paraffin nonwoven fabric/medical paraffin cotton ball | Medical sterile liquid paraffin non-woven fabric is made of non-woven fabric/cotton ball and liquid paraffin. It can be divided into several models and specifications according to formula, material, size, etc., and can be provided in a sterile form; for patients to check the surface of the site and equipment For surface lubrication. Exemption does not include: (1) Products whose indications claim to promote epithelialization, guide tissue regeneration, promote wound healing, relieve pain, antibacterial, anti-infection, anti-virus, hemostasis, dissolve necrotic tissue, reduce scars, prevent adhesions, and act as artificial skin/skin substitutes; (2) Products that claim to be used for internal wounds, third-degree burns, infected wounds, wounds with a lot of necrotic tissue, patients with wound sepsis, etc.; (3) Products containing active ingredients: such as drugs/pharmaceutical active ingredients , Biological products/biologically active ingredients, silver, disinfectants, etc.; (4) Other new products, such as products with new materials, new mechanisms of action, and new functions. | II |  | 700
701 | 14-16-09 | lubricant | Lubricants are usually composed of glycerin, xanthan gum and other ingredients, and do not contain pharmaceutical ingredients. It is clinically used for lubrication when instruments enter the natural orifice of the human body. | II |  | 701
702 | 14-16-10 | Cotton swabs/balls/tablets for disinfection | Cotton swabs/balls/tablets containing alcohol, tincture of iodine or povidone iodine for disinfection of intact skin before injections and infusions. | II |  | 702
703 | 14-16-10 | Sterile cotton swab/ball/tablet | Usually includes absorbent material. For ease of use, some products have components for hand-held use. Does not contain sanitizers. Supplied sterile, single use. Used for cleaning and treating skin and wounds. | II |  | 703
704 | 14-16-13 | Aseptic tube machine | It usually consists of a host, a controller, a welding device, and a display screen. For aseptically splicing two medical tubing together. | II |  | 704
705 | 15-00-00 | ENT examination and treatment table | The otolaryngology inspection and treatment table usually consists of an operation inspection and treatment table, a host and accessories (spray gun, suction gun, blow gun, defogging, spotlight and bracket, tray, cotton ball cylinder, monitor, cold light source, CCD camera, ultraviolet disinfection Lamp, film viewing lamp, etc.), which can be divided into several models and specifications according to structural composition, function and technical parameters; it can be matched with ENT endoscope for inspection, diagnosis and treatment in ENT department. | II |  | 705
706 | 15-01-01 | electric operating table | The electric operating table is mainly composed of a table, a pedestal, an electrical control system, and accessories. It is divided into several models and specifications according to the driving mode, technical parameters, auxiliary functions, and accessories. It can be applied to the head, neck, chest, and It is used to carry patients in abdominal cavity, perineum, extremities and other surgeries, gynecology, urology, ENT, orthopedics, and provide suitable surgical positions. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1106-2008 Electric Operating Table, YY 0570-2005 Medical Electrical Equipment Part 2: Special Requirements for Operating Table Safety. | II |  | 706
707 | 15-01-01 | Electro-hydraulic operating table | The electro-hydraulic operating table is mainly composed of table top, pedestal, electro-hydraulic control system (usually the table top lifting is controlled by hydraulic pressure), supporting parts and other parts. It is divided into several models and Specifications, can be respectively applied to head, neck, chest, abdominal cavity, perineum, limbs and other surgery, gynecology, urology, ENT, orthopedics and other surgical load patients, providing suitable surgical positions. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1106-2008 Electric Operating Table, YY 0570-2005 Medical Electrical Equipment Part 2: Special Requirements for Operating Table Safety. | II |  | 707
708 | 15-01-02 | Manual operating table (hydraulic) | It usually consists of bed body (including support part, transmission part and control part) and accessories. The supporting part includes table top, back board, hip board, board, etc., which can be adjusted. The lifting form is hydraulic lifting type, and the body position adjustment is all manpower operation. | II |  | 708
709 | 15-01-02 | hydraulic operating table | The hydraulic operating table is mainly composed of a table, a pedestal, a hydraulic control system, and accessories. It is divided into several models and specifications according to the driving parts, technical parameters, auxiliary functions, and accessories. It can be applied to the head, neck, chest, and It is used to carry patients in abdominal cavity, perineum, extremities and other surgeries, gynecology, urology, ENT, orthopedics, and provide suitable surgical positions. | II |  | 709
710 | 15-02-01 | Electric diagnosis table and treatment chair | The diagnosis and treatment (examination) table is usually composed of bed frame, bed surface, pillow and so on. The diagnosis and treatment (examination) chair usually consists of a base, a backboard, a seat and a footrest. Active product. | II |  | 710
711 | 15-02-01 | electric examination chair | The electric examination chair is generally composed of a base, a backboard, a seat board, a footrest and a foot control switch. For supporting the patient during examination and treatment. | II |  | 711
712 | 15-03-01 | electric hospital bed | The electric hospital bed is a network power supply or an internal power supply or a combined electric drive device. Stand part (lifting mechanism, central control mechanism, casters, etc.), electric control part (lifting motor, backboard motor, leg board motor, controller, hand control board), infusion stand and other accessories, which are divided into different types according to different functional configurations. Several models are used for carrying and nursing patients in various clinical wards. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0571-2013 Medical electrical equipment Part 2: Special requirements for the safety of hospital electric beds. | II |  | 712
713 | 15-04-01 | Electric Patient Positioning Aids | It usually consists of a host, universal connection accessories, bed clamps, etc. Active product. It is used for fixation and intraoperative adjustment of knee, shoulder, hip and small joints during surgery. | II |  | 713
714 | 15-05-01 | Negative pressure isolation cabin for transporting infectious disease personnel | It usually consists of a negative pressure cabin (movable hatch, bilge) and an exhaust filter device. Usually it is inflated to expand during work, and can be folded when deflated during storage and transportation. It is used for the safe isolation and transfer of infectious disease patients. It can be used in conjunction with stretchers and vital sign detection instruments to monitor infectious disease patients during transportation. | II |  | 714 | NEW
715 | 15-05-02 | Electric carts, stretchers and other equipment | It usually consists of a bed surface (trolley panel, stretcher surface) and a support frame. Infusion stands, guardrails, directional wheel pedals, casters, brakes, hand-cranking mechanisms, etc. can be added. For single or double cars. Active product. | II |  | 715
716 | 15-05-02 | transfer bed | The transfer bed can be composed of bed frame, bed board, casters, mattress, side rails, braking system, hydraulic lifting and control system, push handle, and accessories (such as: restraint belt, infusion stand). It is suitable for the transfer or emergency rescue of patients in hospitals. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0571-2005. | II |  | 716
717 | 15-05-02 | Electric Stretcher | This product usually consists of a bed surface (trolley panel, stretcher surface) and a support frame. Infusion stands, guardrails, directional wheel pedals, casters, brakes, hand-cranking mechanisms, etc. can be added. For single or double cars. Active product. It is used in medical institutions to transport and move patients. | II |  | 717
718 | 15-06-01 | Motorized Anti-Pressure Ulcer (Decubitus) Mat | It usually consists of an air mattress, an airway (connecting tube), an air pump, etc. A pressure sore (decubitus) cushion consists of several air chambers. After the air cushion is inflated by the air pump, the air chamber maintains a certain air pressure, and the soft cushion formed can increase the contact area between the patient's body and the cushion and reduce the local pressure on the body. | II |  | 718
719 | 15-06-01 | Medical anti-decubitus air cushion | The medical anti-decubitus air cushion is mainly composed of airbags, electric air pumps, controllers, massage motors, and bed covers; it can be divided into several models according to different materials, designs, technical parameters, and inflation modes; it is used for preventing and treating bedsores in bedridden patients. | II |  | 719
720 | 16-00-00 | Passive ophthalmic surgical instruments | Passive short-term (≤24 hours) use devices, marketed for many years, non-innovative functional design, not made of new materials, no coating, no degradation, absorption function, no drugs, animal-derived components, generally for one-time use Sterilized products, used to assist in the completion of ophthalmic operations and inspections. | II |  | 720
721 | 16-01-01 | Disposable ophthalmic surgical knife | Disposable ophthalmic surgery knife consists of handle (polycarbonate) and blade (stainless steel), which can be divided into several models and specifications according to material, design, technical parameters, intended use, etc.; provided in sterile form; for display For micro eye surgery. | II |  | 721
722 | 16-02-01 | Disposable Ophthalmic Piercing System | The disposable ophthalmic puncture system may include a puncture knife, a puncture sleeve, a handle, a protective cap, a self-closing valve, and the tip of an infusion tube, which may be made of stainless steel, polycarbonate, polypropylene, silicone, etc.; The material, design, and technical parameters are divided into several models and specifications; the product is provided in a sterile form; it is used for establishing instrument channels in the flat part of the sclera in ophthalmic operations such as vitrectomy. | II |  | 722
723 | 16-03-01 | Optometry Equipment and Appliances | All products under classification code 16-03-01 are exempt from clinical evaluation. It usually consists of spherical lenses, cylindrical lenses, prism lenses, auxiliary lenses and mechanical lens replacement structures. Use the subject's subjective expression of the clarity of the visual target image to measure the refraction state of the subject, and use it in conjunction with the eye chart. It is used to subjectively measure the refractive state of the human eye. | II |  | 723 | NEW
724 | 16-03-01 | Retinoscope | The retinoscope is usually composed of a refracting mirror, a projection system of a central light-transmitting mirror, an observation system and a power supply, and can be divided into a strip light (rectangular cross-section beam, adjustable focus and azimuth) retinoscope and a point light retinoscope (near circular cross-section light beam, image focus adjustable) two types, each type is divided into several models according to power supply, technical parameters, auxiliary functions, etc.; for objective evaluation of eye refractive error and displacement observation when retinal reflection crosses the pupil. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0718-2009 Ophthalmic Instrument Retinoscope. | II |  | 724
725 | 16-03-01 | Prescription lenses | Optometry lenses are mainly composed of positive and negative spherical trial lenses, positive and negative cylindrical trial lenses, prism trial lenses, auxiliary trial lenses (horse-style rod lenses, cracked lenses, pinhole lenses, black lenses, frosted lenses, flat lenses, cross lenses, etc. Optometry lenses, red filters, green filters, polarizers, cross cylinder lenses, flip mirrors, etc., can be divided into several models or types according to the number and type of prescription lenses, and prescription frames can be attached for inspection For the refractive status of the human eye (hyperopia, myopia, presbyopia, astigmatism) and strabismus and other ophthalmological diseases. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB 17342-2009 Optometry lenses for ophthalmic instruments. | II |  | 725
726 | 16-03-01 | Optometry | The optometry instrument consists of the main body, the power supply and the jaw support. The refractor can measure the spherical diopter, cylinder diopter and astigmatism axis of the eye. | II |  | 726
727 | 16-03-01 | optometrist | The optometry head is a device that can provide spherical lenses, cylindrical lenses, prism lenses and other optical devices. It can be controlled manually or electrically to check the refractive status of the human eye. Product performance indicators adopt the applicable parts of the following reference standards: YY 0674-2008 "Optometry Head for Ophthalmic Instruments" | II |  | 727
728 | 16-03-01 | LCD eye chart | The liquid crystal eye chart is mainly composed of a liquid crystal display, a host computer and a power cord. An eye chart or card displayed by a monitor. For visual acuity measurement. | II |  | 728
729 | 16-03-02 | Synoptic machine | It usually consists of two sets of independent optical systems with adjustable spatial orientation and brightness, and a mechanical system that can measure the spatial orientation changes of the two systems, and can be combined with other auxiliary components (such as Haidinger brushes, polarizers) use. It is used to check binocular vision functions such as simultaneous vision, fusion, and stereopsis of the human eye, and to diagnose ophthalmic diseases such as subjective and objective oblique angles, abnormal retinal correspondence, phoria, afterimage, and amblyopia. It can also be used for amblyopia training and treatment. | II |  | 729 | NEW
730 | 16-03-02 | Pupil distance measuring instrument, pupil distance meter | It usually consists of light source, sight mark, reading system, mechanical adjustment system, observation or display system. Use the principle of optical imaging positioning to measure the interpupillary distance of human eyes. Used to measure the distance between the two pupils of the human eye. | II |  | 730 | NEW
731 | 16-03-02 | Ocular Aberrometer, Whole-ocular Wavefront Aberrometer | It usually consists of optical positioning system, aberration measurement system, signal detector and data processing and analysis system. The optical path difference derived by ray tracing is analyzed by the method of obtaining the ZERNIK coefficient through fitting to analyze the wavefront aberration. Used to measure aberrations of the human eye. | II |  | 731 | NEW
732 | 16-03-02 | Perimeter | Perimeter is mainly composed of background screen, stimulation point display device, jaw rest, forehead rest, pupil monitoring device, response device, non-independent lifting platform, control circuit/computer host, etc., which can be different according to design, technical parameters, additional auxiliary functions, etc. Divided into several models; used for judging and diagnosing visual path diseases, such as glaucoma and fundus diseases. Visual field follow-up can also be used to judge the development of certain diseases and guide treatment. The product performance index adopts the applicable parts of the following reference standards, such as: YY 0676-2008 Perimeter for Ophthalmic Instruments. | II |  | 732
733 | 16-03-02 | Manual Keratometer | The manual keratometer is a manual estimation type, which consists of five parts: the curvature meter part, the moving base, the instrument table, the electric box, and the head frame. It can be divided into several models according to the design, technical parameters, and additional auxiliary functions; it is used for checking the cornea. Curvature and astigmatism axis measurement. | II |  | 733
734 | 16-03-02 | Keratometer | The keratometer is composed of a host (optical system, observation system and control system), a movable worktable and a headrest; it can be divided into several models according to the design, technical parameters, additional auxiliary functions, etc.; it is used to measure the patient's cornea and contact lens The central curvature and meridian direction of . Product performance indicators adopt the applicable parts of the following reference standards. Such as: YY 0579-2005 Keratometer. | II |  | 734
735 | 16-03-02 | Eye Electrophysiology Equipment | The product consists of a stimulator, a biosignal amplifier, a display device, etc. It is used to measure ocular electrophysiological signals and provide reference information for doctors to diagnose. A variety of measurement procedures can be performed, such as ERG, VEP, MERG, MVEP, EOG, etc. | II |  | 735
736 | 16-04-03 | Ultrasonic pulse echo equipment for ophthalmology | Including ophthalmic A-ultrasound, B-ultrasound and A/B-ultrasound, it is composed of a host and a probe. The nominal frequency of the probe is generally above 10MHz. It uses the principle of ultrasonic pulse echo to complete the collection, display, and measurement of ophthalmic diagnostic information. Ultrasonic diagnostic equipment dedicated to ophthalmology, which can realize the functions of ultrasonic imaging of the eyeball and orbit, corneal thickness measurement, and eye axial length measurement. Excluding ophthalmic high-frequency ultrasonic diagnostic equipment (commonly known as ultrasonic biomicroscope UBM) | II | Revision management category description. | 736
737 | 16-04-05 | eye camera | It usually consists of lighting system, observation system, imaging system and so on. It can be used with separate viewing software and realize additional functions. Used to take eye images, observe and diagnose retinopathy. | II |  | 737
738 | 16-04-07 | Slit Lamp Microscopy | The slit lamp microscope is an ophthalmic binocular (or trinocular) microscope. The instrument projects a bright slit into the eye, and uses a movable microscope to observe the reflected light from the side, so as to observe and observe the reflective surface. / or measure. A typical structure consists of a microscope and a rotating lighting system that can produce slits (narrow slits). It can be hand-held or desktop. The microscope can be continuously zoomed or discontinuously zoomed; if equipped with a camera and computer, it can also be managed by software The image can be divided into several models due to the difference in design and technical parameters; it is used for checking the anterior segment and other internal eye lesions. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0065-2007 Ophthalmic Instruments Slit Lamp Microscope. | II |  | 738
739 | 16-04-08 | direct ophthalmoscope | Direct ophthalmoscopy consists of an illumination system and an observation system (consisting of a fluoroscopy lens and a diaphragm set), producing a forward or non-inverted image, allowing the examiner to clearly observe details of the retina and other structures/media (cornea, aqueous humor, lens, and vitreous) ), mostly hand-held, battery-powered (also powered by grid power), a typical structure consists of a light source installed inside, a reflector and a dial, the direction of the light source can be changed by the examiner, and the inside of the eyeball is illuminated through the pupil , There is a round hole in the middle of the mirror, the examiner can observe through it, there are compensation lenses and auxiliary lenses on the dial, which can be selected during inspection, and are used to check fundus lesions, abnormalities of refractive media, and retinal positioning. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1080-2009 Ophthalmic Instruments Direct Ophthalmoscope. | II |  | 739
740 | 16-04-09 | indirect ophthalmoscope | An indirect ophthalmoscope consists of a condenser lens (handheld or integral) to produce an intermediate real image that can be viewed visually, usually a primary light source to illuminate the interior of the eye, a lens and filter system, and a Composed of reflective lighting beams, it can be divided into handheld or headband type, which can be powered by internal power supply or network power supply; it is used for intraocular medium and fundus examination. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY0633-2008 Ophthalmic Instruments Indirect Ophthalmoscope. | II |  | 740
741 | 16-04-12 | corneal topograph | The corneal topograph is composed of a data acquisition unit, a data correction/processing unit and accessories, which may include a computer and special software. It can be divided into several models according to the working principle, design, technical parameters, additional auxiliary functions, etc.; through non-contact measurement The surface shape of the human cornea is used to obtain parameters such as corneal curvature and refractive power. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0787-2010 Ophthalmic Instruments Corneal Topography. | II |  | 741
742 | 16-04-13 | corneal pachymeter | It usually consists of an optical emission system and an optical measurement system. Corneal thickness is measured by the deviation method of two reflected light paths or the optical path difference method. Used to measure corneal thickness. | II |  | 742 | SPLIT/NEW
743 | 16-04-17 | Tonometer | All products under classification code 16-04-17 are exempt from clinical evaluation. It usually consists of a corneal shape change generator, a corneal deformation measurement system or a contact corneal device, and a pressure change sensor. Intraocular pressure can be obtained through conversion of corneal shape changes (applanation, indentation, non-contact, etc.) or direct measurement of corneal blood flow pulsation pressure changes. Used to measure pressure inside the eye. | II |  | 743 | NEW
744 | 16-04-17 | Indentation tonometer | The indentation tonometer usually consists of a fixed weight, an indenter, a hammer bow and a pointer; it can be divided into different models according to the type (handheld or desktop); an external force is manually applied to the eye through the instrument, and it is measured by the indenter Examine the degree of corneal ingrowth, and indirectly measure intraocular pressure through conversion. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 1036-2004 Indentation tonometer. | II |  | 744
745 | 16-04-17 | applanation tonometer | The applanation tonometer is an optical applanation type, which consists of a weight control box, a double-image prism, a measuring arm, a measuring knob, mounting components and calibration components, and can be divided into several models according to the connection with the slit lamp microscope; For the measurement of intraocular pressure. | II |  | 745
746 | 16-04-17 | rebound tonometer | The rebound tonometer is composed of a probe, a measuring system and a display system, and can be divided into several models according to the design, technical parameters and additional auxiliary functions; it is used for measuring the intraocular pressure. | II |  | 746
747 | 16-04-17 | non-contact tonometer | The non-contact tonometer is composed of display system, air jet, measuring optical path system and three-dimensional motion system. It can be divided into several models according to design, technical parameters and additional auxiliary functions; it is used for measuring human intraocular pressure. | II |  | 747
748 | 16-04-19 | dry eye detector | It usually consists of a light source, a CCD camera and a control system, and is used in combination with a computer and supporting software. Through interference imaging of the patient's ocular surface tear film, it is used to diagnose the degree of dry eye. | II |  | 748
749 | 16-04-21 | Ophthalmic Diagnostic Aids | All Class II products under the 16-04-21 classification code are exempt from clinical evaluation. | II |  | 749 | NEW
750 | 16-04-21 | Tear secretion test filter paper | Tear secretion detection filter paper is cut from filter paper with fluorescein sodium marking line, and is used for clinical detection of various tear secretion disorders | II |  | 750
751 | 16-05-05 | Ophthalmic Surgical Microscope | This product is generally composed of an optical system, a control system, an illumination system, and a bracket system, excluding optional parts or components with special functions, and is clinically used to provide magnification and illumination for ophthalmic surgery. | II |  | 751
752 | 16-05-06 | Intraocular Illuminator Fiber Optic Probe | Usually consists of an optical fiber and an insertion head. Supplied sterile. Used in connection with ophthalmic lighting sources, it can be inserted directly into the human eye, and can also be used with a trocar system. It is used for intraocular illumination by transmitting the light emitted by the illumination source during intraocular surgery. | III |  | 752
753 | 16-05-06 | glass cutting head | It is usually composed of a vitrectomy head, a joint, a conduit and a handle, and the head is generally made of stainless steel. Used in conjunction with pneumatic devices for removal of the vitreous during ophthalmic surgery. It is equivalent to the product that has been approved for domestic registration, and the product composition materials are mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 753 | NEW
754 | 16-05-06 | Disposable Intraocular Laser Fiber | It is composed of laser connector, optical fiber and handle, and is used to transmit laser light into the eye after being connected with the laser. The light output end of the optical fiber is a straight cut end face, without other functions such as suction or lighting. | III |  | 754 | NEW
755 | 16-05-06 | Contact Ophthalmic Viewing Lenses | An optical device (in contact with the cornea) that assists ophthalmic treatment and surgery through optical principles such as magnification or inversion. Used as an adjunct to ophthalmic therapy and surgery. | II |  | 755
756 | 16-05-06 | Corneal Lamellar Blades | Usually of flake design, with cutting edges. Supplied sterile, single use. It is used in conjunction with the corneal lamellar knife to section and cut the cornea. | II |  | 756
757 | 16-05-06 | intraocular illuminator | Usually consists of an intraocular illuminator, an intraocular lighting fiber optic and an insertion head (supplied sterile, single use), a detachable and sterilizable knob. For the illumination of the eye during ophthalmic surgery. | III |  | 757
758 | 16-06-01 | soft contact lens | Daily wear, single focus design, correct myopia or hyperopia. It is equivalent to the product that has been approved for domestic registration, and the product formula is mature. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 758
759 | 16-06-01 | rigid contact lenses | Daily wear, single focus design, correct myopia or hyperopia. It is equivalent to the product that has been approved for domestic registration, and the product formula is mature. Excluding new structural design, new mechanism of action, drug-device combination products, and rigid gas-permeable contact lenses for orthokeratology. | III |  | 759
760 | 16-06-02 | contact lens saline solution | An aqueous solution with a content of 0.9% sodium chloride, used for washing and storing soft or/and hard contact lenses, and dissolving and diluting tablet care products, etc. It is equivalent to products that have been approved for domestic registration. | III |  | 760
761 | 16-06-02 | Contact Lens Care Products | For cleaning, disinfection, rinsing, storage, etc. of soft or hard contact lenses. The exception is contact lens lubricant. It does not contain components that are not contained in similar listed products in my country. The concentration of active ingredients in the product should not exceed the concentration of active ingredients in products already on the market. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III | Add | 761
762 | 16-06-03 | Medical Light Radiation Protective Glasses | Usually consists of a frame and lenses. Lenses are made of materials that can reflect or absorb radiation, but can transmit certain visible light. For protection against UV, blue, red and infrared hazards during diagnostic or surgical procedures. | II |  | 762 | NEW
763 | 16-06-03 | laser safety glasses | It is generally made of glass or plastic lenses with absorbents added. Used to prevent damage to human eyes from laser radiation during diagnosis or surgery. | II |  | 763
764 | 0 | prosthetic eye platform | A spherical or ellipsoidal object implanted in the body. Made of hydroxyapatite material. Drug-device combination products are excluded. It is suitable for filling after loss, removal or atrophy of the eyeball and orbital wall, and for intraorbital support. The product is equivalent to the product that has been approved for domestic registration. New structural design, new mechanism of action, and drug-device combination products are excluded. | III |  | 764
#N/A
765 | 16-07-15 | Disposable foldable intraocular lens injector | The disposable foldable intraocular lens injector can be made of suitable materials such as polypropylene, and can be composed of two parts: a propeller and an implanter; it can be divided into several models according to different materials, designs, technical parameters, etc.; in sterile form Provided; used for bolus injection of foldable intraocular lens during intraocular lens implantation. | II |  | 765
766 | 17-00-00 | retention fiber | The main component of the retention fiber is a polymer material such as polyethylene terephthalate or glass fiber, which can be divided into several types according to the material, design, physical properties, size, etc.; it is used for the temporary restoration of loose teeth and missing teeth , dentition gap maintenance, temporary or permanent reinforcement and repair of fixed bridges, etc. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 766
767 | 17-00-00 | denture cement | The denture stabilizer can be a paste made of sodium carboxymethylcellulose and polymethylvinyl ether sodium salt, calcium salt, malic acid mixture and other auxiliary materials, which can be divided into several types according to different formulations and technical parameters. species; for denture restoration, coated between the denture tissue surface and the mucous membrane gap. | II |  | 767
768 | 17-00-00 | Disposable Oral Anesthesia Catheter | The product consists of a hose, a hose connector, a handle, and a sleeve, with or without a disposable sterile injection needle. Used in conjunction with the oral anesthesia booster, it is suitable for delivering anesthetics during oral local anesthesia. | III |  | 768
769 | 17-01-01 | Periodontal pocket detection equipment | It usually consists of a probe, handle, foot switch and power supply. Periodontal pocket depth was measured by palpation of the periodontal periodontal with a pressure-sensitive electronic probe. Used to detect periodontal pocket depth. | II |  | 769
770 | 17-01-06 | dental examination light | It usually consists of lighting device and detection and observation device. The lighting device usually includes a lighting handle, a power supply, and patient goggles; the detection observation device usually includes a viewing glass. It is used for oral lighting and detection observation, and assists in enhancing the visualization of abnormal mucosa and oral lesions in oral examination. | II |  | 770 | SPLIT/NEW
771 | 17-02-02 | Disposable sterile mouth mirror | Usually consists of a handle, a mirror with or without a connecting rod. Mirrors are generally made of stainless steel or glass. Supplied sterile, single use. | II |  | 771
772 | 17-03-01 | Dental unit | Electrically driven equipment for dental diagnosis, treatment, and surgery. It can be composed of a dental treatment machine and a dental patient chair. The dental treatment machine can be composed of a three-purpose spray gun, a film viewing lamp, an instrument tray, a saliva suction component, a strong suction component, a medical foot switch, a floor box, and a water purification system. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 1043.1 Dental treatment machine, YY/T 0058 Dental patient chair, etc. | II |  | 772
773 | 17-03-02 | Dental Patient Chair | Used to support a patient in a seated position for dental examinations, treatment and/or monitoring of surgical procedures. Height is usually adjustable to allow healthcare workers to operate while standing. Typically includes a headrest and armrests, the reclining chair back can be tilted from a vertical position to a horizontal position, and can have a swivel function; dental examination and/or treatment (such as lights, irrigation devices) instruments can be fixed as components of the surgical chair, or Standalone wall or pendant mounted units which may be electrically driven. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0058 Dental patient chair. | II |  | 773
774 | 17-03-02 | Dental Chair | It usually consists of a base support system, a control system, a foot switch, a seat surface and a headrest. Active product. It is used to carry patients (dental chair) and medical staff (physician chair) during dental clinical diagnosis and treatment. | II |  | 774
775 | 17-03-03 | Ultrasonic scaler | Ultrasonic dental scaler applies ultrasonic energy to the teeth through the tip of the handle for cleaning the tooth surface, root canal, periodontal and other parts; it usually consists of an ultrasonic generator, a handle, and various action heads (tips) matched with the handle , can have a flushing part; it can be divided into several models according to the design, technical parameters, whether it has a working head, additional auxiliary functions, etc.; it is used for clinical tartar removal. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0460-2009 Ultrasonic dental cleaning equipment. | II |  | 775
776 | 17-03-03 | Pneumatic Scaler | Pneumatic dental scaler is composed of pneumatic host/handpiece, connector, cleaning head, accessories, etc. It can have functions such as auxiliary lighting and water spray, and can be divided into several models according to design, technical parameters, auxiliary functions, and intended use; It can be used to clean the tooth surface, root canal, periodontal and other parts. | II |  | 776
777 | 17-03-03 | Pneumatic dental blasting device | Pneumatic dental sandblasting device is composed of handle, connection device, etc., and can be divided into several types according to material, design, technical parameters, etc.; with the help of artificially controlled high-pressure air flow and the tiny alumina particles contained in it, the tooth surface and Adhesive tartar and stains on periodontal, root canal and other parts or polishing tooth surface. | II |  | 777
778 | 17-03-04 | dental handpiece | A dental handpiece is a handpiece used to deliver the energy needed to work with dental tools or appliances. The products cover high-speed gas turbine handpieces meeting the requirements of YY1045.1 "Dental handpieces Part 1: High-speed gas turbine handpieces", and dental contra-angle handpieces meeting the requirements of YY 1045.2 "Dental handpieces Part 2: Straight handpieces and contra-angle handpieces" or Dental Straight Handpieces, and Polishing Handpieces. | II |  | 778
779 | 17-03-04 | Dental Low Voltage Electric Motors | Dental low-voltage electric motors are controlled by a host computer or as a stand-alone mobile device. It is electrically connected to the host through the motor cable and is powered and controlled by the latter. Dental low-voltage electric motors are used to provide the driving force for dental handpieces to perform dental operations. Product performance indicators are formulated with reference to YY0836-2011. | II | Merged with the original 759. | 779
780 | 17-03-04 | Dental Air Motor | The dental air motor is controlled by the dental unit. Dental air motors are used to provide the driving force for dental handpieces to perform dental procedures. Product performance indicators are formulated with reference to YY0837-2011. | II |  | 780
781 | 17-03-04 | Dental Handpieces & Accessories | According to different power sources, it can be divided into pneumatic and electric. Among them, the air motor generates power by pushing the blades to rotate through compressed air, and the electric motor generates power through electromagnetic principles. For use with dental handpieces. Used to drive dental handpieces. | II |  | 781
782 | 17-03-06 | Dental light curing machine | The blue light emitted by the product is used to irradiate the repair material to cure it. It is divided into two types: the light curing machine with the quartz tungsten halogen lamp as the light source and the LED light curing machine with the light-emitting diode (LED) lamp as the light source. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0055 light curing machine series standards. | II |  | 782
783 | 17-03-07 | dental implant | A dental implant usually consists of a host, a motor and a foot switch. For dental implant surgery. | II | Add | 783
784 | 17-03-09 | root canal preparation machine | The root canal preparation machine is an electric drive type root canal surgery equipment. According to its different technical structure, there are usually two compositional structures. The first type is composed of host, motor cable, motor, mobile phone, foot switch, etc.; the second type is composed of host, battery, charger, mobile phone, etc. Neither root canal file is included in the composition. The root canal preparation machine is clinically used in root canal treatment, and the root canal preparation stage is shaped and cleaned. The product does not include the root apex positioning function (or root canal length measurement function). | II |  | 784
785 | 17-03-09 | Dental filling instrument | The gutta-percha filling instrument is an electric-driven root canal surgery device. The product usually consists of a host, a handle (or a hand piece), a working tip (a channel for pushing the gutta-percha), a power supply, etc. The gutta-percha filling instrument is clinically used in the root canal filling stage of root canal surgery. The intended purpose is to soften the gutta-percha under pressure and fill the gutta-percha into the prepared root canal. The product does not include the root apex positioning function (or root canal length measurement function). | II |  | 785
786 | 17-03-09 | root apex locator | It usually consists of a host, a lip hook, a measuring lead, a probe and a power supply. It is used to assist in determining the working length when used in dental clinical root canal treatment. | II | Add | 786 | NEW
787 | 17-03-10 | Oral anesthesia injection equipment | It usually consists of a booster tube and a host. The speed and flow of injection are controlled by setting programs to realize automatic injection. For the injection of oral anesthetics. | II |  | 787
788 | 17-03-12 | tooth grinder | Usually consists of a grinder and a disposable sterile grinding container. Usually electric. It is used to grind tooth fragments into bone powder during oral surgery for use in surgery. | II |  | 788
789 | 17-04-01 | Soft Tissue Circumcision Knife | For use with dental handpieces, constructed of suitable materials such as stainless steel, titanium, or supplied sterile. For shaping the gums in dental implant surgery. The raw materials used in the product must meet the requirements of relevant national and industry standards, and the product performance indicators must meet the applicable parts of the following standards, such as: YY91010 "Dental Rotary Instrument Fitting Dimensions", etc., and the basic principles, scope of application, performance and composition are the same as those already on the market The product is the same. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. Supplied sterile. | II |  | 789
790 | 17-04-06 | Machine Root Canal Instruments | One of the dental endodontic instruments used in conjunction with active medical devices. It consists of a handle and the working end of the head, and the body material is usually stainless steel, nickel-titanium alloy and other applicable materials. Used for root canal preparation and auxiliary filling in dental root canal treatment. It can be divided into several models and specifications according to the material, type and size of the handle and head. The material of the product body must meet the requirements of relevant national and industry standards, and the performance indicators must meet the applicable parts of the following standards, such as: YY 0803.1 "Dental Endodontic Instruments Part 1: General Requirements and Test Methods", YY/T 0803.2 "Dental Scientific Root Canal Instruments Part 2: Expanders, etc., and the basic principles, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 790
791 | 17-04-06 | dental file | Generally, mechanical instruments made of metal materials such as stainless steel and nickel-titanium alloy. The working end has structural design such as engraving or spiral cutting edge, which plays the role of cutting and smoothing. The product is connected to a mobile phone and powered by the host. Used with active instruments, it is used for cutting, leveling, cleaning and shaping dental bone and root canal in dental treatment. | II |  | 791
792 | 17-04-07 | Dental Bur | Dental bur is a kind of dental rotary instrument, which consists of a handle (steel or suitable material) and a head working end (made of steel or hard alloy, diamond, etc.). It can be divided into several models and specifications according to the material, type, and size of the handle and head. Product performance indicators adopt the applicable parts of the current versions of the following industry standards, such as: YY 0761.1 Dental diamond rotary instruments Part 1: Dimensions, requirements, marking and packaging, YY 0302.1 Dental rotary instrument burs Part 1: Steel and hard Quality alloy bur. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 792
793 | 17-04-07 | Drills for root canal post preparation | A type of dental rotary instrument, which consists of a handle and a working end of the head. The body material is usually stainless steel, titanium alloy and other applicable materials. It is used in the process of oral root canal post restoration to remove part of the root canal filling material and prepare the root canal shape that matches the root canal post. It can be divided into several models and specifications according to the material, type and size of the handle and head. The material of the product body must meet the requirements of relevant national and industry standards such as YY/T0294.1 "Metal Materials for Surgical Instruments Part 1: Stainless Steel", and the performance indicators must meet the applicable parts of the following standards, such as: YY 91064 Dental Rotary Instrument Steel and Hard Steel Quality alloy dental drill technical conditions, YY 91010 dental rotary instrument matching dimensions, etc., and the basic principle, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 793
794 | 17-04-07 | Dental drill for implant surgery | A type of dental rotary instrument, which consists of a handle and a working end of the head. The body material is usually stainless steel, titanium alloy and other applicable materials. It is used in the process of oral implant restoration to prepare the implant cavity that matches the type of implant used. It can be divided into several models and specifications according to the material, type and size of the handle and head. The material of the product body must meet the requirements of relevant national and industry standards such as YY/T0294.1 "Surgical Instrument Metal Materials Part 1: Stainless Steel", and the performance indicators must meet the applicable parts of the following standards, such as: YY 91064 "Dental Rotary Instrument Steel and Carbide Dental Drill Technical Conditions", YY91010 "Dental Rotary Instrument Matching Dimensions", etc., and the basic principles, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 794
795 | 17-04-07 | Dental drill for implant attachment extraction | A type of dental rotary instrument, which consists of a handle and a working end of the head. The body material is usually stainless steel, titanium alloy and other applicable materials. It is used to take out the abutment, break the screw in the implant and clean the thread during the implant surgery. It can be divided into several models and specifications according to the material, type and size of the handle and head. The material of the product body must meet the requirements of relevant national and industry standards such as YY/T0294.1 "Surgical Instrument Metal Materials Part 1: Stainless Steel", and the performance indicators must meet the applicable parts of the following standards, such as: YY 91064 "Dental Rotary Instrument Steel and Carbide Dental Drill Technical Conditions", YY91010 "Dental Rotary Instrument Matching Dimensions", etc., and the basic principles, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 795
796 | 17-04-07 | Dental burs (non-metallic) | One of the dental rotary instruments, made of resin, polyether ether ketone (peek) and other materials. It consists of a handle and the working end of the head. It can be divided into several models and specifications according to the material, type and size of the handle and head. It is used to cut tooth tissue, remove diseased tissue, treat drilling or prepare tooth. Product performance indicators adopt the applicable parts of the current versions of the following national standards, such as: YY 0761.1 Dental diamond rotary instruments Part 1: dimensions, requirements, marking and packaging, YY 0302.1 Dental rotary instrument burs Part 1: Steel and hard Quality alloy bur. | II |  | 796
797 | 17-04-07 | Drill for orthodontic anchorage | A type of dental rotary instrument, which consists of a handle and a working end of the head. The body material is usually stainless steel, titanium alloy and other applicable materials. Before implanting orthodontic anchorage, drill holes in the jaw to determine the implantation position of orthodontic anchorage. It can be divided into several models and specifications according to the material, type and size of the handle and head. The material of the product body must meet the requirements of relevant national and industry standards such as YY/T0294.1 "Metal Materials for Surgical Instruments Part 1: Stainless Steel", and the performance indicators must meet the applicable parts of the following standards, such as: YY 91064 Dental Rotary Instrument Steel and Hard Steel Quality alloy dental drill technical conditions, YY 91010 dental rotary instrument matching dimensions, etc., and the basic principle, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 797
798 | 17-04-10 | Intraoral Grinding and Polishing Instruments | A type of dental rotary instrument used in conjunction with a dental handpiece. It consists of a metal handle and a working end, which is disc-shaped, cup-shaped, etc. It is used for polishing and grinding of restorations, removal of excess parts, and cleaning of implants. It is also used for descaling and polishing of the tooth surface. It can be divided into several models and specifications according to the material, type and size of the metal handle and working end. The raw materials used in the product must meet the requirements of relevant national and industry standards, and the product performance indicators must meet the applicable parts of the following standards, such as: YY91010 "Dental Rotary Instrument Fitting Dimensions", etc., and the basic principles, scope of application, performance and composition are the same as those already on the market The product is the same. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 798
799 | 17-04-10 | Grinding and polishing cleaning utensils | Divided into manual and motorized two categories. Manual products are usually composed of a hand-held part and a working end; motorized products are usually composed of a metal handle and a working end, the metal handle is connected to the dental handpiece, and the working end is in the shape of a brush or a cup. Used for polishing, grinding and removing excess parts of restorations, or cleaning of implants. It is also used for descaling and polishing of the tooth surface. Appliances used only in dental laboratories are not included. | II |  | 799
800 | 17-04-11 | Sterile sinus lift balloon | This product is composed of a balloon made of latex material, a cap, a catheter, a connector and a disposable syringe, and is used in dental implant surgery to lift the maxillary sinus. | II |  | 800 | NEW
801 | 17-04-11 | implant installation aids | It is usually used for maxillary sinus lifting balloons, etc., and is used in dental implant procedures. passive product. Supplied sterile. For dental implant procedures. | II |  | 801
802 | 17-04-14 | Disposable dental flush needle | Disposable dental flushing needles with flat head are made of polymer materials and high-grade stainless steel materials. Class; provided in a sterile form; used together with a syringe for oral pressure washing after dental diagnosis and treatment. | II |  | 802
803 | 17-04-14 | Oral Cleaning Appliances | Passive product for rinsing the mouth. Supplied sterile, single use. Used to remove debris or debris from the mouth. | II |  | 803
804 | 17-04-15 | Disposable sterile nozzle | It is generally made of materials such as stainless steel or plastic. When in use, it is installed on the front end of the spray gun of the dental comprehensive treatment unit, and is the outlet port of the air flow and liquid of the spray gun. Supplied sterile, single use. | II |  | 804
805 | 17-04-15 | Disposable sterile saliva suction tube | It is generally made of materials such as stainless steel or plastic. Typically used with the suction unit of a dental unit. Supplied sterile, single use. | II |  | 805
806 | 17-04-15 | spray gun | It usually consists of a gun body, a handle and a spray rod. According to the way of holding it in use, it can be divided into two types: curved type and straight type. According to needs, compressed air, water or air-water spray can be delivered to a certain part of the oral cavity in the form of spraying. Usually connected to the dental unit for use. For oral cleaning and drying. | II |  | 806
807 | 17-05-01 | Glass ionomer cement for light-cured orthodontic bonding/dual-cure glass ionomer cement for orthodontics | Generally, it is powder liquid or paste. As a temporary material placed into the human body, it is used for bonding of orthodontic products, short-term bonding or temporary filling. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 807
808 | 17-05-01 | Zinc Oxide Eugenol Cement | Generally, it is powder liquid or paste. Powders are metal oxides or metal salts that can precipitate ions, such as zinc oxide, glass powder, calcium hydroxide, etc.; liquids are acid solutions or chelates, such as phosphoric acid, polycarboxylic acid, salicylic acid, and eugenol; Some cements contain polymerizable resin components. Most of them are cured by acid-base reaction, and those containing resin components can be cured by chemical reaction or light curing reaction. As a temporary implanted material in the human body, it is used for bonding of orthodontic products, short-term bonding or temporary filling. | III |  | 808
809 | 17-05-01 | Zinc oxide cement without eugenol | It is mainly composed of zinc oxide and mineral oil. For dental temporary restorations and underlayments. Products exempted from clinical practice must at least meet the relevant requirements of YY 0272 Dentistry Zinc Oxide/Eugenol Cement and Zinc Oxide Cement Without Eugenol. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 809
810 | 17-05-01 | zinc polycarboxylate cement | Cement cured by reacting zinc oxide with aqueous solution of polyacrylic acid or polyalkenic acid or reacting zinc oxide/polyalkenic acid powder with water. Used for bottoming, lining and cementing of cavities. Products exempted from clinical practice must at least meet the applicable parts of YY 0271.1 Dental water-based cements Part 1: powder/liquid acid-base cements and YY 0272 Dentistry zinc oxide/eugenol cements and eugenol-free zinc oxide cements. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 810
811 | 17-05-01 | zinc phosphate cement | A cement cured by the reaction of oxide powder (mainly composed of zinc oxide) with aqueous phosphoric acid. Used for bottoming, lining and cementing of cavities. Products exempted from clinical practice must at least meet the relevant requirements of YY 0271.1 Dental water-based cement Part 1: powder/liquid acid-base cement. | III |  | 811
812 | 17-05-01 | glass ionomer cement | Cement cured by reacting aluminosilicate glass powder with alkene acid aqueous solution, or by reacting aluminosilicate glass powder/polyacid powder with water or tartaric acid aqueous solution. It is used for restoration and filling of foundation, hole lining, cementation, deciduous teeth and permanent teeth cavities (class III holes, class V holes, wedge-shaped defects). Products exempted from clinical practice must at least meet the relevant requirements of YY 0271.1 Dental water-based cement Part 1: powder/liquid acid-base cement. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 812
813 | 17-05-01 | calcium hydroxide cement | The main ingredients are calcium hydroxide and chelating agent. For low strength underlayment, liners, indirect liners or pulp capping. Products exempted from clinical practice must at least meet the applicable parts of YY 0271.1 Dental water-based cements Part 1: powder/liquid acid-base cements and YY 0272 Dentistry zinc oxide/eugenol cements and eugenol-free zinc oxide cements. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 813
814 | 17-05-01 | Resin cement for post and core | It is generally composed of bisphenol A-glycidyl methacrylate (Bis-GMA) and the like. Used in conjunction with root canal posts, it is used for filling and repairing tooth tissue during post-core restoration. The performance, structure, composition, etc. belong to the resin cement that can be fully covered by YY 1042 "Polymer-based Restorative Materials in Dental Science", and the basic principle, scope of application, performance and composition are the same as those of the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions compared with products already on the market. | III |  | 814
815 | 17-05-01 | Resin cement for cementing | It is generally composed of bisphenol A-glycidyl methacrylate (Bis-GMA) and the like. It is used for cementation or retention of restorations such as crowns and bridges. The performance, structure, composition, etc. belong to the polymer-based cementing materials that can be fully covered by YY 1042 "Polymer-based Restorative Materials in Dentistry", and the basic principle, scope of application, performance and composition are the same as those of products that have been marketed. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 815
816 | 17-05-02 | orthodontic adhesive | It is mainly composed of resin, filler, solvent and so on. Bonding treatment for orthodontic treatment. | II |  | 816
817 | 17-05-02 | Dental Adhesives | It is a viscous liquid, mainly made of methacrylate, etc. It is used for bonding between restorations or restoration materials and tooth bonding and restoration materials. Performance, structure, composition, use, etc. belong to the dental adhesives that can be fully covered by YY/T 0518 "Polymer-based Adhesives for Dental Restorations", and the basic principles, scope of application, performance and main components are the same as those of products already on the market same. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 817
818 | 17-05-03 | gutta percha | It is mainly composed of gutta-percha, zinc oxide, barium sulfate, etc. It belongs to dental solid root canal filling material. Used in conjunction with root canal paste for root canal filling during root canal treatment. Products exempted from clinical practice must at least meet the relevant requirements of YY 0495 root canal filling tips. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 818
819 | 17-05-06 | Silver amalgam | Powder-liquid two-component or capsule type. The liquid is mercury, and the powder is silver alloy. Silver amalgam is formed by the amalgamation reaction. For direct filling restoration of tooth defects. Products exempted from clinical practice meet the relevant requirements of YY 1026 dental mercury and silver alloy powder. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 819
820 | 17-05-07 | temporary filling material | Solid, one-component or two-component pastes. The solid is usually gutta-percha; the paste usually consists of a resin matrix, surface-treated inorganic fillers, an initiating system, etc., or calcium sulfate/zinc, zinc oxide, a resin matrix, etc. For temporary restoration of tooth defects. | II | Add | 820 | NEW
821 | 17-05-08 | Light curing calcium hydroxide indirect pulp capping agent | Light-curing calcium hydroxide indirect pulp capping agent is a light-curing radio-blocking calcium hydroxide indirect pulp capping agent, the main components are usually calcium hydroxide, urethane dimethacrylate (UDMA), triethylene dimethacrylate alcohol dimethacrylate (TEDMA), etc. The product is not used for direct pulp capping, but as an indirect pulp capping agent for other filling materials. The exemption does not include products that use new materials, new technologies, or have new mechanisms of action or functions compared with products already on the market. | III |  | 821 | NEW
822 | 17-06-00 | Metal materials for dental additive manufacturing | Pure titanium, titanium alloy, and cobalt-chromium alloy metal materials for dental additive manufacturing are used to process non-implanted restorations such as crowns and bridges through laser selective melting. The material and performance indicators of the product body must meet at least YY∕T 1702-2020 "Dentistry Additive Manufacturing Laser Selective Melting Metal Materials for Oral Fixation and Removable Restoration", GB 17168 "Metal Materials for Dental Fixation and Removable Restoration" and other relevant national and industry standards. The exemption does not include products used to make maxillofacial prostheses, implants, abutments and accessories, and products that use new materials, new technologies, or have new mechanisms of action and new functions compared with products already on the market. | III |  | 822 | NEW
823 | 17-06-01 | cast alloy | Materials for custom-made dentures shall comply with GB 17168 Dentistry Standards for Metal Materials for Fixed and Removable Restorations and YY 0621.1 Dentistry Compatibility Experiment Part 1: Gold Pottery-Ceramic System. Precious metals shall meet the relevant standards of YY 0620. Dental casting gold alloys; nickel-chromium-based alloys, cobalt-chromium alloys, cobalt-chromium alloys, etc. Pure titanium and titanium alloys cannot be used for customization of abutments. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 823
824 | 17-06-01 | Titanium and titanium alloys for dental cutting | Titanium and titanium alloy metal materials for dental cutting are used to manufacture dental crowns, bridges and other restorations through CAD/CAM processing. Metallic Materials for Movable Restoration", and applicable parts of GB/T 13810 "Titanium and Titanium Alloy Processing Materials for Surgical Implants". The exemption does not include products used to make implants and accessories. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 824
825 | 17-06-01 | Dental attachments | It is usually composed of two parts, the female type and the male type. According to the structure, it can be divided into bolt type, bar type, snap type and ball cap type. Auxiliary retention for removable partial dentures, overdentures and other restorations. The raw materials used in the product must meet the requirements of relevant national and industry standards. The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. Exemption does not include products that use new materials, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 825
826 | 17-06-01 | Alloys for Dental Cutting | Cobalt-chromium alloys and nickel-chromium alloys for dental cutting. It is used to manufacture restorations such as crowns and bridges through CAD/CAM processing. The material and performance indicators of the product body must meet the requirements of relevant national and industry standards, such as: the applicable part of GB17168 "Metal Materials for Fixed and Removable Restoration in Dentistry". The exemption does not include products used to make implants and accessories, and does not include products using new materials, active ingredients, new technologies, new designs, or products with new mechanisms of action and new functions. | II |  | 826
827 | 17-06-02 | Porcelain powder for customized dentures | Materials for custom-made dentures, type I powder ceramics, can be divided into several categories according to components, intended uses, etc. , staining, transparency, addition, and glazing. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0716-2009 Dental Ceramics. | II |  | 827
828 | 17-06-02 | Zirconia blocks for all-ceramic dentures | The main components of zirconia blocks for all-ceramic dentures are: zirconia, yttrium oxide, aluminum oxide and other oxides, among which, zirconia (including hafnium oxide, yttrium oxide) (ZrO2+HfO2+Y2O3) material composition (weight percentage) Should be greater than 99%. It can be divided into several models and specifications according to the composition, molding process, shape, etc.; it is used for cutting and making dental fixed dentures by using CAD/CAM (computer-aided design/computer-aided manufacturing) methods, such as crowns, bridges, inlays, and veneers. The product performance index adopts the applicable part of the current version of the following national standards, such as: GB30367 dental ceramic materials. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 828
829 | 17-06-02 | All-ceramic materials for dental ceramic restoration | Materials for custom-made dentures, type I powder ceramics, can be divided into several categories according to components, intended uses, etc. , staining, transparency, addition, and glazing. For the production of all-ceramic single crowns, long bridges and inlays and all-ceramic restorations. The product performance index adopts the applicable part of the current version of the following national standards, such as: GB30367 dental ceramic materials. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 829
830 | 17-06-02 | Glass-ceramic for all-ceramic dentures | Materials for custom-made dentures are usually block-shaped, and the main components are silicon dioxide, phosphorus pentoxide, aluminum oxide, lithium oxide, potassium oxide, zirconium oxide, and other oxides. It is used for the production of all-ceramic veneers, inlays, crowns and bridges. Generally, it is made by powder coating, computer aided design/manufacturing (CAD/CAM), hot die casting and other processes. Products exempted from clinical practice must at least meet the relevant requirements of GB 30367 Dental Ceramic Materials. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 830
831 | 17-06-02 | Staining solution for all-ceramic dentures | It is usually a liquid, and its main components are polyethylene glycol, hydrogen chloride, ferric chloride, erbium chloride, etc. Used for staining ceramic materials in denture manufacture. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 831
832 | 17-06-03 | denture cement | Mainly composed of polymethyl vinyl ether sodium salt, sodium carboxymethyl cellulose, etc. It is used to apply on the surface of denture tissue to make the denture fit closely with the gums and maintain the stability of the denture. The basic principle, scope of application, performance and composition of the product are the same as the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 832
833 | 17-06-03 | Temporary Crown and Bridge Materials | The product is a two-component paste or powder-liquid, or a one-component paste, or a prefabricated product. The main components are polyacrylates, polyvinyl esters, polycarbonates, etc. For the fabrication of dental temporary restorations. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY/T 0826 dental temporary polymer-based crown and bridge materials. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 833
834 | 17-06-03 | Synthetic resin teeth | Mainly made of acrylic polymers, it is suitable for the production of partial dentures and complete dentures. The material and performance indicators of the product body must meet the requirements of relevant national and industry standards, such as: YY 0300 Relevant requirements for artificial teeth for dental restoration. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 834
835 | 17-06-03 | denture base polymer | It is generally composed of two-component paste or powder liquid, powder, or one-component paste , mainly composed of polyacrylate, polyamide and other materials. For the production of denture bases. Product performance indicators must at least meet the relevant requirements of YY 0270.1 "Base Polymers in Dentistry Part 1: Denture Base Polymers", and the basic principle, scope of application, performance and composition are the same as those of products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III | Revised product description. | 835
836 | 17-06-03 | Polymer-based crown and bridge materials | Two-component paste or powder-liquid, powder, or one-component paste, the working principle is that the initiator and catalyst are mixed and solidified (self-setting). The main component is acrylate. For the production of anterior crowns with or without contact with metal surfaces, not for the restoration of tooth defects, and not for the restoration of posterior teeth in load-bearing areas. The basic principle, scope of application, performance and composition are the same as the products already on the market. Products exempted from clinical practice must at least meet the relevant requirements of YY/T 0710 "Polymer-based Restorative Materials in Dentistry". The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 836
837 | 17-06-03 | Dental powder and dental water | It is mainly made of methyl methacrylate (MMA) homopolymer or acrylate copolymer and methyl methacrylate monomer, and is used in the production of dentures. Products exempted from clinical practice must at least meet the relevant requirements of YY 0270.1 Base polymers in dentistry Part 1: Denture base polymers. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 837
838 | 17-06-04 | custom dentures | It is generally made of cobalt-chromium alloy, nickel-chromium alloy, pure titanium, titanium alloy, precious metal alloy, porcelain block, porcelain powder, base resin, synthetic resin teeth and other materials, depending on the needs. Can be fixed or removable such as crowns, bridges, inlays, veneers, post and core, removable partial or complete dentures, etc. All medical device materials used in the production process are materials with registration certificates. | II |  | 838
839 | 17-06-04 | custom fixed dentures | Custom-made fixed dentures are mainly made of metal, ceramics, and polymer materials, and are used for fixed restoration of missing teeth and dentition, including crowns, bridges, veneers, inlays, etc. The denture materials used in the production process are all materials with registration certificates. The exemption does not include customized implants, abutments and accessories, and products that use new materials, new technologies, new designs, or have new mechanisms of action and new functions compared with products already on the market. | II |  | 839
840 | 17-06-04 | Custom removable dentures | Customized movable dentures include partial dentures and complete dentures, mainly made of metal, ceramics, and polymer materials, and are used for the active restoration of missing dentition. The denture materials used in the production process are all materials with registration certificates. The exemption does not include customized implants, abutments and accessories, and products that use new materials, new technologies, new designs, or have new mechanisms of action and new functions compared with products already on the market. | II |  | 840
841 | 17-06-05 | metal pile | Made of metal material. It is fixed in the root canal or crown to ensure the retention of the restoration and prevent the crown or filling restoration from breaking or falling off due to insufficient retention. It is used for the reconstruction of the crown core when the tooth has a large area of defect after root canal treatment. For retention, see Guidelines. It is often used together with resin or post and core materials. Products that are exempted from clinical practice must at least meet the relevant requirements of YY/T 0517 Dental Prefabricated Root Canal Posts. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 841
842 | 17-06-05 | Retaining nail | Generally made of metal materials. It is often used in conjunction with resin and other materials to assist in the restoration of tooth defects and provide retention for the restoration of teeth. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | III |  | 842
843 | 17-06-05 | fiber post | Made of polymer such as epoxy resin, it is a non-customized product with standard dimensions. Refer to the guidelines for reliable retention of the reconstructed coronal nucleus when the tooth has a large area of defect after root canal treatment. Products that are exempted from clinical practice must at least meet the relevant requirements of YY/T 0517 Dental Prefabricated Root Canal Posts. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | III |  | 843
844 | 17-07-01 | Orthodontic brackets | Orthodontic brackets can be made of suitable materials such as stainless steel and ceramics, and can be divided into several models and specifications according to different materials, tooth positions, technical parameters, and use parts; they are matched with orthodontic wires for orthodontic treatment. Product performance indicators adopt the applicable parts of the current version of the following industry standards, such as: YY/T 0915 orthodontic brackets and buccal tubes. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 844
845 | 17-07-02 | Orthodontic wire | Orthodontic wire is made of stainless steel, nickel-titanium alloy, titanium alloy, titanium-molybdenum alloy, copper-nickel-titanium alloy, and can be divided into several models and specifications according to material, elasticity, cross-sectional shape (circular and rectangular) and size. Orthodontic brackets, etc., are used for orthodontics. Product performance indicators adopt the applicable parts of the current versions of the following industry standards, such as: YY/T 0625 Dentistry orthodontic products: Orthodontic wire. | II |  | 845
846 | 17-07-02 | Orthodontic ligature wire | Orthodontic ligature wire can be made of applicable materials such as stainless steel, and can be divided into several types according to material, physical properties, wire diameter/length, etc. Phase ligation, in order to achieve the purpose of fixed and traction tooth movement. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 846
847 | 17-07-03 | orthodontic buccal tube | Generally made of metal materials. The buccal tube is tubular, consisting of a main body and a retractor. It can be divided into several models and specifications according to different materials, designs, and technical parameters; it is matched with orthodontic brackets, orthodontic wires, etc., and is used to fix orthodontics in orthodontic treatment. Silk can also transmit the correction force. Product performance indicators adopt the applicable parts of the current version of the following industry standards, such as: YY/T 0915 orthodontic brackets and buccal tubes. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 847
848 | 17-07-03 | orthodontic band | It is made of suitable materials such as stainless steel and titanium alloy, and can be divided into several types according to different materials, physical properties, and sizes. Used in orthodontic treatment to provide support points for orthodontic wires, expanders, elastomers or other accessories to correct teeth. The raw materials used in the product must meet the requirements of relevant national and industry standards. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 848
849 | 17-07-04 | orthodontic base polymer | It generally consists of powder and liquid. It is mainly made of methyl methacrylate (MMA) homopolymer or acrylate copolymer and methyl methacrylate monomer, and is suitable for the production of orthodontic bases. Products that are exempted from clinical practice must at least meet the relevant requirements of YY 0270.2-2011 Base Polymers in Dentistry Part 2: Orthodontic Base Polymers. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 849
850 | 17-07-06 | Orthodontic Elastomer Attachments | Orthodontic elastomer accessories include orthodontic elastic traction rings, orthodontic elastic wires/tubes, orthodontic rubber chains, orthodontic ligatures, orthodontic spacers, rotating rubber pads, made of natural or synthetic rubber, and can be pressed Different materials, shapes, sizes, etc. are divided into several models and specifications; in the mouth and outside the mouth, it can cooperate with fixed or movable appliances for correction. Product performance indicators adopt the applicable parts of the current versions of the following industry standards, such as: YY/T 0624 Dentistry orthodontic products: Orthodontic elastic body accessories. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 850
851 | 17-07-07 | free tow hook | The free traction hook can be made of suitable materials such as stainless steel, and can be divided into several types according to different materials, physical properties, and sizes. It can be matched with other orthodontic products for orthodontic use. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 851
852 | 17-07-07 | Spiral expander | The spiral arch expander is made of suitable materials such as stainless steel. It can be composed of a screw rod, a nut, a guide rod, and an adjustment rod. It can be divided into several types according to different materials, designs, technical parameters, and applicable parts. It is used to expand the dental arch of those with narrow dental arches gap etc. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 852
853 | 17-07-07 | Anterior maxillary distraction device | The extraoral device for frontal jaw traction is composed of forehead bracket, adjustable bracket, adjuster and mandibular pocket, etc. The forehead bracket and mandibular pocket are made of polypropylene resin, and the adjustable bracket is made of stainless steel; according to material, design, technology It can be divided into several types according to different parameters; it can be used as an auxiliary traction for human maxillary correction with intraoral orthodontic devices. Intraoral orthodontic appliances are not included. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 853
854 | 17-07-07 | molar thruster | The molar propeller is made of suitable metal materials such as stainless steel, and is composed of various types of struts or struts and springs. It can be divided into several types according to different materials, designs, and technical parameters; it is matched with brackets, belt rings, and implant nails. It is used to push molars to obtain the gap required for alignment of malocclusions. Planting nails, brackets, and bands are not included. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 854
855 | 17-07-07 | Tongue buckle | Usually made of stainless steel, titanium alloy and other materials. For orthodontic treatment, it is used together with orthodontic brackets, orthodontic buccal tubes, orthodontic wires, etc. The raw materials used in the product must meet the requirements of relevant national and industry standards, and the basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. The exemption does not include products that use new materials, new technologies, new designs, or have new mechanisms of action or functions compared with products already on the market. | II |  | 855
856 | 17-07-07 | Orthodontic retainer | Products are made of medical grade PVC material, polyethylene terephthalate (PET) , thermoplastic polyurethane or medical silicone, and are non-customized products with standard sizes. It is used to consolidate the curative effect after the completion of orthodontic treatment. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II | Revised product description. | 856
857 | 17-08 | Custom Dental Implant | Usually consists of a base plate and a guide ring. It is used for the positioning and guidance of implants in dental implant surgery, so as to assist in the installation of implants. The substrate is usually made of epoxy or acrylic material, and is made by 3D printing technology according to the CAD model. The guide ring is made of pure titanium titanium alloy and other metal materials. The material and performance indicators of the product body must meet the requirements of relevant national and industry standards, such as: GB/T 3620.1 "Titanium and Titanium Alloy Grades and Chemical Composition", GB/T 13810 "Titanium and Titanium Alloy Processing Materials for Surgical Implants", and the applicable part of YY/T0911 "Dentistry Polymer-Based Substitute Materials". The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II | Revised product description. | 857
| #N/A | Guides
858 | 17-08-02 | protective cap | The protective cap for implant abutment is made of polymethyl methacrylate and plastic polyoxymethylene, cap-shaped, used for temporary restoration process after implant surgery in stomatology, to protect the superstructure or abutment of the implant. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 858
859 | 17-08-02 | Straight Abutment, Healing Abutment, Temporary Abutment and Accessories | Straight abutments, healing abutments made of titanium and titanium alloy materials, also include temporary abutments. Accessories include central screws, abutment screws, closure screws, cover screws, healing caps, sleeves and screws, abutment caps. It is a matching component for dental implants implanted in the jawbone after tooth loss. It is installed on the implant platform anchored in the bone for connecting, supporting and retaining the prosthesis or the superimposed structure of the implant. Products made of titanium and titanium alloy materials that are exempt from clinical evaluation must at least meet the relevant requirements of YY/T 0520 "Titanium and titanium alloy dental implant accessories", and the body materials of titanium and titanium alloy materials must comply with GB/T 3620.1 " Titanium and titanium alloy grades and chemical composition", GB/T 13810 "Titanium and titanium alloy processing materials for surgical implants" relevant requirements. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | III | Revised product name and product description. | 859
860 | 17-08-04 | Implant sealant | Available in one-part paste or other forms. Usually composed of polysiloxane and anhydrous silicic acid, but can also be composed of other ingredients. For closing the gap between the implant and the central screw. Prevent pathogenic bacteria from infiltrating into the implant and cause secondary infection of surrounding tissues. The basic principle, scope of application, performance and composition are the same as the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 860
861 | 17-08-05 | Transfer Caps/Impression Caps | Refers to the aseptically provided transfer cap or impression cap, which is a kind of auxiliary device for dental implant restoration and is made of suitable materials such as titanium alloy. It is used to transfer the position of the implant and abutment in the oral cavity to the working model for implant restoration. The raw materials used in the product must meet the requirements of relevant national and industry standards. The basic principle, scope of application, performance and composition of the product are the same as those of the products already on the market. Exemption does not include products that use new materials, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 861
862 | 17-09 | tooth desiccant | Tooth desiccant is composed of ethanol, acetone, distilled water, etc.; it is applied on the surface of the teeth to be treated to make it clean and dry quickly, so as to facilitate further diagnosis and treatment operations. | II |  | 862
863 | 17-09-01 | calcium hydroxide root canal disinfectant | Calcium hydroxide root canal disinfectant is a one-component or two-component paste, and the ingredients do not have pharmacological effects; the single component is composed of calcium hydroxide and excipients, and the two-component is composed of calcium hydroxide and propylene glycol aqueous solution Composition; it can be divided into several types according to the composition, technical parameters and accessories; it is only for short-term use in the oral cavity, and is used for root canal disinfection. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 863
864 | 17-09-01 | Root canal enlargement solution/root canal cleaning solution/root canal lubricant | It usually consists of ethylenediaminetetraacetic acid/edetic acid (EDTA), carboxymethylcellulose, etc. Used for lubricating, cleaning and irrigating root canals during root canal preparation and enlargement. The basic principle, scope of application, performance and main components of the product are the same as those of the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions compared with products already on the market. | II |  | 864
865 | 17-09-02 | Disposable sterile dental paper tips | Usually paper or cotton fibrous tapered tip. It has good water absorption, is hard and flexible, and is easy to put into the root canal. Supplied sterile, single use. | II |  | 865
866 | 17-09-03 | dental etchant | Usually one-component or two-component liquid or gel. Generally phosphoric acid, lactic acid, citric acid, hydrofluoric acid, oxalic acid, polyacrylic acid, etc. Utilizes the corrosiveness of acid to play a role. When used for intraoral restoration or orthodontic treatment, the corrosion of the etchant is used to treat the surface of the tooth, metal, ceramics and other restorations to remove the contamination layer, rough surface, and improve its surface properties. The mechanism of action and main components must be the same as those of the products already on the market. Product performance indicators adopt the applicable parts of the current versions of the following industry standards, such as: YY 0769 Phosphoric Acid Etching Agent for Dentistry. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 866
867 | 17-09-04 | A silane coupling agent | One-component liquid. Usually acrylate functional monomers with hydrophilic groups, silane coupling agents or other ingredients. For surface treatment of teeth, resins, ceramics, metal restorations, etc. Use its chemical modification to change the surface properties of teeth and restorations. The basic principle, scope of application, performance and composition of the product are the same as the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 867
868 | 17-09-04 | pretreatment agent | Pretreatment agents are usually acrylate functional monomers with hydrophilic groups, silane coupling agents or other components. It is used for surface treatment of teeth, resins, ceramics, metal restorations, etc., and its chemical modification can be used to change the surface properties of teeth and restorations to improve the bonding force between restorations or in orthodontic treatment. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. Pretreatment agents, orthodontic pretreatment agents, acrylate functional monomers, usually hydrophilic groups, silane coupling agents, or other ingredients. It is used for surface treatment of teeth, resins, ceramics, metal restorations, etc., and uses its chemical modification to change the surface properties of teeth and restorations to improve the bonding force between restorations. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 868
869 | 17-09-05 | Gum retraction line | Cotton threads without vasoconstriction or hemostatic function can be divided into several types according to material, thread diameter, length, etc., and are used to drain the gums when preparing teeth, taking impressions, or cementing crowns. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 869
870 | 17-09-05 | Gum Removal Cream | Paste-like material that does not contain vasoconstrictor or hemostatic properties. Generally composed of aluminum chloride, kaolin and auxiliary materials, etc., it is used to displace the gums when preparing teeth, taking impressions or cementing crowns. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 870
871 | 17-09-05 | Gingival retraction cords (retraction cords with aluminum sulfate/aluminum chloride) | It is usually made of suitable materials such as cotton thread with braid structure, and contains ingredients such as aluminum sulfate or aluminum chloride. It is used to reduce gingival damage and expel gingiva before tooth preparation, making impressions or gluing crowns. The basic principle, scope of application, performance and composition are the same as those of the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions compared with products already on the market. | III |  | 871
872 | 17-09-06 | Intraoral abrasive polishing material | Intraoral grinding and polishing materials are mainly made of silicon dioxide or sodium bicarbonate and auxiliary materials. They are generally powders and pastes, and can also be in other forms. They can be divided into several types according to the composition, abrasive particle size, loading, accessories, etc.; used to grind and polish tooth tissue or restorations in the oral cavity to make the surface smooth and even. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 872
873 | 17-09-06 | Tooth cleaning powder/sandblasting powder | It is used in conjunction with a special sandblasting machine for dentistry, and is usually composed of sodium bicarbonate, silicon dioxide, aluminum hydroxide, calcium carbonate, calcium fluoride, sodium bisulfate, etc. Used to remove plaque, pigment and tartar on teeth. The basic principle, scope of application, performance and composition of the product are the same as the products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 873
874 | 17-09-07 | Agar-based hydrocolloid impression material | The agar-based hydrocolloid impression material is a reversible gel material, which can be divided into several models and specifications according to different technical parameters such as consistency, and is used for tooth and oral tissue impressions. It can be reused if it is used to replicate the model in the laboratory. Product performance indicators adopt the applicable parts of the current versions of the following industry standards, such as: YY 0494 dental agar-based hydrocolloid impression material. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 874
875 | 17-09-07 | Elastomeric impression material | Elastomer impression material is a material that is elastic after being plasticized, such as silicone rubber and polyether rubber. It is usually composed of a matrix, a crosslinking agent, a catalyst, and a filler. It can be divided into several models and specifications according to the material, formula, and impression technical parameters (such as elasticity, chemical stability, and solidification time). After polymerization, it is used for taking impressions of teeth and oral tissues. Product performance indicators adopt the applicable parts of the current editions of the following industry standards, such as: YY 0493 Dental Elastomer Impression Materials. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 875
876 | 17-09-07 | dental alginate impression material | Dental alginate impression material is in powder form, with a uniform appearance and no impurities. The material after reconciliation should be uniform, not agglomerated, not granulated, and the surface is smooth, which can form a smooth plastic body. It is a one-time use product and has no oral cavity. Mucous membrane irritation, no sensitization toxicity, the main components are alginate, calcium sulfate, diatomaceous earth, talcum powder, etc., can be divided into several specifications according to the composition, mixing time, total working time, curing time, etc., for supply For making impressions of teeth and oral tissues. Product performance indicators adopt the applicable parts of the current editions of the following industry standards, such as: YY 1027 Dental alginate impression material. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 876
877 | 17-09-07 | Dental optical powder | Simple powders, usually composed of alcohol, fluorinated hydrocarbons, and mint flavor, are sprayed on the patient's teeth to aid in imaging. It is used to make a negative model that records the shape and relationship of various tissues in the oral cavity, or to assist in obtaining a clear 3D image of teeth. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 877
878 | 17-09-07 | Dental Optical Developing Paste | Products are usually composed of polydimethylsiloxane, barium sulfate, silicon dioxide, etc. Through coated dentures or non-radio-opaque components, under an oral scanner (CT or DVT) in order to obtain a 3D image of the teeth. For denture-guided implant planning and dental diagnosis. | II |  | 878
879 | 17-09-12 | Silicone bite recording material | The main components of silicone rubber occlusal recording materials are silicon dioxide and polydimethylsiloxane, which are used in oral restorative treatment to record the occlusal relationship between the upper and lower dentition. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II |  | 879
880 | 17-09-12 | Denture Fitting Point Indicators | It is usually composed of a substrate, a catalyst and a polymerization inhibitor, and the main components are vinyl silicone rubber, silicon dioxide, etc. It is used for coating on the tissue surface of dental restorations, and after occlusal, it can be used to detect bad contact parts of restorations through curing, discoloration, deformation, etc. Product performance indicators must meet the requirements of relevant national and industry standards, and the exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 880
881 | 17-09-12 | bite wax | Two-component paste or powder liquid or tablet. Usually composed of materials such as silicone rubber or soft plastic. The contained ingredients do not have pharmacological effects, and the contained ingredients cannot be absorbed by the human body. The occlusal examination material is indicative of occlusion by curing, discoloration, deformation, or staining of the portion where force is applied. It is used in oral restorative treatment to record the occlusal relationship of the upper and lower dentition, and according to the recorded results, make a negative model of the oral tissue shape and relationship, or assist in obtaining a clear 3D image of the teeth. | II |  | 881
882 | 17-10-03 | desensitizer | Products are usually in the form of one-component or two-component liquids, pastes, and gels. The main ingredient is fluoride and so on. It is used to relieve the symptoms of dental hypersensitivity caused by exposed dentin. The main components, basic principles, and performance must be the same as those of products already on the market. The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 882
883 | 17-10-04 | Fluorine anti-caries material | It is generally made of fluorine-containing materials. For the prevention of dental caries. Product performance indicators adopt the applicable parts of the current version of the following national standards, such as: YY/T 0823 "Dental fluoride anti-caries materials". The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 883
884 | 17-10-04 | Pit and fissure sealant | Generally made of resin-based materials, it is used to seal the pits and fissures of teeth and prevent the formation of caries in the pits and fissures of teeth. The basic principle, performance and main components must be the same as the products already on the market, and the product performance indicators adopt the applicable parts of the current version of the following national standards, such as: YY 0622-2008 "Dental resin-based pit and fissure sealant". The exemption does not include products that use new materials, active ingredients, new technologies, new designs, or have new mechanisms of action or new functions. | II |  | 884
885 | 17-10-05 | Dental film | Dental diaphragms can be made of suitable materials such as polyethylene, vinyl acetate, edible coloring, single-layer sheet structure or multi-layer sheet structure, which can be divided into several types according to different materials, shapes, sizes, etc., for making Teething pads, sports mouthguard (occlusal) pads, retainers, etc. The exemption does not include products that are different from the marketed products in terms of materials, technology, design or mechanism of action, function, etc. | II | Revised product description. | 885
886 | 18-01-02 | Disposable umbilical cord clipper | Disposable umbilical cord clippers are usually made of polymer materials (such as polyethylene, etc.) Supplied sterile for cutting and sealing newborn umbilical cord stumps. | II |  | 886
887 | 18-01-03 | Disposable cervical forceps | High polymer plastic product, composed of male and female pliers and gill shaft, used to pull the cervix. Sterile, single use. | II |  | 887
888 | 18-01-04 | Disposable sterile umbilical cord clamp | Disposable aseptic umbilical cord clips are made of polymer materials, and can be divided into several models according to different materials, designs, technical parameters, etc., and are provided in sterile form; they are used for closing umbilical cord stumps of newborns. Cord scissors not included. | II |  | 888
889 | 18-01-05 | Disposable sterile vaginal dilators | Disposable sterile vaginal dilators can be made of suitable polymer materials such as polypropylene, mainly composed of upper leaves, lower leaves and handles, and can be divided into several models and specifications according to different materials, types and sizes; the product is sterile Provided in the form; for vaginal examination in obstetrics and gynecology. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0336 Disposable sterile vaginal dilator. | II |  | 889
890 | 18-01-05 | Disposable cervical dilator | Disposable cervical dilation rods are generally made of polymer materials and can be divided into several specifications according to different outer diameters. The products are provided in sterile form. It is used for cervical dilatation before uterine examination in clinical obstetrics and gynecology. | II |  | 890
891 | 18-01-05 | Disposable Cervical Dilation Balloon Catheter | The product consists of a catheter, a balloon, and an insufflation head, and the materials of the catheter, balloon, and insufflation head are polymer materials. This product is a sterile product for single use. This product is primarily used to mechanically dilate the cervix. | II |  | 891
892 | 18-01-06 | Disposable fetal head aspirator | The disposable fetal head aspirator is made of suitable polymer materials such as polypropylene, and is composed of a handle and a suction cup; it can be divided into several models and specifications according to different materials, designs, and sizes; it can be provided in a sterile form; When the cause cannot be delivered normally, the negative pressure generated by the suction cup on the scalp of the fetal head is used to assist vaginal or cesarean delivery. Pneumatic negative pressure pump tire head suction devices are not included. | II |  | 892
893 | 18-01-07 | Disposable gynecological irrigation treatment head | Disposable gynecological irrigation treatment head is made of plastic through injection molding, consisting of a tube body and a joint; it can be divided into several models and specifications according to different materials, technical parameters (such as length, tube diameter, etc.); the product is provided in a sterile form ; Connected to the treatment or flushing handle of the medical ozone flushing therapy device, used to establish a channel for transporting ozone gas or liquid in the vagina. | II |  | 893
894 | 18-01-07 | Disposable vaginal applicator | The disposable vaginal applicator is made of suitable materials such as polypropylene, and can be composed of a cartridge (with graduations) and a push rod; the product is usually provided in a sterile form; it can be divided into several models according to different materials, designs, capacities, etc. And specifications; for female vaginal administration. | II |  | 894
895 | 18-01-08 | Disposable uterine curette | It usually consists of a curette head, a curette neck and a curette handle, and is generally made of polymer materials. Supplied sterile. It is used to extract or remove intrauterine substances by scraping and rubbing. | II |  | 895
896 | 18-01-09 | hysteroscopic lifter | The product is a hysteroscope accessory, which can be composed of swing head, swing head seat, obturator, connecting rod, hook, handle and injection tube and other components. It is used to move the position of the uterus during hysteroscopic or combined hysteroscopic/laparoscopic operations. Supplied sterile. | II |  | 896
897 | 18-01-09 | uterine manipulator | This product is usually a rod-shaped instrument made of stainless steel and polymer materials. Supplied sterile. It is used to change the position of the uterus during abdominal surgery, and can also be used to control and elevate the uterus. | II |  | 897
898 | 18-01-09 | uterine positioning handle | The product consists of handle, trigger, shaft and clip-on joint. Used for aiming the uterus and positioning the uterus, fallopian tubes and ovaries in laparoscopic surgery, and for dye delivery in operations requiring stained fallopian tube perfusion. | II |  | 898
899 | 18-01-10 | Tubal fluid diagnostic instrument | The fallopian tube fluid diagnosis instrument is mainly composed of a control system, a power supply system, a booster system, a printer, etc., and can be divided into several models according to the design, technical parameters, and additional auxiliary functions; it is used for salpingography injection diagnosis and tubal dredging. | II |  | 899
900 | 18-01-10 | Disposable hysterography catheter | Disposable intrauterine fluid contrast tubes can be made of silicone rubber and other polymer materials, mainly composed of balloons, catheters, inflatable joints, and liquid joints; they can be divided into several models and specifications according to materials, lumen outer diameters, etc. ; The product is provided in a sterile form for clinical human hysterography. | II |  | 900
901 | 18-01-10 | Disposable salpingography catheter | Disposable salpingography catheter can be made of plastic, stainless steel wire, etc. It is composed of elbow guide tube, traction rod, inner catheter, guide wire; it can also be made of silica gel and plastic, etc. Catheter base (can include valve and tee), balloon; catheter can be marked with X-ray opaque; can be divided into several models and specifications according to material, design, technical parameters (such as tube diameter, balloon pressure); The product is provided in a sterile form; it is inserted into the fallopian tube through the operation channel of the hysteroscope or other uterine interventional instruments, and the pigment solution or contrast agent is injected to evaluate the patency of the fallopian tube, and the fluid injection is performed to unblock the fallopian tube. | II |  | 901
902 | 18-02-01 | Ultrasound Doppler Fetal Monitor | Ultrasonic Doppler fetal monitor uses the principle of ultrasonic Doppler to continuously monitor the fetus during the perinatal period and provide timely alarm information when abnormalities occur. It usually consists of an ultrasonic probe (generally a plum-blossom probe), a uterine contraction pressure sensor and a host connected to it; it can be divided into several models according to the model, probe configuration, technical parameters, additional auxiliary functions, software processing, and intended use. ; Mainly used for continuous monitoring of perinatal fetal heart rate and maternal uterine contraction pressure. | II |  | 902
903 | 18-02-02 | Ultrasound Doppler fetal heart rate monitor | Ultrasound Doppler fetal heart rate meter is an ultrasonic instrument that obtains fetal heart rate information from the abdomen of pregnant women according to the Doppler principle. It usually consists of a probe (generally a unit probe), an ultrasonic transmitting/receiving circuit, and a signal output part. It can be divided into several models according to the model, probe configuration, technical parameters, additional auxiliary functions, software processing, and intended use; it is used for fetal heart rate measurement. Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0448-2009 Ultrasonic Doppler Fetal Heart Rate Apparatus. | II |  | 903
904 | 18-02-03 | single use uterine probe | Disposable uterine probes are made of polypropylene or applicable materials, and consist of a handle and a thin rod-shaped probe needle body with graduations. It can be divided into several types according to the material, presence or absence of X-ray-opaque marks, needle diameter, and length. Model and specifications; provided in sterile form; used for detecting the direction and depth of the uterine cavity during diagnosis and treatment of the uterine cavity. | II |  | 904
905 | 18-03-02 | optical colposcopy | It usually consists of an observation system and an illumination system. The observation system is a stereoscopic optical microscope system with a short working distance of eyepieces and objective lenses, and can be connected to an external image acquisition and display system. Use the principle of microscopic magnification to observe the details of objects. It is used for observation and imaging of vulva, vagina and cervix. | II |  | 905 | NEW
906 | 18-03-02 | electron vaginal microscope | It consists of an observation system and a lighting system, and can be connected to an external image acquisition and display system. The structural type can be portable, split, or integrated, and is used for magnified observation of the cervix, vagina and vulva. | II |  | 906
907 | 18-03-03 | hysterectomy endoscopy | Generally, it consists of a rigid optical endoscope, a sheath, a manipulator, and surgical electrodes. Under the observation of the endoscope, high-frequency current is used to resect the uterine cavity tissue. | III | Revised product management category description. | 907
908 | 18-03-03 | Hysteroscopy | Rigid optical endoscopes generally consist of an optical imaging system and an illumination system. The optical imaging system consists of three parts: objective lens, transfer system and eyepiece. The lighting system is typically fiber optic. The inverted image of the observed object formed by the objective lens is converted into a positive image through the image conversion system, and transmitted to the eyepiece, and then enlarged by the eyepiece or camera system for observation. For imaging in uterine cavity observation. | II |  | 908
909 | 18-03-04 | Disposable gynecological scraper | The disposable gynecological scraper is made of bamboo, wood, polymer materials, etc. It consists of a handle and a head end. The head end (patient end) is a crescent-shaped smooth flat plate; it can be divided into several specifications according to different sizes; the product is sterile Provided in the form of cervical sampling for gynecological examination. | II |  | 909
910 | 18-03-04 | cervical sample collector | For the collection of cervical epithelial cells, cell clusters and cell masses. Sterile, single use. | II |  | 910
911 | 18-04-01 | Disposable balloon uterine stent | It usually consists of a catheter, a balloon, and a catheter adapter. It is used for intrauterine placement after intrauterine manipulation and surgery to reduce uterine bleeding. The retention time in the human body is less than 24 hours. | II |  | 911
912 | 18-04-02 | Urinary Incontinence Sling / Urethral Sling | Includes a sling and implant tools for implantation in the body. The sling is made of polypropylene material, which can be partly absorbable, and the implant tool consists of a handle and a metal rod. The principle of action is to implant near the urethra to form support for the urethra. For stress urinary incontinence in women due to excessive movement of the urethra and/or sphincter dysfunction. It is equivalent to products that have been approved for domestic registration. Excluding new structural design and drug-device combination products. | III |  | 912
913 | 18-04-02 | Pessary | Generally made of metal or non-metallic materials. Placed in women's vagina (cervix) to support and support. The residence time in the human body is more than 24 hours and less than 30 days. | II |  | 913
914 | 18-04-02 | Pessary | Generally made of metal or non-metallic materials. Placed in women's vagina, cervix, uterus or fallopian tubes to support and support. The residence time in the body exceeds 30 days. | III |  | 914
915 | 18-05-01 | delivery bed | The delivery bed is generally composed of a bed surface, a bed frame, an electrical control system, and accessories. It is used by pregnant women before, during and after childbirth. The performance indicators of this product adopt the applicable parts of the following reference standards: YY0571-2013 Medical Electrical Equipment Part 2: Special Requirements for Safety of Electric Beds in Hospitals. | II |  | 915
916 | 18-05-02 | Electric obstetrics and gynecology operating table | It usually consists of table top, lifting column, base, control system (control surface adjustment), lifting column and motor. According to the transmission principle, it can be divided into three transmission structures: hydraulic, mechanical and pneumatic. Active product. | II |  | 916
917 | 18-06-01 | IUD and Retrieval Devices | Usually hook-shaped, forceps-shaped or looped instruments. Generally made of polymer materials. Supplied sterile. For IUD/vaginal gripping, insertion and/or removal. | II |  | 917
918 | 18-06-03 | Natural Latex Rubber Condoms | Natural latex rubber condoms are made of natural latex, and can be divided into flat type, floating point type, screw type, and special-shaped according to the difference in model; it can be divided into several types according to different added ingredients; it can be divided into several specifications according to different nominal widths; color type Add food-grade pigments, add food flavors for scented types, and add appropriate amount of wet lubricant silicone oil for lubricating types; nominal width, model difference and added ingredients can be combined to a certain extent. Used by men for contraception and protection against sexually transmitted diseases. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 7544 Technical requirements and test methods for natural latex rubber condoms. | II |  | 918
919 | 18-06-03 | Polyisoprene synthetic rubber condom | Made of polyisoprene rubber, it can be divided into flat type, floating point type, screw type, and special-shaped according to the model difference; it can be divided into several types according to the added components; it can be divided into several specifications according to the nominal width; the color type is added Food-grade pigments, food flavors are added for flavor type, and appropriate amount of wet type lubricant silicone oil is added for lubricating type; nominal width, model difference and added ingredients can be combined to a certain extent. Used by men for contraception and protection against sexually transmitted diseases. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 7544 Technical requirements and test methods for natural latex rubber condoms. | II |  | 919
920 | 18-06-03 | female polyurethane condom | Female polyurethane condoms are composed of inner and outer rings and a polyurethane film capsule. The inner ring is closed, and the inner and outer surfaces are coated with dimethyl silicone oil. It can be divided into several types according to different designs and technical parameters; it is used for contraception and prevention by women sexually transmitted diseases. | II |  | 920
921 | 18-06-06 | Ultrasound-guided uterine cavity surgery monitor | Ultrasound-guided uterine surgery monitor consists of a host, a dedicated convex array probe and a monitor. The ultrasound imaging part can be analog or digital imaging. The host can include a negative pressure suction system for abortion and related special accessories. The equipment can be designed according to the design and technical parameters. It is divided into several models according to different functions, auxiliary functions, and accessories; it is used for real-time monitoring of obstetrics and gynecology operations such as artificial abortion, taking and placing birth control rings, and abdominal ultrasound examination. | II |  | 921
922 | 18-06-06 | Disposable Uterine Tissue Suction Tube | Disposable uterine cavity tissue suction tubes are generally made of polymer materials, and can be composed of handles and tube bodies. They can be divided into several specifications according to different tube diameters, and are provided in sterile form for clinical cooperation with aspirators to absorb intrauterine cavity tissue. for organization. | II |  | 922
923 | 18-06-06 | Disposable abortion suction tube | Disposable abortion suction tubes are generally made of polymer materials, consisting of a tube body and joints. The suction tube may have a scale or can be a cannula. It can be divided into several models and specifications according to different materials, designs, and technical parameters; The product is provided in a sterile form; matched with a suction device, it is used for clinical suction method termination of pregnancy. | II |  | 923
924 | 18-06-06 | abortion aspirator | It usually consists of a suction pump, a switch, a safety valve, a check valve, a liquid storage bottle, and a control circuit, and is used in conjunction with suction pipes and abortion suction pipes. It can be used under the supervision of other imaging equipment (such as ultrasonic diagnostic equipment). It is used to perform artificial abortion on pregnant women in early pregnancy. | II |  | 924
925 | 18-07 | Embryo Transfer Catheter Aid | This product is intended to complement and assist the insertion of approved embryo transfer devices into the uterus for the placement of in vitro fertilization (IVF) embryos into the uterine cavity. The product is composed of hard core, catheter and handle etc. | II |  | 925
926 | 18-07-01 | Disposable Embryo Transfer Catheter | The disposable embryo transfer catheter is composed of a guide catheter, a transfer catheter with a Luer connector, and a sheath; it is generally made of polymer materials and stainless steel materials; it can be divided into several models according to materials, designs, and technical parameters. Specifications; provided in a sterile form; matched with a syringe, it is used for transvaginally transferring embryos or fertilized eggs that have undergone in vitro fertilization (IVF) into the uterus. | II |  | 926
927 | 18-07-01 | Disposable Sheathed Embryo Transfer Tube | The disposable sheathed embryo transfer catheter consists of a protective sleeve, a guide catheter and a transfer catheter. It is generally made of polymer materials and stainless steel materials; it can be divided into several models and specifications according to different materials, designs, technical parameters, etc.; it is provided in a sterile form; it cooperates with a syringe and is transferred to the uterus through the vagina after in vitro fertilization (IVF). embryos or fertilized eggs. | II |  | 927
928 | 18-07-01 | assisted reproductive duct | It usually consists of a guide catheter, an outer cannula, a graft catheter, connectors, and a protective sleeve. Generally made of polymer materials. Supplied sterile. The insemination catheter is used to insert into the uterine cavity through the vagina, inject semen, and perform artificial insemination. Embryo transfer catheters are used for the vaginal transfer of embryos that have undergone in vitro fertilization (IVF) into the uterus. | II |  | 928
929 | 18-07-02 | Assisted reproductive puncture egg retrieval/sperm needle collection | It usually consists of a needle tube, a needle sleeve, a needle handle, a hose and a rubber plug interface, a negative pressure connector, a suction tube, a vacuum tube, and a flushing tube. Supplied sterile. For transvaginal puncture of ovaries and aspiration and flushing of oocytes from ovarian follicles, or for percutaneous testicular/epididymal aspiration of sperm. | II |  | 929
930 | 18-07-03 | Assisted Reproductive Microtools | Microtools for assisted reproduction, including microtubular and needle tools for IVF micromanipulation for injection, egg holding, stripping, biopsy, assisted hatching, and sperm retrieval. Also includes vessels and plates in contact with germ cells and embryos. Supplied sterile. For manipulation or storage of human germ cells and embryos in vitro, including sperm microinjection, embryo biopsy and assisted hatching. | II |  | 930
931 | 18-07-04 | Oil for in vitro assisted reproductive culture | Prepared from paraffin oil (mineral oil), it is used to cover culture fluid or culture droplets in in vitro assisted reproductive technology to establish a protective barrier between gametes, zygotes or embryos and the external environment. The product is equivalent to the product that has been approved for domestic registration. New mechanism of action and drug-device combination products are excluded. | III |  | 931
932 | 18-07-04 | Sperm Density Gradient | It is generally composed of silane-coated silica gel particles, inorganic salt components, energy substrates, acid-base buffer components, amino acids, human serum albumin, antibiotics, water and other components. It does not contain components that are not contained in similar listed products in my country. For separation of spermatozoa as a density gradient separation medium. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III |  | 932
933 | 18-07-04 | sperm brake fluid | It is generally composed of polyvinylpyrrolidone, inorganic salt components, energy substrates, acid-base buffer components, amino acids, human serum albumin, antibiotics, water and other components. It does not contain components that are not contained in similar listed products in my country. Used to reduce sperm motility during micromanipulation of sperm. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III |  | 933
934 | 18-07-04 | Granulocyte removal solution | Generally, it is a solution composed of hyaluronidase, inorganic salt components, acid-base buffer components, antibiotics, human serum albumin, water and other components. Used in assisted reproductive technology to remove egg granulosa cells. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III |  | 934
935 | 18-07-04 | Operation solution for in vitro assisted reproduction | Short-term contact solutions for egg retrieval, sperm retrieval, gamete/embryo washing, and manipulation, such as egg retrieval fluid, follicle rinsing fluid, gamete buffer, operating fluid, sperm washing fluid, utensil washing fluid, etc. It does not contain components that are not contained in similar listed products in my country. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III |  | 935
936 | 18-07-04 | Freezing/Thawing Solution for In Vitro Assisted Reproduction | Generally, it is composed of cryoprotectants, inorganic salt components, energy substrates, acid-base buffer components, amino acids, human serum albumin, antibiotics, water and other components. It does not contain components that are not contained in similar listed products in my country. The product is equivalent to the product that has been approved for domestic registration. It is intended to be used for freezing/thawing of human gametes, embryos, and oocytes, including programmed and vitrified freezing/thawing. The product is equivalent to the product that has been approved for domestic registration. Novel mechanisms of action are excluded. | III |  | 936
937 | 18-07-05 | Embryo incubator | It usually consists of an incubator host, a temperature control system, and an air supply control system, which can provide a temperature-controllable (close to human body temperature) environment. Used for culturing gametes, embryos and other cells in assisted reproductive technology (ART). | II |  | 937
938 | 18-07-05 | jet lag incubator | It usually consists of an incubator host, a temperature control system, and an air supply control system. Able to provide a temperature-controlled (closer to human body temperature) environment. The time-lapse incubator may have a time-lapse camera system, image capture and analysis application software, server, client, etc. It is used for culture, dynamic monitoring and analysis of cells, gametes and embryos in assisted reproductive technology (ART), and to evaluate the developmental ability of embryos. | II |  | 938
939 | 18-07-05 | Constant temperature platform for assisted reproduction, program freezer, program-controlled cooling device, in vitro fertilization (IVF) ultra-clean workbench | It usually consists of operating table, heating table or controller, freezing device, etc. | II | Revised classification codes. | 939
940 | 19-01-07 | Cassette Hearing Aids | Cassette hearing aids are mainly composed of input transducers, signal conditioning units, output transducers and power supplies. Most of them are analog hearing aids. , for hearing compensation for patients with air conduction hearing loss. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 14199-2010 General Specification for Electroacoustic Hearing Aids. | II |  | 940
941 | 19-01-07 | behind-the-ear hearing aids | Behind-the-ear hearing aids are mainly composed of input transducers, signal conditioning units, output transducers and ear hook shells, etc., and can be divided into several models according to design, technical parameters, intended use, additional auxiliary functions, etc. (analog / digital / adjustable Programming, etc.); experience matching, for hearing compensation for patients with air conduction hearing loss. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 14199-2010 General Specification for Electroacoustic Hearing Aids. | II |  | 941
942 | 19-01-07 | In-Ear Hearing Aids | In-the-ear hearing aids include concha cavity and ear canal hearing aids, which are mainly composed of input transducers, signal conditioning units, output transducers and power supplies, and can be customized according to design, technical parameters, intended use, additional auxiliary functions, etc. It is divided into several models (mostly digital programmable hearing aids now); it is equipped with experience and is used for hearing compensation for patients with air conduction hearing loss. The product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 14199-2010 General Specification for Electroacoustic Hearing Aids. | II |  | 942
943 | 19-02-01 | Gait training equipment | It usually consists of a weight-reducing device, a host machine, a treadmill, a control device, a fixing device, etc. Promote rehabilitation by training the patient's gait, and can be equipped with a gait assessment function. It is used for gait rehabilitation training for patients with walking impairment of lower limbs. Does not include limb linkage equipment, does not include exoskeleton rehabilitation training equipment. | II |  | 943
944 | 19-02-02/05 | Physical rehabilitation training equipment | It usually consists of a bed frame, straps, mechanical support components, electric control devices, fixed protection devices, etc. The patient is trained to promote rehabilitation by changing the body position and standing angle, and it is used for limb movement rehabilitation training for patients with limb motor dysfunction caused by cerebral apoplexy, traumatic brain injury, and cerebral apoplexy. Does not include bioelectrical signal acquisition, feedback and electrical stimulation functions. Does not include limb linkage equipment, does not include exoskeleton rehabilitation training equipment. | II |  | 944
945 | 19-02-03 | balance training equipment | It usually consists of a measurement platform, an auxiliary bracket, and balance training software. Balance training, usually with patients standing or sitting on a test platform, may be accompanied by a balance assessment. It is used for rehabilitation training for patients with balance impairment. Does not include limb linkage equipment, does not include exoskeleton rehabilitation training equipment. | II |  | 945
946 | 19-02-04 | vibration training equipment | It usually consists of training platform, control device, fixed frame and so on. Through periodic mechanical vibration, the purpose of muscle or joint rehabilitation is achieved. It is used to improve muscle function, balance and coordination in patients with motor impairment. Does not include limb linkage equipment, does not include exoskeleton rehabilitation training equipment. | II |  | 946
947 | 19-02-05 | Joint continuous passive motion instrument (CPM instrument) | Joint continuous passive movement instrument (CPM instrument) is an active device, which is usually composed of a host, a control part and a joint fixation part. It can be classified into several models according to different fixed joints, technical parameters, additional functions, etc.; through the electromechanical structure, the patient's corresponding joint Repeated flexion and extension, continuous passive movement can stimulate the joints, so that the physiological synovial fluid can produce a virtuous circulation, reduce joint swelling, and prevent or treat the reduction in the range of motion of the joints and joint stiffness caused or possibly caused by trauma or lack of exercise. and venous thrombosis. | II |  | 947
948 | 19-03-01 | manual wheelchair | Manual wheelchairs are driven by human power and can be composed of frame, front wheels, rear wheels, seat, back pad, armrests, hanging feet, brakes, pedals, etc., according to the driving mode, wheelchair material, configuration, function, size, applicable Different ranges are divided into several models, which are used as transportation tools for the disabled, the sick, and the elderly and infirm who have difficulty in moving. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB/T 13800-2009 Manual wheelchairs. | II |  | 948
949 | 19-03-01 | electric wheelchair | An electric wheelchair is a two-wheeled personal mobility device with a supporting chair support system, and one or more electric motors provide driving force. The speed of travel is controlled by a motor, and the direction of travel is manually or electrically controlled by an accompanying person or passenger; it can be divided into several models according to different materials, functions, accessories, technical parameters, etc.; ) for transportation. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 12996-1991 Electric Wheelchairs and GB/T 18029 Wheelchair Electric Wheelchair Series. | II |  | 949
950 | 19-04 | nasal septum clip | It is inserted into the nasal cavity after nasal cavity surgery to fix the nasal septum. | II |  | 950
951 | 19-04-01 | Orthotics | It is usually made of polymer materials, fabrics, metals and other materials. Worn on the head, torso or extremities to correct or prevent deformities. | II |  | 951
952 | 20-02-04 | Cupping equipment | It usually consists of electric negative pressure source, catheter, tank and so on. Negative pressure is generated in the tank through a negative pressure source, thereby adsorbing on the muscles. | II |  | 952
953 | 20-02-05 | Fumigator | The fumigation instrument is composed of liquid medicine tank, liquid inlet/outlet pipeline, heating device, temperature/liquid level and other control units, display unit, treatment head, etc. It is divided into several models according to different auxiliary functions; it can be used in conjunction with liquid medicine for partial fumigation of the human body. | II |  | 953
954 | 20-02-05 | fumigation bed | The fumigation bed can be composed of a bed body, a control unit, a medicine tank, a temperature controller, a liquid level controller, a heating device, a liquid crystal display, etc. It is divided into several models; it can be used with liquid medicine for fumigation treatment of human trunk and other parts. | II |  | 954
955 | 20-03-01 | acupuncture needles | Acupuncture needles are composed of a needle body and a needle handle. The needle body material can be made of suitable materials such as austenitic stainless steel. The needle handle can be made of stainless steel wire (tube) or copper wire (tube) or aluminum wire (tube) or polypropylene, etc. According to whether it is brought into the needle tube, the type of needle handle (such as ring handle needle, flat handle needle, flower handle needle, plastic handle, etc.), needle body diameter and length, it is divided into several types and specifications; it can be sterilized or non-sterilized. Supplied in a sterilized state; it is used by professionals to invasively stimulate acupuncture points (including auricular points) on the human body at different depths. Product performance indicators adopt the applicable parts of the following reference standards, such as: GB 2024-1994 Acupuncture Needles. | II |  | 955
956 | 20-03-03 | small needle knife | Small needle-knife therapy is a kind of closed lysis between surgical method and non-surgical therapy. Its operation is characterized by piercing deep into the lesion to cut at the treatment site, stripping harmful tissues, so as to relieve pain and relieve pain. disease purpose. A small needle knife is a kind of acupuncture tool that looks like a needle and a knife. It is divided into three parts: the handle, the needle body, and the needle knife. Usually, the knife is the same width as the needle and has a sharp edge. It can be divided into several parts according to the shape and length. Models and specifications; can be provided in sterile/non-sterile form; for soft tissue injury lesions and osteoarthritis lesions release under direct vision. | II |  | 956
957 | 20-03-04 | skin needle | It usually consists of a dial, a needle body, a needle tip and a needle handle. It looks like a small hammer, with a lotus-shaped dial at one end, and short stainless steel needles are regularly embedded under the dial. Depending on the number of needles, they are called plum blossom needles (five needles), Qixing needles (seven needles), and Arhat needles (ten needles). For tapping acupuncture points and other parts of the skin. Supplied sterile . | II | Add | 957
958 | 20-03-05 | needle rolling | It usually consists of brackets, rollers, stainless steel needles, handles, etc. For local stimulation of specific parts of the body surface, needle rolling therapy is implemented. Supplied sterile . | II | Add | 958
959 | 20-03-06 | press needle | Press needles are usually composed of a needle body and a dial, and can also be pasted on a conformable applicator to facilitate indwelling. The needle body and needle seat are mostly made of austenitic stainless steel wire and other suitable materials after grinding, polishing, and winding. It is divided into several models and specifications according to its use, structure, needle diameter, and length; most of them are provided in a non-sterile state; they are used by professionals for superficial invasive stimulation of human body surface acupuncture points (including auricular points). Product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0105-1993 push pin. | II |  | 959
960 | 20-03-06 | intradermal needle | The intradermal needle is composed of two parts: the needle tip and the needle disc. It can also be pasted on a conformable applicator to facilitate indwelling on the body surface. Different types of structure, needle diameter, length, etc. are divided into several models and specifications, which can be provided in sterilized and non-sterilized states, and are used by professionals for superficial invasive indwelling stimulation of human body surface acupoints (including ear acupoints). The product performance index adopts the applicable parts of the following reference standards, such as: YY 0105 Press needle. | II |  | 960
961 | 20-03-07 | Disposable embedding needle | It usually consists of a lining seat, a needle seat, a needle tube, a lining core and a protective sleeve. It is used for puncture and embedding of acupuncture points. | II |  | 961
962 | 20-03-10 | Disposable floating needle | It usually consists of needle core, needle seat, hose and protective sleeve. for floating acupuncture | II |  | 962
963 | 21-00-00 | trace analysis software | The blot analysis software is used to save the membrane strip information scanned by the flatbed scanner as an electronic picture, identify the position of the membrane strip, and identify the specificity of the strip. Does not include auto-diagnostics. | II |  | 963
964 | 21-00-00 | Clinical Management Software | It is used for the management of ECG and obstetric diagnosis and treatment activities . | II | Revised product description. | 964
965 | 21-00-00 | Clinical Management Software | It is used for the management of tumor, critical care and anesthesia diagnosis and treatment activities . | III | Revised product description. | 965
966 | 21-01-02 | Radiation therapy recording and verification system software | It is used for the verification of treatment parameters before radiotherapy and the record of treatment parameters during treatment. | III |  | 966
967 | 21-01-02 | radiation therapy contouring software | It is used to outline tumors and related organs and tissues before making radiation therapy plans. | III |  | 967
968 | 21-02-01 | Medical image storage and transmission system software | For the transmission, display, output and storage of medical images. | II |  | 968
969 | 21-02-02 | Medical Image Processing Software | It is used to process medical images from single modality or multimodality. | II |  | 969
970 | 21-03 | data processing software | It is used for the transmission of monitoring data, centralized real-time display, alarm or the transmission, display and processing of physiological data. | II |  | 970
971 | 22-01-02 | Blood cell analysis instrument | It usually consists of blood cell detection module, hemoglobin measurement module, mechanical module, electronic module, computer system and so on. The principle is generally electrical impedance method, colorimetry, flow laser scattering technology, etc. It is used for quantitative and qualitative analysis of formed components in blood/body fluids and provides relevant information. | II |  | 971
972 | 22-01-04 | Coagulation analysis instrument | It usually consists of a preheating module, a sample adding module, a timing module, a sample delivery and processing module, a detection module, and a computer system. The principle is generally coagulation method, chromogenic substrate method and immunoturbidimetric method. It is used to analyze the functions of coagulation and anticoagulation, fibrinolysis and antifibrinolysis of blood. | II |  | 972
973 | 22-01-05 | Platelet Analysis Instrument | It usually consists of a liquid circuit module, a sample processing module, a detection module, and a computer system. The principle is generally turbidimetry. It is used to analyze related functional parameters such as platelet quantity, volume and aggregation rate in blood samples. | II |  | 973
974 | 22-01-06 | Hemorheometer/Viscometer | Generally divided into capillary viscometer and rotational viscometer. A capillary viscometer is usually composed of a capillary, a sample cell, a temperature control device, a drive device, a timer, etc.; a rotational viscometer is usually composed of a sampling module, a sample sensor, a speed control and adjustment module, a torque measurement module, and a constant temperature module. The principle is generally Poiseuille's law or the law of viscosity. It is used clinically to analyze the rheological properties of whole blood, plasma or blood cells. | II |  | 974
975 | 22-01-08 | Flow Cytometry Instruments | It usually consists of a flow chamber and a liquid flow system, a laser source and an optical system, a photoelectric cell and a detection system, a computer and an analysis system. The principle is generally to allow a single particle to pass through the flow chamber through a liquid flow system and analyze the fluorescent labeling signal of the single particle to realize the analysis of its biological characteristics. | II |  | 975
976 | 22-02-01 | Biochemical Analysis Instruments | It usually consists of one or more of samplers, sampling devices, reaction pools or reaction pipes, incubators, detectors, microprocessors, etc. The principles are generally spectrophotometry, turbidity colorimetry, ion selective electrode method, fluorescence method, reflection photometry, differential potential method, etc. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 976
977 | 22-02-02 | Blood glucose and blood glucose related parameters analysis instrument | It usually consists of host module, power module, software module and so on. The principle is generally electrochemical method, light reflection technology, colorimetric method and so on. Does not include blood collection equipment and suitable reagents. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 977
978 | 22-03-01 | Electrolyte Analyzer | It usually consists of an electrode module, a measurement module, a pipeline module, a circuit module and a data output module. The principle is generally the ion selective electrode method. Used to analyze the electrolyte content in blood and body fluids. | II |  | 978
979 | 22-03-02 | blood gas analyzer | It usually consists of an electrode module, a measurement module, a pipeline module, a circuit module, and a data output module. The principle is generally ion selective electrode method and microelectronics and biochip technology. It is used to measure blood gas parameters such as pH, partial pressure of carbon dioxide and partial pressure of oxygen of blood and body fluids. | II |  | 979
980 | 22-03-03 | Electrolyte blood gas analyzer | It usually consists of an electrode module, a measurement module, a pipeline system circuit module and a data output module. The principle is generally the ion selective electrode method. | II |  | 980
981 | 22-03-04 | Detection electrode for electrolyte, blood gas and biochemical analysis | It usually consists of single or multiple electrode blocks. It is used in conjunction with electrolyte analyzers, blood gas analyzers and biochemical analyzers with electrolyte modules for analysis of electrolytes, blood gases or metabolites. | II |  | 981
982 | 22-04 | Automatic immune analyzer | It is usually composed of sampling center, processing center, waste liquid and supply center, system control center, etc., based on the immunological reaction of antigen-antibody binding, using enzyme labeling or chemiluminescence to label antigen-antibody, through a series of cascade amplification The reaction is to link the optical signal or electrical signal with the concentration of the analyte, etc., and analyze the antigen or antibody to be tested in the human sample. It is used for quantitative, semi-quantitative or qualitative detection of various analytes in human body fluids, such as tumor markers, pathogen antigen antibodies, etc. | II |  | 982
983 | 22-04 | Allergen semi-quantitative analysis software | The allergen semi-quantitative analysis software contains a standard curve, which can be used with ordinary scanners. Test strip for the semi-quantitative detection of allergen-specific antibodies and autoimmune antibodies in immunoblotting and immunochromatography. Does not include auto-diagnostics. | II |  | 983
984 | 22-04 | Automatic Immunoluminescence Apparatus | It usually consists of one or more of the sample loading module, reaction module, optical detection module (photomultiplier tube), data processing module, incubation temperature control module, cleaning and separation module, etc. The principle is generally to convert the light signal emitted by the chemiluminescent reaction into a digital signal, and the concentration value is calculated by the data processing system. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 984
985 | 22-04-01 | Enzyme-linked immunoassay instrument | It usually consists of one or more of the transmission module, reagent filling module, incubation module, optical module, cleaning module and data processing module. The principle is generally that after the monochromatic light is absorbed by the specimen, the optical signal is converted into an electrical signal by the photodetector, and the concentration value is calculated by the data processing system. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 985
986 | 22-04-02 | Chemiluminescence Immunoassay Analyzer | It usually consists of one or more of the sample loading module, reaction module, optical detection module (photomultiplier tube), data processing module, incubation temperature control module, cleaning and separation module, etc. The principle is generally to convert the light signal emitted by the chemiluminescent reaction into a digital signal, and the concentration value is calculated by the data processing system. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 986
987 | 22-04-03 | Fluorescence Immunoassay Instrument | It usually consists of one or more of the sample loading module, reaction module, optical detection module (fluorescence), data processing module, incubation temperature control module, cleaning and separation module, etc. The principle is generally to convert the fluorescent signal into a digital signal, and the concentration value is calculated by the data processing system. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 987
988 | 22-04-04 | Colloidal gold test paper analyzer | It usually consists of a photoelectric detection module, a mechanical scanning control module, a control motherboard module, and an information collection module. The principle is generally to convert the reflectance signal of the detection reagent card into a photoelectric signal through the sensor, and convert the photoelectric signal into the corresponding concentration value or threshold value through the calibration information, and analyze the object to be tested. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 988
989 | 22-04-05 | Western blot | It usually consists of a peristaltic pump module, a sampling module, an incubation module, and a temperature control module. The principle is generally to transfer the protein to the solid-phase membrane by means of electrotransfer and other methods, and finally perform immunological detection. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 989
990 | 22-04-06 | Specific protein immunoassay analyzer | It usually consists of an optical module, a detection module, a computer system, etc. The principle is generally immunoturbidimetric method. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 990
991 | 22-05-03 | Nucleic Acid Amplification Analysis Instrument | It usually consists of control components, thermal cover components, thermal cycle components, photoelectric components, transmission components, embedded software and analysis software, power supply components, etc. The principle is generally to use temperature control to provide a suitable environment for the in vitro amplification of nucleic acids, and to collect and analyze the optical and electrical signals generated during the amplification process. Used in conjunction with adapter reagents, it is used for in vitro amplification and analysis of nucleic acid of sample genes. | III |  | 991
992 | 22-05-04 | PCR amplification instrument | It usually consists of control components, thermal cover components, thermal cycle components, optoelectronic components, transmission components, embedded software and power supply components. The principle is generally to use temperature control to provide a suitable environment for the in vitro amplification of nucleic acids. Used in conjunction with adapter reagents for in vitro amplification of nucleic acid from sample genes. | II |  | 992
993 | 22-05-05 | Nucleic acid molecular hybridization instrument | It usually consists of a temperature control module and a control panel module. The principle is generally the principle of base complementarity. For hybridization of nucleic acid molecules. | II |  | 993
994 | 22-06-02 | Microbial culture monitoring instrument | It usually consists of an incubation module, a detection module, a control/alarm module, a display module, and random software modules. It can also include air filtration, barcode scanning, etc. The principle is generally to determine the concentration of bacteria suspended in liquid culture medium by measuring light scattering, optical density, electrical impedance, pressure sensing, color production (carbon dioxide) or direct counting of bacteria . It is used for clinical culture and detection of aerobic bacteria, anaerobic bacteria, fungi and mycobacteria in specimens such as blood and body fluids. | II |  | 994
995 | 22-06-06 | Microbial Identification Drug Sensitivity Analysis Instrument | It usually consists of one or more of an automatic inoculator module, an incubation module, an identification module, a drug susceptibility analysis module, a computer module, a barcode reader module, and a random software module. The principle is generally to identify infectious and/or pathogenic microorganisms isolated from biological samples (such as: blood, urine, cerebrospinal fluid, saliva or feces) by means of morphology, growth, physiology and clinical chemistry. Drug susceptibility part: determine the drug susceptibility of bacterial pathogens isolated from clinical samples by cooperating with reagents containing different concentrations of antimicrobial agents. For the identification and/or drug susceptibility analysis of clinically isolated microorganisms. | II |  | 995
996 | 22-06-07 | Bacterial endotoxin/fungal dextran detection instrument | It usually consists of optical channel module, automatic thermostat module and temperature control module. The principle is generally gel method, photometry and so on. For the detection of bacterial endotoxin and/or (1,3)-β-D glucan. | II |  | 996
997 | 22-06-08 | 13C, 14C breath analyzer | It usually consists of a sampling module, an air purification and drying module, an optical module, a measurement module, an electrical control module and a gas collection module. The principle is generally the breath test detection method. For clinical diagnosis of diseases caused by Helicobacter pylori infection. | II |  | 997
998 | 22-07-02 | image scanning instrument | It usually consists of optical imaging system, image acquisition system, computer , software and so on. The principle is generally to present cells on slides or slices as scanned images in a photographic scanning manner . It is used to scan and observe microscopic images of clinical samples . | II |  | 998
999 | 22-07-03 | Image Analysis Instrument | It usually consists of optical imaging system, image acquisition system, computer, image analysis software, etc. The principle is generally to present the cells on the slide as images by means of photographic scanning, and to classify, mark and analyze the cells on the images. It is used to observe, screen, mark and analyze microscopic images of clinical samples. | II |  | 999
1000 | 22-08-02 | liquid scintillation counter | It usually consists of a detection module, a measurement module and a computer system. The principle is generally to use a liquid scintillator (scintillation fluid) to receive rays and convert them into fluorescent photons. It is used to measure the beta radioactive rays emitted by radionuclides such as tritium and carbon 14. It is clinically used in conjunction with radioimmunoassay kits for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 1000
1001 | 22-09-01 | Dry Chemistry Urine Analyzer | It usually consists of mechanical modules, optical modules, circuit modules, etc. The principle is generally reflective photometry. Used in conjunction with suitable reagents to measure urine protein, glucose, urine pH, ketone bodies, urobilinogen, bilirubin, nitrite and other parameters. | II |  | 1001
1002 | 22-09-02 | Urine Formed Components Analyzer | Generally divided into flow cytometry and imaging. Flow cytometry usually consists of an optical detection module, a liquid flow module, an electrical impedance detection module and a circuit module. The principle is generally flow cytometry. The imaging system usually consists of a specimen processing module, an optical counting pool module, a microscopic camera module, and a data processing module. The principle is generally the principle of automatic identification of digital imaging. For the identification and analysis of formed components in urine. | II |  | 1002
1003 | 22-09-04 | stool analysis workstation | It usually consists of a sample processing module, a microscope module, and a result processing software module. The principle is generally chemical method, immunological method and microscopy method. It is used for the detection of formed components, occult blood and pathogenic microorganisms of stool samples. | II |  | 1003
1004 | 22-09-05 | Sperm Quality Analyzer | It usually consists of a microscopic image scanning module, a temperature control system module, a counting pool module, a computer system, software, etc. For the analysis of sperm quality. | II |  | 1004
1005 | 22-10-02 | Trace element analyzer | It usually consists of a host module, a micro-analysis workbench module, software, and a computer system. The principles are generally electrochemical analysis methods, potential stripping methods, atomic absorption methods, and mass spectrometry. It is used to detect various trace elements in blood, urine, hair and other samples. | II |  | 1005
1006 | 22-10-04 | Glycated hemoglobin analyzer | It usually consists of host module, power module, software module and so on. The principle is generally electrochemical method, light reflectance technology, colorimetric method, immunoturbidimetric method, etc. Does not include blood collection equipment and suitable reagents. It is used in conjunction with suitable reagents for qualitative and/or quantitative analysis of analytes in human samples. | II |  | 1006
1007 | 22-10-05 | Glycosylated hemoglobin chromatography column | It usually consists of columns and fixed parts. It is used in conjunction with analysis equipment to separate the analytes in human samples. | II |  | 1007
1008 | 22-10-06 | Freezing Point Osmometer | Generally divided into freezing point osmotic pressure measurement and colloid osmotic pressure measurement. The freezing point osmolarity tester is usually composed of a refrigeration module, a stirring module, a temperature sensor module and a computer system; the principle is generally the crystallization of the vibration principle. The colloid osmotic pressure measuring instrument usually consists of a semipermeable membrane and its fixture, a reference liquid chamber, a voltage sensor and a computer system. The principle is generally the principle of osmosis. Used to measure crystal osmotic pressure and colloid osmotic pressure of samples such as urine and blood. | II |  | 1008
1009 | 22-10-08 | Biochip Analysis Instrument | It usually consists of a host module, a photoelectric signal collector module, and a computer system. The principle is generally to collect the optical and electrical signals on the biochip and analyze them through software. It is used in clinical laboratories for qualitative, semi-quantitative or quantitative testing of various medical testing items. | II |  | 1009
1010 | 22-11-01 | Arterial and venous blood collection needles and connectors | It usually consists of an arteriovenous blood collection needle, a blood collection device, a protective cover and other components. For collecting arterial and venous blood samples. If the product has special properties or structure, it cannot be exempted from clinical practice. The exemption does not include products using new materials, new mechanisms of action, or new functions. | III |  | 1010
1011 | 22-11-02 | Peripheral blood collection needle | It usually consists of needles, needle handles, protective sleeves, etc. An activation device (such as a spring, etc.) may be included. Supplied sterile. Single use. It is used for skin puncture in clinical medicine to collect peripheral blood samples from human body. | II |  | 1011
1012 | 22-11-04 | Disposable vacuum blood collection container | Disposable vacuum blood collection containers are usually composed of a tube and a head cap. The inner wall of the tube is attached or not with additives or additives. It is used in conjunction with a disposable blood collection needle for the collection, transportation and storage of human venous blood. The product performance indicators adopt the applicable parts of the following reference standards, such as: YY 0314 "Disposable Venous Blood Sample Collection Container". | II |  | 1012
1013 | 22-11-05 | Peripheral blood collection tube | It usually consists of capillaries, straws, connectors, etc. When provided sterile, the inner wall of the tube may be provided with or without additives; when provided non-sterile, the inner wall of the tube may have additives. It is used for the collection and storage of human peripheral blood. | II |  | 1013
1014 | 22-11-06 | Peripheral blood collection container | Usually consists of a container (tube or bottle or vial or test tube), a cap and additives. It is used for the collection, transportation and storage of human peripheral blood samples. | II |  | 1014
1015 | 22-11-08 | Disposable occult blood sampling capsule | Disposable occult blood sampling capsule is composed of medical hollow capsule, medical absorbent cotton and cotton thread, which is used to extract gastric juice for occult blood examination. | II |  | 1015
1016 | 22-11-09 | Sterile sample sampling swabs, disposable sterile microbial swabs, disposable sterile sampling swabs | Usually composed of a handle and a sampling head, the swab can be divided into different models and specifications according to the material, design, size, etc., and can be equipped with a sampling container; usually provided in a sterile form; for the natural cavity of the human body, such as the nasal cavity, oral cavity, It is used for testing biological samples from the vagina, urethral opening and other parts. | II |  | 1016
1017 | 22-11-11 | heel lancing device | It usually consists of a spring, a blade, a pop-up structure and a casing, and the blade is generally made of stainless steel. Supplied sterile. It is used for heel pricking of premature infants or newborns. | II |  | 1017
1018 | 22-14-01 | Medical culture/incubator | It usually consists of temperature and humidity, gas concentration control system, electronic display system, box and so on. For the cultivation of samples of human origin. | II |  | 1018
1019 | 22-14-02 | Anaerobic culture system | It usually consists of sampling room, operation room, anaerobic tank, cultivation room and so on. For the cultivation of anaerobic and facultative anaerobic microorganisms. | II |  | 1019
1020 | 22-15-03 | Automatic sampling system | It is usually mainly composed of a precision sampling system, which may include other functional connections such as a transmission system, a cleaning system, an incubation system, a mixing system, and a software system. It is used for precision adding of reagents or samples before analysis of clinical laboratory analysis instruments. | II |  | 1020
1021 | 22-15-04 | cryogenic storage equipment | It usually consists of refrigeration device, heat insulation box, electric control mechanism and other components. For cryogenic storage or transport of tissues, cells, blood and blood products, etc. | II |  | 1021
1022 | 22-15-04 | Organ Refrigerator | Passive appliances, usually composed of ice cubes and other refrigerants and boxes, are used for refrigerated transport of human organs. | II |  | 1022
1023 | 22-16-01 | biological safety cabinet | The product generally consists of a cabinet body, front window operation opening, supporting legs and casters, fan, liquid sump, filter, control panel, ultraviolet lamp and lighting source. The operation port of the front window is required to flow into the airflow with the lowest average flow rate. The descending airflow is filtered by a high-efficiency filter and then sent to the work area. The polluted airflow can be discharged to the next link after being filtered by a high-efficiency filter. All biologically contaminated parts in the safety cabinet are under negative pressure. pressure state or surrounded by negative pressure channels and negative pressure ventilation system. The biological safety cabinet is a safety cabinet with a front window operation port. The operator can operate in the safety cabinet through the front window operation port to protect personnel, products and the environment during the operation. | III |  | 1023
1024 | 22-16-02 | clean bench | It consists of a cabinet, fan, high-efficiency filter, casters, lighting, ultraviolet germicidal lamp and control system. The cleanliness of the working area of this product can reach ISO level 5, which is used for clinical laboratory tests and experiments, so that the local operating environment can reach a certain level of cleanliness. | II |  | 1024
1025 | 通用02-14-01；眼用06-01-07；口腔17-04-14 | Disposable Surgical Irrigation Needle | Disposable flushing needles for surgery are made of suitable polymer materials and metal materials, and are composed of a needle seat and a flat-headed needle tube. They can be divided into several models and specifications according to different materials, capacities, and needle diameters; they are provided in sterile form; and The syringe is used as a complete set for washing wounds in microsurgery, ophthalmology, stomatology, plastic surgery and other operations. | II |  | 1025